## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2004-13

Agency: IOS, Ocean Science and Productivity, Sidney, B.C.

Location: Gulf of Alaska / Bering Sea / Chukchi Sea

Project: Joint Western Arctic Climate Study

Party Chief: van Hardenberg B. / Grebmeier J.

Platform: Sir Wilfrid Laurier

Date: 8 July 2004 –22 July 2004

Processed by: Germaine Gatien

Date of Processing: 17 August 2005 – 25 August 2005

Number of original CTD casts: 69 hex files, 56 bl files, 48 btl files

Number of casts processed: 51 CTD files, 44 CHE files

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (#0293) was mounted with Transmissometer 139, SBE 43 Dissolved Oxygen Sensor S/N #0052 and Seapoint Fluorometer S/N #2336. The deck unit type and S/N are unknown. The pressure sensor was moved before cast #10. The fluorometer was unpumped. There was a touch-down sensor.

# SUMMARY OF QUALITY AND CONCERNS

No recalibration has been applied to the salinity data. While the CTD is lower than bottles by an average of 0.002psu, there was a lot of scatter in the comparison. The post-cruise calibration indicates no significant drift had occurred in either temperature or salinity.

It is recommended that salinity calibration samples not be taken in deep water in areas of high turbidity. During this cruise and cruises in the Strait of Georgia it has been observed that such samples are often outliers in comparison to the CTD. It is believed that the presence of mud in the samples affects the titration.

The file names were non-standard with the event number concatenated with a station name. These all had to be changed to standard format. When there were several files created at one site, the same event number was assigned to each of them, the only difference in name being something following the station name. Separate event numbers should be used so that standard names can be assigned.

The daily log book contained no information about the crew and equipment. This led to some delays in processing since it was not clear which channels needed to be converted.

Fluorescence:URU:Seapoint – The data are uncalibrated and unedited and frequently off-scale. There was no chlorophyll sampling data available.

Oxygen:Dissolved:SBE –The DO data looked bad. The problems varied through the cruise and are detailed in section 19. This data will not be archived.

The pressure sensor is considered ±1db. An offset of +0.6db was applied to all data during conversion.

The salinity calibration is considered ±0.005psu except in areas of rapid temperature change where the salinity has been edited heavily and should be considered ±0.02psu.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

The file names had to be changed to standard format. In some cases there were two or more files for a cast. One was named as usual and the other was given a “9”, “8”, “7” or “6” in the 1st digit of the event number. Later it will be determined which file should be used or if several are needed.

The Daily Log was obtained; no rosette log sheets were available. There is a note at the end of the log indicating that the internal clock was about 11 minutes ahead of GPS time. Problems were noted with the data from the DO channel.

The equipment was not listed in the Daily Log. There was a transmissometer entered in the CON file, but no serial number is given. However, the calibration co-efficients are the same as those listed for #197 in 2003, so it is assumed that is the serial number.

Salinity bottle data were available.

The configuration file used at sea contained pre-cruise calibrations. There was a post-cruise calibration which indicates no significant drift in either temperature or salinity. The con files include 3 User Polynomial channels, a pH channel, as well as dissolved oxygen, transmissometer and fluorometer channels. It is believed that one of the User Poly or pH channels contains the bottom sensor data, but the others are not believed to contain any signal.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

There was a change of configuration before cast #10 when the DO sensor was moved.

File 2004-13-CTD1.CON (a copy of the con file used for cast #1) was used to convert the hex files for casts 1-9.

File 2004-13-CTD2.con (a copy of the file used for cast #10) was used for all other casts.

There are 3 userpoly channels in the con files; one looks like a bottom sensor, but the others seem to have some signal, but nothing recognizable – this may just be noise. The bottom sensor channel will not be converted since it is not needed for archiving. None of the unidentified channels will be converted since Bon van Hardenberg says no other instruments were mounted on the CTD.

Three of the hex files were empty when converted, all were extra files from event #30, and there were no con files associated with them. File 2004-13-9019 did not produce a bottle file. The station name is TEST and the data in the CNV file is from the surface only so it will not be processed further.

The temperature, conductivity, transmissivity and fluorescence look reasonable. The dissolved oxygen does not. For the first 9 casts it looks like the inverse of the temperature channel. Thereafter, it looks like the oxygen voltage in shape, but not in values. I tried stripping the DO values and re-deriving them, but that did not help. Both oxygen channels will be kept and examined later.

A study was then made of the surface pressure to determine if an offset was needed. Ten casts were examined and the surface pressures were estimated by when conductivity, fluorescence and transmissivity went from out-of-water to in-water values. This is a very rough estimate. For both downcast and upcast all surface pressures were less than 0. The value ranged from -0.2db to -1.2db. The average for the downcast surface is -0.6db. Since the sensor is considered ±1db, these are within the expected error. But since all values are negative, it seems reasonable to apply an offset of +0.6db. The same offset was applied when this CTD was used in 2003. The con files were adjusted and the files were reconverted with that offset applied.

Two casts deeper than 500db were checked for hysteresis and the differences between downcast and upcast surface pressures were 0.23db and 0db, so there is no problem.

Rosette files were converted taking the scan range data from the BL files and using a 5s window around firing time. The files that were named 2004-13-6030, 7030 and 8030 could not be converted; those were the files for which there were no configuration files.

A few casts were checked and all expected channels look ok. The descent rate is generally quite steady and high, but there are a few casts during which it is noisy. The files were put through CLEAN to add an event number based on the file name. (Output: BOT)

4. WILDEDIT

This step was be skipped as it has not been found to be useful in the past and there is no evidence of isolated “wild” points. If problems are noted later a return could be made to this stage for the casts in question.

5. WFILTER

Cosine filters were run on pressure (window size 5), temperature (window size 5) and conductivity (window size 5) based on the results of 2001 Arctic cruises which used an SBE25. Plots were made of casts #1 and 10 and the filter was found to have been effective. It smoothed the small steps in pressure and made a significant improvement in salinity. What problems remain may disappear after SHIFT.

6. CELLTM

Tests were run using a variety of parameter choices for CELLTM on casts #1 and 25. For both the best choice was (α,1/β) = (0.03,9). CELLTM was run on all casts using (0.03,9.0).

7. DERIVE

Program DERIVE was run to calculate salinity.

8.  Conversion to IOS Headers

The IOSSHELL routine was used to convert both the CNV and ROS files to IOS Headers. There was an error message for file 2004-13-9048.ios; it lacked latitude and longitude entries, so that information was copied from file 2004-13-0048.ios. The same step was applied to the ROS file.

9. Checking Headers

CLEAN was used to add event numbers and to replace pad values in the pressure channel using linear interpolation based on record number.

The header check was run and no errors were found.

Track plots were produced and no errors found.

The surface routine was run and the average found to be +0.07db. The CTD was often turned on out of water, so this is a reasonable result.

The header summary was run and the output checked against the log book. For the first 6 casts the start time appears to be in PDT. ADD TIME CHANNEL was used to add 7 hours to the headers for those 6 casts and the output was used to replace the CLN files. After cast 7 the times agree approximately with the log. For all casts the log times are generally about 11 minutes later than the header times, and there is a note at the end of the log saying that a check indicated that the internal clock was ahead of GPS time by 11 minutes. It is assumed that the header times are correct, but it is not completely clear how the header time was entered; however, the error is not significant in any case.

There were no errors in positions but some of them differ slightly. The differences are larger than we usually see, but given strong currents, that is likely due to the drift between the time the CTD was turned on and when the entries were made in the logs. This was seen during the 2003 cruise in the same area using the same ship. The information from the headers will be used.

##### 10. Test plots

Plots were made to check for any problems with processing. The fluorescence was frequently off-scale; the range was 0-5 ug/l. The temperature and salinity traces are reasonably close though the upcast is offset relative to the downcast, probably due to wake effects or the pump not operating as well during the upcast. The transmissivity looked good for the early casts with upcasts very similar to downcasts. There were then a few deep casts that showed considerable hysteresis, with downcasts showing higher values than the upcasts. After that most casts showed some hysteresis.  The DO channel looks very strange. For the first 9 casts the shape is the reverse of the temperature profile and the values are unreasonably high at depth (for example, 8.6ml/l for cast #9 at 600db.). From cast #10 onwards the shape of the DO profile is like that of the oxygen voltage, but both look bad with many sections of zero values followed by sudden shifts to reasonable values. If there is any calibration information, it may be possible to recover some sections of useful data, possibly from upcasts, but it seems doubtful that anything can be retrieved. The descent rate was kept fairly high, on average; a few casts had quite a noisy descent rate, but most look quite steady.

Plots were made of scan number versus pressure to ensure there were no initial partial downcasts before the full downcasts as is sometimes done during Arctic cruises due to problems with the equilibration of the pressure sensor. There were no such events during this cruise. The CTD was soaked for from 0.5 to 2 minutes at about 2db before running the casts.

- SHIFT

Casts #7 (1000db), and # 10(100db) were selected to study alignment.

Conductivity

Tests were run using shift values from +0.8 to +1.2 records, and the results examined on a T-S curve after removal of the upcast data; the aim is to minimize instabilities without oversmoothing. The best were with +1.0 records, the same value as used for the 2003 data from this instrument.

SHIFT was run on all casts advancing the conductivity by +1.0 records.

Dissolved Oxygen Sensor

The data looks bad and not amenable to testing the time response, but the same value as used for 2003-18 and 2003-19 was applied in case it is later found there are useful sections.

SHIFT was run on all casts advancing the DO channel by 8 records.

Fluorescence

Study of the fluorescence alignment is limited by the fact that it went off-scale often and because there were so many stops for bottles during upcasts that it is hard to judge the offset between upcast and downcast. However, for the first few casts there are some areas where that offset could be compared with the offset in the temperature trace and no significant difference was seen. So the fluorescence will not be shifted. It usually needs shifting when pumped, but not when unpumped as was the case for this cruise.

- FILTER

A boxcar filter, size 5, was applied to the salinity channel. It removed many unstable features in the T-S plots, but not all.

- DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min and Salinity less than 20psu in the top 20db.

Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 9 points) was deleted from 10db to 10db above the maximum pressure.

The sample interval was taken from the header.

COMMENTS ON WARNINGS: The only warnings were for 3 files that contained only upcast data.

At this point all the repeat casts were examined to ensure the best downcast data was selected for the archive. For casts #30 and 48 the repeats contained only upcast data so they will not be processed further. For cast #24 the computer froze early in the upcast of the first file. The next attempt produced a partial upcast, and finally a full cast was run. That final attempt will be archived. So file 2004-13-8024.del was renamed 2004-13-0024.del.

- DETAILED EDITING

CLIP was used to remove the top 2db since the CTD was stopped at about that level before the full cast was run; the data is very noisy in the top 2db, is not considered reliable from this ship and would take a lot of time to edit. All CLIP files were copied to EDT files

Page plots were produced and examined for spikes and instabilities and used to guide the use of CTDEDIT.  These were produced from the DEL files, i.e. before CLIP was run, so the noise at the surface is worse in the test plots than in the files to be edited.

CTDEDIT was used. Records were removed that appeared to be corrupted by shed wakes. Salinity was cleaned where there were spikes or “overshoots” suspected to be due to misalignment of T and C. This data comes from a region with some very high temperature gradients, resulting in large errors in salinity. For example temperature changed by 3 to 4Cº in a metre or two and more than 8Cº in one 5m segment; alignment of T and C cannot be precise enough to produce reliable salinity under such conditions, and there may also be other temperature-change related affects on the sensors. In handling the 2003 data an attempt was made to fine-tune the alignment to improve the data, but anything that helped in a sharp gradient area, made other parts of the data worse. Many casts were in a region of active mixing and there were some instabilities that looked like they could be real interleaving features; these were left unedited. The salinity gradients were usually fairly low, so even though the editing was at times extremely heavy, the resultant error would generally be within ±0.02psu. Away from the large gradients the salinity is considered ±0.005psu.

For many casts there were unstable regions near the bottom; unless there was a likely instrumental cause for this, such records were left unedited. Records were removed from the bottom if the descent rate had slowed suddenly to near-zero, since shed wakes are likely to account for the instability.

All casts required some editing. Note was made of the editing details in the relevant files.

The edited files (ED1) were bin averaged using 1db bins. (Separate files with 0.5m bins were prepared for the use of Bon van Hardenberg, but they will not be put in the DATA_LIBRARY.)

- COMPARE

A spreadsheet (2004_13_Rosette_Chemistry.xls) was obtained which included salinity bottle values. The file was saved as 2004-13-sal.csv after removal of extraneous information. The headers were changed to standard channel names. There are no sample numbers. Those cannot be added until the bottle files have been edited.

The three problem casts were treated as follows:

- The file named 8024.IOS contains the full 12 bottles. The other versions (0024 and 9024) were deleted and 8024 was renamed as 0024.

- Files 0030.IOS contains bottles 1-2 and 9030.IOS contains 1-12, but the first two are dummy values. The data from 30.IOS was copied and used to replace the first two bottles for 9030. File 0030.IOS was then deleted and file 9030 was renamed as 0030.IOS. The scan numbers are not in order, but the files can be run through routine BIN AVERAGE so this seems ok.

- Files 0048.IOS and 9048.IOS were treated in the same way as the files for cast #30 except that 3 files were copied from the first attempt. File 0048.IOS was then deleted, and 9048 was renamed as 0048.IOS.

CLEAN was run to fix the headers and to add event numbers to the IOS files. (Output:BOT)

The BOT files were then averaged on bottle number and the COMPARE lookup routine was run. The ADDSAMP file was then edited to add sample numbers; these were assigned consecutively. Sample numbers were then added to the BOT files. (Output: SAM)

The SAM files were bin-averaged on bottle number. (Output: SAMAVG)

That ADDSAMP file was printed and the information used to add sample numbers to the spreadsheet 2004-13-sal.csv. The spreadsheet was then converted to individual SAL files. Those files were merged with the SAMAVG files. (Output: MRG)

COMPARE was run using MRG and SAM files as input. There is a lot of scatter in the comparison between bottles and CTD, even at 600db. Most of the sampling is very shallow, but still at depths with low salinity gradient, so they should be useful. Where there is a significant halocline, it is usually around 5 to 20db and quite sharp, so the shallow data should mostly be included in studying the calibration. When differences >0.02psu were excluded there is little pressure dependence and the CTD is found to be low by an average of 0.0019psu. The differences from the 5 deepest casts are as follows:

| Pressure | Sal_CTD - Sal_BOT |
| --- | --- |
| 600.845 | -0.09662 |
| 598.343 | -0.0215 |
| 597.84 | 0.000744 |
| 505.906 | -0.01143 |
| 501.54 | 0.002728 |

A plot of differences versus file pair numbers does not suggest that there is any time-dependence in the calibration.

A plot of differences versus Niskin Bottle number shows an unexpectedly large variation for bottle #1. This bottle was used more than any other bottle, but given it was used for the deepest part of each cast, the result is still surprising. There are two possible explanations:

- The Niskin Bottle may have leaks that lead to variable errors. We would expect that to lead to low salinity in the bottles, which is not the case. In general the CTD is lower than the bottles, not higher.

- It has been observed in the Strait of Georgia that the samples nearest the bottom do not compare well with the CTD. Frank Whitney confirms that turbid waters in the samples could lead to faulty salinity values. To see if this could be the problem, transmissivity was plotted and from casts 24 to the end the transmissivity is quite low near the bottom, especially for casts 30, 33 and 43-48. All of those bottles had salinity higher than the CTD by more than the average value, and five of them were higher by at least 0.03psu. It is possible that this method misses some bottles that were affected by turbidity. The water might have cleared at the CTD sensor level, but the bottle might not have flushed well enough to replace the turbid water.

The turbidity factor probably explains some of the scatter, but problems in analysis and the limitations of the CTD are likely to contribute as well. Whatever the explanation, the differences are not large enough, nor consistent enough to justify recalibration. Moreover, when the T and C sensors were recalibrated there was mention of damage to the cell, but the drift in calibrations were insignificant.

It is likely that the CTD salinity is well within 0.002psu of the bottles, which is the best we can expect from this instrument.

Flags were entered in the MRG files for the bottom bottle from the following casts:  15 (sample #41), 30(175), 43(309), 44(320), 45(332), 46(344), 47(356), 48(368). Niskin #1 for casts 15 and 47 were flagged “d” and for the rest a “c” flag was entered; comments were entered in the header explaining why the flags were assigned.

- Other Intercomparisons

Sensor History – This CTD was used for 2003-18 and 2003-19 but it has been recalibrated since. During 2003 the salinity was found to be high by from 0.001 to 0.002psu. That data was not recalibrated.

Comparison of nearby casts – We have two downcast files for cast #24 with a 40 minute difference in start times. The positions were close. The differences along σt-lines were on the order of 0.2 to 0.3Cº and 0.03psu in the top 20m. Between 20 and 70db the differences were an order of magnitude smaller. This is an area of active mixing so the differences may well be real, rather than a measure of errors.

Historic Ranges – Climatology was available for casts #1 to 24, but it covers a large area so the ranges are broad. It is also unlikely to represent near-shore casts well. Most data fell within the ranges. Exceptions were cast #11 for which the deep (500-600m) temperatures were above the historic maximum and cast 24 which had lower temperature and salinity throughout. Both of those casts were near shore and in areas of great variability, so this is not considered indication of calibration problems, particularly since the temperature excursions are in the opposite direction for the two casts.

- Quality Control

Profile plots were produced with groups of nearby casts plotted together. A few questionable features were noted, and in one case the editing was redone. In the other case the feature, while unstable, has no obvious instrumental cause and may well be real.

- DO sensor data

As mentioned previously the data from the dissolved oxygen sensor look bad. The problems varied through the cruise. The FIL files were put through REVERSE and DELETE to see if the upcast data was better.

- For the first 9 casts the signal looks like the inverse of the temperature channel. While the DO voltage has a non-zero value, the variations are very small, at the noise level. The upcasts look the same.

- For the deep casts #10 and 11 there is a zero value from 0 to about 275db, the value then shoots up to >8ml/l and then gradually goes down to a value of <1 at 600db. The upcast looks the same. Casts #18-20 (deep casts) look similar though the maximum values are lower.

- For casts #12 to 30 (excluding 18-20) there was only noise in the DO channel. The same was true for the upcasts.

- From cast #31 onwards there is a DO signal but it looks like the DO voltage in shape and the upcasts look totally different from the downcasts. During the upcasts the DO variability is extreme reflecting the ascent rate in a strange way. When the ascent rate is high the DO values are high, when it stops the values decrease markedly. This is unlike the temperature and salinity channel behaviour. There are some places where the DO signal seems to reflect quite small variations in the ascent rate.

There was no calibration information available at the time of processing. The DO data will not be archived. The oxygen channels will be saved in files that are kept elsewhere in case anyone finds a way to retrieve useful information.

- Recalibration

No recalibration will be applied. (A pressure offset was applied at the time of conversion.)

- Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Scan_Number, Conductivity:Primary, Descent Rate, Oxygen:Dissolved:SBE, Oxygen:Voltage:SBE and Flag.

- HEADER EDIT and CTD file creation

Header Edit was used to fix formats and channel names and to add the following note to the headers.

Fluorescence:URU:Seapoint – The data are uncalibrated and

unedited and frequently off-scale.

Transmissivity: The data are nominal and unedited.

The pressure sensor is considered ±1db.

The salinity calibration is considered ±0.005psu except in areas

of rapid temperature change where the salinity has been edited

heavily and should be considered ±0.02psu.

The dissolved oxygen data was removed because the data look bad

and there was no calibration sampling.

The Standards Check routine was run and Header Edit adjusted until no errors remained.

The final files were named CTD.

A cross-reference listing was produced.

- Producing final bottle files

The MRG files were put through CLEAN to remove Sea-Bird headers, SORT to order the records on pressure, REMOVE to remove Scan_Number, Conductivity:Primary, Descent Rate, Oxygen:Dissolved:SBE, Oxygen:Voltage:SBE and Flag.

HEADEDIT was used to add a comment and to fix formats and channel names.

The Standards Check routine was run and Header Edit adjusted until no errors remained.

The final files were named CHE.

Particulars from log book

1. All bottles fired, but no samples taken.

8. Touchdown weight lost while out of water.

10. Problems noted in DO channel. DO sensor cleaned, cable changed, sensor moved to different position and con file changed before this cast.

18. Still problems in DO.

20. Lost touchdown weight.

23. Bottles misfired.

24. There were three attempts at this station. One was a downcast and partial upcast. One was a partial upcast and one was a full upcast.

25. Double trip at 15m

30. New file started during upcast ~50m. (4 hex files exist)

48. New file for upcast.

## Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-13

Dates: Start: 8 July 2004 End: 22 July 2004

Location: Bering Sea / Bering Strait / Chukchi Sea

Vessel: Sir Wilfrid Laurier

Party Chief: van Hardenberg B.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0293 | Yes | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0293		Cruise ID#:		2004-13

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4115 | 15 May 04 | Factory | 10/12/04 | Factory |
| Conductivity | 2607 | 14 May 04 | Factory | 20/01/05 | Factory |
| Fluorometer | 2336 | ? | IOS | ? |  |
| Oxygen SBE43 | 0052 | 06Aug01 | Factory | ? |  |
| Transmissometer | 139 | 23Apr01 | IOS | ? |  |
| Pressure | 0464 | 27Dec02 | Factory | 27Dec02 | factory |

** calibration information received from Bon van Hardenberg.