Nutrient Report

Cruise 2004 – 16 Seawater nutrient analysis aboard the Louis S. St. Laurent

Linda White, September 1, 2004

Fresh seawater samples were analyzed on a three channel Technicon Auto Analyzer.

Reagents were prepared on board as needed.  Sigma sodium chloride, 3.2% w/v, was prepared daily to maintain as nitrate free baseline wash as possible.

A Nanopure water system was installed to provide ultra-clean water for standard and reagent preparations and for cleaning labware.

It was fed with ship’s distilled water.

Standards of a Low, Medium and High nutrient concentration were prepared using fresh daily 3.2% Sigma salt solution and analyzed at the beginning, middle of each day’s run when more than 2 stations were analyzed and always at the close for each day.

The standard ranges bracket the expected nutrient levels in the samples.

A medium standard for each nutrient was analyzed between stations.

Proportioning pump tubing was changed after ~ 500 samples.

One cadmium column was used for all samples.

Wako 20 µm/l nitrate calibration standard, and the Japanese AS and AT series of reference samples were analyzed at the end of each day.

Deionized water (dmq) was analyzed before the initial standards and after the last standard set.

Japanese Reference Material Intercalibration

AS and AT reference samples were analyzed daily.  The RM opened on the previous day was analyzed with newly opened reference samples the next day.

AG5, an onboard reference sample was analyzed each day also.

Phosphate Turbidity

Surface samples of salinities < 27 o/oo were analyzed for phosphate turbidity.

No corrections have been made to the data at this time.  Two samples may require correction.

Nitrate + Nitrite

Sigma sodium chloride, 3.2% w/v was used for standard preparations and sample rinse.

The amount of nitrate in Sigma salt is only slightly lower than some surface samples therefore no adjustments were made to surface nitrate sample values.

The autoanalyser was rinsed with 3N NaOH and 1N HCl for ~ 5 – 7 minutes and rinsed with deionized water (dmq) for > 20 minutes after all reagents and salt were disconnected at the end of the day.

Data was logged by analog (chart) and Newget program (computer).

| Quality Control and Assurance |  |  |  |
| --- | --- | --- | --- |
| Nutrient | Nitrate + Nitrite | Silicate | Phosphate |
| Sample Replicates |  |  |  |
| Sp | 0.07 | 0.14 | 0.01 |
| n | 99 | 98 | 98 |
| Medium check standard |  |  |  |
| Calibrated value | 16.1 uM | 40.2uM | 2.01uM |
| Average | 16.2 +/- 0.16 | 40.2 +/- 0.01 | 2.01 +/- 0.01 |
| n | 27 | 29 | 29 |
| AG5 onboard reference | 14.5 +/- 0.25 | 17.2 +/- 0.3 | 1.11 +/- 0.02 |
| N | 16 | 16 | 16 |
| Wako 20uM | 20.1 +/- 0.2 |  |  |
| n | 11 |  |  |
| Quadratic C2 | 0.028 +/-0.006 | 0.048 +/-0.01 | 0.004+/-0.003 |
| n | 45 | 43 | 43 |

Final data file submitted to Janet is: NutrfinalJanet200416Aug29.xls.

Method of nutrient analysis from the Institute of Ocean Sciences, Janet Barwell-Clarke and Frank Whitney.