## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 7 Feb. 2019 | Bottle spreadsheet converted to searchable BOT files. |
| 11-Jan-2006 | Surface loop data was added to the archive. The data was acquired from John Morris at PBS. The original spreadsheet file from John and more detailed processing notes can be found in the “Cruise_Data\Documents” directory. Any questions regarding this data should be directed to John Morris. J.L. |
| 22 Dec. 2004 | Recalibrated CTD salinity based on post-cruise calibration; see note below. |

## PROCESSING NOTES

Cruise: 2004-17

Agency: OSAP

Location: N. W. Pacific

Project: High Seas Salmon

Party Chief: Trudel M.

Platform: W. E. RICKER

Date: June 10, 2004 – June 22, 2004

Processed by: Germaine Gatien

Date of Processing: 19 September 2004 –22 September 2004

Number of original CTD casts: 47

Number of casts processed: 47

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit was a SeaBird model 11 (S/N 0471.) The salinity was analyzed using Portasal Model #8410 (S/N 59724).

# SUMMARY OF QUALITY AND CONCERNS

The CTD Daily Log Book was generally in good order. The CTD is identified as having serial number 0550, but the pressure sensor is the one from CTD #0506. It will be assumed that the CTD is #0506.

There was no bottle sampling below 10m.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and was in good order. No problems are noted in the log.

Bottle salinity and extracted chlorophyll data were obtained. Nutrients were not yet available.

The configuration files were obtained and the calibration constants were checked.

An error in the calibration constants and date for the pressure sensor were corrected, the pressure offset removed, and the date of the transmissometer recalibration corrected. The resulting file was named 2004-17-CTD.con.

The sensor history was found. All temperature and conductivity sensors were last used for cruise 2004-03 and were used for 2004-19 in July 2004.

##### Conversion of Raw Data

The raw data were converted using configuration file 2004-17-CTD.con.

All expected channels were present. The differences between the pairs of channels look fairly good for downcasts but fairly large for upcasts, especially for conductivity. The transmissivity looks much better than for 2004-19, but for a few casts there is a significant difference between upcast and downcast.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### ALIGNCTD

ALIGNCTD was not run even though the deck unit was one of the older model since recent cruises suggest that even the secondary conductivity is being advanced. If this is not so, SHIFT can be used later to do this correction.

##### CELLTM

Tests were run on three casts running CELLTM with choices of (0.01,7), (0.01,9), (0.02,7), (0.02,9), (0.03,7),(0.03,7) and (0.0245,9.5) for (alpha, 1/beta). The best choice for the primary sensors was (0.01,9) or (0.01,7). The secondary data looks slightly improved by using (0.03,7). CELLTM was run using (0.01,9) and (0.03,7) for the primary and secondary, respectively, for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Three casts that sampled to about 200m were plotted to check for agreement between the pairs of T and C sensors. The differences were about 0.001Cº, 0.0008 and 0.0008psu for temperature, conductivity and salinity, respectively; the primary was higher than the secondary in each case. Given that none of these casts were very deep, these differences are reasonably close to the results of 2004-03 and look better than those of 2004-19 and 2004-28 when the same equipment was used.

The transmissivity and fluorescence look much better than for the two latter cruises as well.

The first 4 casts did not have station names in the headers. These were added to the CNV files in the DERIVE folder using information from the log book.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header summary and a header check were produced. No errors were found.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 4.1db, with a minimum of 3.0db; these values are as expected for the Ricker when Hugh Maclean is operating the CTD. The mixed-layer depth calculation shows that none of the surface bottles are likely to be useful for calibration purposes.

T0, T1, S0 and S1 were plotted for all casts. The pairs of sensors are reasonably close for the downcasts but during the upcasts they differ significantly; this has been observed frequently in other data collected from this CTD on W.E. Ricker cruises.

##### SHIFT

Fluorescence

A shift of +24 records has been used for most cruises in the past, including the most recent uses of this particular equipment. A test was run on one cast and +24 records worked well, so it was applied to all casts.

Conductivity

For two recent cruises using this equipment a setting of -1.2 records was found effective in minimizing salinity spiking without oversmoothing. A test was run on cast #92 from this cruise and the same result was found. The secondary salinity has not been archived in recent cruises because it was noisy, and no tests of alignment were done. So tests were run using advancements of from -1 to +1.4 records on cast #92. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were with a shift to the secondary conductivity +0.8 records

All data were put through SHIFT using -1.2 for the primary and +0.8 for the secondary.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Salinity: 5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maximum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The only warnings referred to near-surface upcast data.

##### DETAILED EDITING

The primary sensors were selected for editing.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case.

The descent rate was varied greatly, both in noise and average rate.

The following casts required heavy editing: 16, 19, 21, 30, 33, 36, 39.

The following casts required no editing: 104.

All other casts required light editing.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that do not require editing.

##### Niskin bottle data and comparison to CTD

Bottle data were received from the analysts in spreadsheet format:  2004-17sal.csv and 2004-17chlhugh.xls.

Salinity

After discussion with the analyst some of the flags were removed from the salinity file. Flags were kept for 2 samples which required 4 measurements before stable values were achieved. The information from the spreadsheet was converted into SAL files.

There were no deep bottles and there are fairly high gradients in salinity around 10db, so the bottle comparison is of limited value. To determine which casts are likely to provide the best calibration information all casts were put through REVERSE and clipped to save just the data from 5 to 20db. Plots of pressure versus scan number were used to determine the depth of the bottle stop. Salinity was examined at 5db above that. There were 14 casts that had reasonably low gradients and noise in salinity at the level of the Niskin bottles. These were examined in detail to determine the salinity from the CTD as it passed through that level and to make a quality judgment. A spreadsheet was prepared with the bottle salinity, CTD salinity and the quality judgment (based on the noise and gradient at the level of the Niskin bottle.)

When all 14 casts were used the CTD primary salinity was found to be low by 0.001psu and when only the 7 best casts were used it was high by 0.0003psu. The one bottle with the smallest gradient and noise level in the salinity was low by 0.0006psu. (See 2004-17-sal-comp.xls)

Since the secondary sensors are lower than the primary by about 0.001psu, they would be lower than the bottles by from 0.001psu to 0.003psu.

Chlorophyll versus Fluorescence

All chlorophyll samples were nominally from 10m depth. The actual depths of the Niskin bottles were estimated as described in the salinity comparison section. There were many chlorophyll samples, so it was not practical to go through all the upcast data to judge which would make the best comparison. Moreover, there is some reason to expect the downcast data to be better since the flow-rate is steadier and the descent rate higher, on average, than near the surface of the upcast just after a bottle stop. So the edited downcast CTD files were bin-averaged with 0.5db bins, and thinned to depths from 10 to 15db. Then the value was recorded from the level closest to the bottle stop estimate. There is much room for error in this approach, so this must be considered a very rough comparison.

The fit of CHL versus FL is close to 1 for CHL<4.  For CHL from 4 to 10 the fit is about FLUOR=0.6 * CHL. When CHL>10 the data looks random, and in at least one case the CHL was above the range of the fluorometer. (See 2004-17-chl-fluor-comp.xls.)

To determine how big a difference there is between upcast and downcast fluorescence, the data files were reversed, put through DELETE, bin-averaged and thinned. The results were compared with the downcast data. On average the downcast values are lower than the upcast by about 6% between 10 and 15db. This is a small difference given the noise level.

##### Other comparisons

Previous experience with these sensors – This equipment has been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter. Despite these limitations a trend emerges. The primary sensor changed from being high by about 0.006psu in 2002, to low by about 0.001 in February 2004. The secondary sensors produced salinity high by about 0.0035psu in mid-2003, and gradually shifted downwards. During 2004-03 the salinity was low by about 0.0025psu.

Historic ranges – There were many excursions from the historic ranges in near-surface temperature. The deepest excursions were at casts #74 (from 18-35m) and #92 (0–40m) in Queen Charlotte Sound. Similar observations have been made during a number of cruises in 2004. There is no indication of instrumental error. The salinity was low from the surface to about 35m for two casts in the northernmost line. This has not been noted before, but as this is likely to be an area of great variation and not heavily represented in the climatology this has not been interpreted as a sign of instrument problems.

##### Recalibration

Based on the bottle data and the history of the instrument, the primary salinity is believed to be slightly lower than the bottles, probably within 0.001psu. No recalibration was applied.

##### Special Fluorometer Processing

The EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

Because there are no rosette files, the upcast files were clipped to 50db after SHIFT, REVERSE and DELETE. These were put through REMOVE and HEADEDIT and named UPFCTD. These data are unedited.

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and header entries and to add the following comments:

Transmissivity and Fluorescence – The data are nominal and unedited, except that some records were removed in editing T and S.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-17

Dates: Start: June 10, 2004 End: June 22, 2004

Location: N. W. Pacific

Vessel: W.E. Ricker

Party Chief: Trudel M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2004-17

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory |  |  |
| Conductivity | 2424 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2374 | 20/06/02 | “ |  |  |
| Secondary Cond. | 2399 | 16/04/02 | “ |  |  |
| Transmissometer | 197 | 16/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |

Dec. 22, 2004: All CTD files were recalibrated using file 2004-17-recal2.ccf to add 0.0130psu to the primary salinity based on post-cruise calibration of the conductivity sensor. It was assumed that the drift was linear with time. G. Gatien