## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2004-18

Agency: Stock Assessment

Location: West Coast Vancouver Island

Project: WCVI Trawl

Party Chief: Fargo J.

Platform: W. E. RICKER

Date: May 25, 2004 – June 10, 2004

Processed by: Germaine Gatien

Date of Processing: 14 November 2004 –18 December 2004

Number of original CTD casts: 53

Number of casts processed: 53

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228) on the primary pump. The deck unit was a SeaBird model 11 (S/N 0471.)

# SUMMARY OF QUALITY AND CONCERNS

- The only calibration samples gave results that cast doubt on their quality.

- The file names were non-standard.

- The secondary salinity was extremely noisy and occasionally the data was bad.

- The log book was in good order except that the sensor numbers recorded are incorrect and the bottle sampling was not noted in the log except in a general comment at the beginning.

- The upcast data is very poor with large differences between pairs of sensors and between upcast and downcast values. The secondary salinity was extremely noisy.

- The temperature and conductivity sensors had not been recalibrated since April 2002; a recalibration in November 2004 shows that there was a large drift in conductivity calibration. It is recommended that whatever sensors are used on the Ricker be recalibrated often, since the bottles comparisons from that ship are very limited in number and quality.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

The serial numbers recorded in the log for the two conductivity sensors do not correspond to what is in the configuration file, and no such conductivity sensors are known. There are temperature sensors with those numbers but those are not what are found in the configuration file. One of the sensors was used on another cruise, 2004-09, which overlapped with this cruise. This is puzzling; perhaps something inappropriate was written on the sensors or CTD.

A note in the log indicates that salinity bottles were to be collected at the end of each day at the surface and 10m above the CTD maximum sampling depth. There is note in the log of only one sample collected.

There were 4 deep salinity samples. No surface salinity samples were available at the time of processing.

The configuration files were obtained and the calibration constants were checked. Errors in the calibration constants and two calibration dates were corrected and the pressure offset removed. The resulting file was named 2004-18-CTD.con.

The files had non-standard names, such as CTD2004-18-LB03.dat. These all had to be changed individually. For cast #2 the station is given in the log as LB04. There is no file by that name, but the file for LB05 is at the right place and time. A note in the log mentions that event #2 is actually at LB04 due to miscommunication, so the file named LB05 will be renamed as cast #2 and the station name in the header will be corrected later.

The data for casts #39 to 46 was missing, stations LF08 to LF01; the con file for LF09 was also missing (cast #38.) All the data was found and a new file was found for LF09 with more data in it. There was an electrical glitch during this cast that probably accounts for the existence of two versions.

The sensor history was found.

##### Conversion of Raw Data

The raw data were converted using configuration file 2004-18-CTD.con.

The pressure was checked to see if the offset (-1db) that was in the original configuration files should have been kept. The pumps come on at between 2.7 and 4db which is reasonable for the Ricker. Using an offset of -1db would make this less believable. This will be checked more carefully later.

All expected channels were present. The differences between the pairs of channels look fairly good for the downcast, but as has been noticed in other recent cruises they are large for the upcast, particularly in the conductivity. The transmissivity and fluorescence look reasonable. The descent rate was extremely noisy.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Tests were run on three casts running CELLTM with choices of (0.01,7), (0.01,9), (0.02,7), (0.02,9), (0.03,7),(0.03,7) and (0.0245,9.5) for (alpha, 1/beta). The best choice for the primary sensors was (0.01,7) and for the secondary (0.0245, 9.5). CELLTM was run on all casts using those values.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Three casts that sampled to 1250m were plotted to check for agreement between the pairs of T and C sensors during the downcast. The differences were very noisy.

| Cast # | T1-T0 | C1-C0 | S1-S0 | Descent rate |
| --- | --- | --- | --- | --- |
| 11 | 0 to -0.0005 | -0.00075 | -0.009 | Low, noisy |
| 38 | -0.0005 | -0.0007 | -0.007 | Low, noisy |
| 47 | avg-0.0002 X Noisy | -0.0006 | -0.007 | Med, X. noisy |

These differences are reasonably consistent. The upcast differences are not significantly different.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header summary and a header check were produced. The only problem noted was an error in the station name for cast #2. (There was an entry in the log book about this error.)

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 3.0db, with a minimum of 0.8db. While these values seem slightly low for the Ricker, the method used to deploy the CTD might differ somewhat from that used by Hugh Maclean.

T0, T1, S0 and S1 were plotted for all casts. The secondary salinity looks extremely noisy, especially near the surface in high temperature gradient areas. Cast #39 has a large excursion of secondary salinity to large values from about 200db to 400db of the downcast at which point it abruptly returns to values close to the primary salinity.

##### SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles for a few casts were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. A shift of +24 records has been used for most cruises in the past and seems reasonable for this one, so it was applied to all casts.

Conductivity

When this CTD was last used a value of -1.2 records was used for the primary and -0.6 for the secondary. However, ALIGN had been run the on the secondary, which was not done this time. Tests were run using advancements of -1, -1.2 and -1.5 records on the primary conductivity channel for cast #48. For the secondary, tests were run using +.7 to -1.5 records. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were with a shift of -1.2 records for both channels.

All data were put through SHIFT using -1.2 for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Salinity: 5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maximum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The only warning referred to cast #39 during the downcast when the pressure jumps from 350.6 to 354.6db, and there are unbelievable spikes in many variables. There are 7 records missing but the pressure jumps by more than would be expected in the missing time. It has already been noted that the secondary salinity was very bad for this cast from 200 to 400db.

##### DETAILED EDITING

The secondary sensors did not perform well. The salinity is very bad especially in areas of high temperature gradient, and for cast #39 the data is unusable. The primary sensors have some noise but the quality of the data is reasonably good.

Page plots were produced using T0, S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case.

The descent rate was very noisy for most of the casts, but the average rate was generally kept high minimizing the corruption by shed wakes.

All casts required some editing.

Note was made of the editing details in the relevant files.

##### Comparisons

Bottle comparisons –.

There were 4 bottles fired during deep casts. There is unfortunately some doubt about the depth and manner of firing. A general note in the log book says they were to be fired 10m above the bottom, but the CTD did not stop except at the bottom. So if the bottle was fired anywhere other than the bottom, this was done “on the fly”. So it is not clear what salinity to compare the bottles with. The salinity analysis sheet gives a depth approximately equal to the maximum pressure for the casts. This suggests that they were actually fired when the CTD was stopped at the bottom. Typically that would mean the Niskin bottle was about 4 or 5m above the bottom. The following comparison assumes that the salinity in the Niskin bottle is from the bottom of the cast.

| Salinity:Bottle | Sample_Number | Number of points | Flag:Salinity:Bottle | Sal0 CTD | Sal1 CTD | Sal0-Sal_bot | Sal1-sal_bot |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 34.4139 | 47 | 2 |  | 34.4845 | 34.4782 | 0.0706 | 0.0643 |
| 34.4501 | 38 | 2 |  | 34.4782 | 34.4704 | 0.0281 | 0.0203 |
| 34.4214 | 30 | 2 | c | 34.4570 | 34.4485 | 0.0356 | 0.0271 |
| 34.4973 | 11 | 2 |  | 34.4945 | 34.4855 | -0.0028 | -0.0118 |

The comparisons from 3 of the bottles show very large errors of the opposite sign to what is expected based on post-cruise calibrations and previous experience. Cast #11 has the expected sign. If the bottles were fired “on the fly” when the CTD was about 10m above the bottom then we would expect errors due to poor flushing of the Niskin bottle. If fired during the downcast we would expect the CTD to look higher than the bottles, and if it was during the upcast we would expect it to look lower. So the first 3 bottles might make sense if they were fired on the fly 10m above the bottom during the downcast.

If they were fired while the CTD was stopped at the bottom, there would be some error due to the Niskin being a little above the CTD, but the local gradients don’t begin to explain these differences. Perhaps some other factor affected the flushing of the Niskin bottles. If so, this must have happened after cast #11. Another possibility is that the CTD was not working well. There is some evidence of flow-rate problems in the secondary sensors. This could cause differences that are not due to sensor calibration. However, it is not likely that both pumps would be affected in the same way, as appears to be the case for 3 of these bottles.

Whatever the cause this data is inconsistent with other information, so is not useful for calibration.

Previous experience with these sensors – These conductivity sensors have been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter. Also, they were used with different temperature sensors in 2002. The method for bottle sampling from the RICKER is inexact and flow-rate problems appear to plague at least the secondary system. So, little weight can be put on the bottle comparisons.

The primary salinity appeared to be high (but gradually drifting lower) until the temperature sensor was changed in 2003. After that the salinity always appeared to be low by from 0.001 to 0.007psu with no trend. There were few deep bottles and a lot of scatter in the results. During 2004-17 in June 2004 the salinity was estimated to be low by up to 0.002psu, but there was no deep sampling and a lot of scatter in the shallow sampling.

Similarly the secondary sensors produced high salinity until the temperature sensor was changed in 2003. After that it appears to have been slightly high at first, then trending to lower values. In February 2004 the salinity was found to be low by about 0.0025psu and in June 2004 the secondary appeared to be low by from 0.001 to 0.003psu.

Post-cruise calibration - There are post-cruise calibrations available for all instruments.

- The primary temperature is estimated to have drifted upwards by about 0.0004Cº since its last calibration and the secondary temperature by 0.0002C. Neither is considered a significant error.

- The conductivity drifts are significant. If the drift is assumed to be steady then the primary salinity would be low by 0.0125psu and the secondary low by 0.015psu. This will require re-calibration.

Historic ranges –The only data that fell outside the local climatology was in the top 20m of most of the stations with bottom depths between 100 and 200m where the temperatures were slightly above the maxima. This has been seen from other cruises in 2004 and is not considered an indication of instrumental error.

##### Recalibration

If we assume that the drift was steady, then the post-cruise calibration implies that the primary temperature is high by 0.0003Cº and the salinity low by about 0.0125su. This is the most reliable information available. So, using file 2004-18-recal1.ccf, these values were applied to all files.

##### Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and header entries and to add the following comments:

Transmissivity and Fluorescence – The data are nominal and unedited, except that some records were removed in editing T and S.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

39 – All channels show some evidence of electrical problems; secondary conductivity data was very bad from about 200 to 400db.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-18

Dates: Start: May 25, 2004 End: June 10, 2004

Location: West Coast Vancouver Island

Vessel: W.E. Ricker

Party Chief: Fargo J.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2004-18

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory | 18/Nov/04 |  |
| Conductivity | 2424 | 16/04/02 | “ | 19/Nov/04 |  |
| Secondary Temp. | 2374 | 20/06/02 | “ | 18/Nov/04 |  |
| Secondary Cond. | 2399 | 16/04/02 | “ | 19/Nov/04 |  |
| Transmissometer | 197 | 16/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |