## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 22 Dec. 2004 | Recalibrated CTD salinity based on post-cruise calibration; see note below. |

## PROCESSING NOTES

Cruise: 2004-19

Agency: OSAP

Location: N. W. Pacific

Project: Copra/Sardines

Party Chief: Andrews, B.

Platform: W. E. RICKER

Date: July 20, 2004 – August 2, 2004

Processed by: Germaine Gatien

Date of Processing: 10 September 2004 –23 September 2004

Number of original CTD casts: 21

Number of casts processed: 21

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit was a SeaBird model 11 (S/N not recorded.)

# SUMMARY OF QUALITY AND CONCERNS

- The CTD Daily Log Book contained no information about the equipment or scientific crew.

- There were two deep bottle salinity samples, but they could not be located and were probably never analysed.

- The transmissivity data has very different upcast and downcast sections for many casts. The downcasts look reasonable, but should be used with caution.

- Fluorescence data looks poor near the surface in the upcasts. The downcasts seem ok.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Two deep salinity bottles were collected during the cruise, but appear not to have been analysed.

The configuration files were obtained and the calibration constants were checked. Errors in the calibration constants and two calibration dates were corrected and the pressure offset removed. The resulting file was named 2004-19CTD.con.

The sensor history was found.

##### Conversion of Raw Data

The raw data were converted using configuration file 2004-19CTD.con.

The pressure was checked to see if the offset (-1db) that was in the original configuration files should have been kept. The pumps come on at about 3 to 3.5db which is slightly higher than the usual data from the Ricker. Using a -1db offset would make this less believable. The surface data in the downcast might suggest that  the pressures are low, but from the upcast this does not appear to be the case. This will be checked more carefully later.

All expected channels were present. The differences between the pairs of channels look fairly good for temperature, but fairly large for conductivity especially during the upcast. The transmissivity looks very different between upcast and downcast.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5	Points per block = 50

##### ALIGNCTD

ALIGNCTD was used to advance the secondary conductivity by 0.073s since it is likely that this deck unit was one of the older models. Fine-tuning of the alignment will be done later using SHIFT.

##### CELLTM

Tests were run on three casts running CELLTM with choices of (0.01,7), (0.01,9), (0.02,7), (0.02,9), (0.03,7),(0.03,7) and (0.0245,9.5) for (alpha, 1/beta). The best choice was for the primary sensors was (0.02,9) for one cast, (0.01,9) for another and the third looked best with no CELLTM applied. As the third was in the calmest conditions and reasonably deep, it is probably the best indicator. The same result was found the last time this equipment was used during 2004-03.

The secondary data looks very poor, with upcast and downcast so different there is no evidence of any change with any setting of CELLTM.

CELLTM was skipped.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Three casts that sampled to 500m were plotted to check for agreement between the pairs of T and C sensors. All the differences were noisier than ever seen before and vary greatly from one cast to another. Looking at the data on a T-S surface shows that the secondary data is frequently unstable, especially at depth. So the comparison is not helpful.

The transmissivity and fluorescence are suspicious. The upcasts are much different from the downcasts.

The column in the log book which indicates that the transmissometer was cleaned has no entries for this cruise.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header summary and a header check were produced. The header of cast #29 had scrambled times and positions; that was fixed in the CNV file and it was reconverted and CLEAN rerun. The station names were wrong for 2 cast; these were fixed in the headers of the CLN files.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 2.9db, with a minimum of 0.8db; these values seem low for the Ricker. An examination of the data file with the minimum of 0.8db shows that the CTD really was close to the surface with very low conductivity values. The unusual values are clearly due to the method used to deploy the CTD which differs from that used by Hugh Maclean when he samples from the Ricker. The pressure calibration is not a problem.

T0, T1, S0 and S1 were plotted for all casts. The secondary salinity looks poor, especially near the surface of the upcast when the values fall well below the historic range. The temperature sensors look reasonably close. The fluorescence looks reasonable on the downcast, but for some casts the upcast values in the top 30m look much too low and are sometimes extremely spiky.

##### SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles for a few casts were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. For this cruise the upcasts are plagued with problems, so the estimate is extremely rough. A shift of +24 records has been used for most cruises in the past and seems reasonable for this one, so it was applied to all casts.

Conductivity

Tests were run using advancements of +0.5, 0, -0.5, -1, -1.2 and -1.4 records on cast #37. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were with a shift to the primary conductivity -1.2 records. This is the same result as found during 2004-03 when the same equipment was used.

All data were put through SHIFT using -1.2 for the primary. No shift was applied to the secondary data since it is unlikely to be archived.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Salinity: 5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maxiumum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The only warnings referred to bottom or upcast data.

##### DETAILED EDITING

The secondary sensors did not perform well. The salinity is very bad for some of the downcasts (#7 and #13) and most upcasts. The primary sensors have some noise but are reasonable.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case.

The descent rate was very noisy for most of the casts, but for most the average rate was kept high minimizing the corruption by shed wakes.

All casts required some editing.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that do not require editing.

##### Other comparisons

Previous experience with these sensors – This equipment has been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter. Despite these limitations a trend emerges. The primary sensor changed from being high by about 0.006psu in 2002, to low by about 0.001 in February 2004. The secondary sensors produced salinity high by about 0.0035psu in mid-2003, and gradually shifted downwards. During 2004-03 the salinity was low by about 0.0025psu. For 2004-17 which took place shortly before this cruise there was only surface sampling, but it suggests that the CTD primary salinity was probably within 0.001psu of the bottles and the secondary low by from 0.001 to 0.003psu.

Historic ranges – There were some excursions from the historic ranges in near-surface temperature only. The temperatures were high in the top 15m at stations A2, A3 and B6. Similar observations have been made by a number of cruises in 2004. There is no indication of instrumental error.

##### Recalibration

The primary salinity is believed to be slightly lower than the bottles, probably within 0.001psu. No recalibration was applied.

##### Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and header entries and to add the following comments:

Transmissivity and Fluorescence – The data are nominal and unedited, except that some records were removed in editing T and S.

The  upcast and downcast transmissivity data were very different for many casts.

The downcast data looked reasonable, but should be used with caution.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-19

Dates: Start: July 20, 2004 End: August 2, 2004

Location: West Coast Vancouver Island

Vessel: W.E. Ricker

Party Chief: Andrews B.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2004-19

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory |  |  |
| Conductivity | 2424 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2374 | 20/06/02 | “ |  |  |
| Secondary Cond. | 2399 | 16/04/02 | “ |  |  |
| Transmissometer | 197 | 16/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |

Dec. 22, 2004: All CTD files were recalibrated using file 2004-19-recal2.ccf to add 0.0135psu to the primary salinity based on post-cruise calibration of the conductivity sensor. It was assumed that the drift was linear with time. G. Gatien