## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 23-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 11-Apr-2010 | Added Lisa Miller’s Dissolved Inorganic Carbon and Alkalinity data to the rosette files. These two parameters were already in the rosette files so they were removed prior to adding the new set. J.L. |
| 31-May-2005 | Added Chlorophyll data to rosette files. J.L. |
| 18-Apr-2005 | Added Carbon data to rosette files. J.L. |

## PROCESSING NOTES

Cruise: 2004-20

Agency: OSAP

Location: Gulf of Alaska

Project: Line P

Party Chief: Whitney F.

Platform: John P. Tully

Date: August 20, 2004 – September 2, 2004

Processed by: Germaine Gatien

Date of Processing: September 23, 2004 – 3 December 2004

Number of original CTD casts: 61 (There are 63 files – two casts were split into 2 files each)

Number of CTD casts processed: 61

Number of rosette casts: 19 full rosette casts and 41 with 5m Niskin

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with Wetlab transmissometer (#723), Altimeter OA-916D (#1024), a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0424) and the logging computer was #FS02. The salinometer was a Portasal model 8410 (#58879). The deck unit was #508. A mid-ship winch was used. A SeaBird 25 CTD mounted with Wet Star fluorometer (S/N 713P) was used for loop sampling.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in good order.

The choice of which sensor pair to archive was not obvious. The secondary salinity differed from the bottles by more than 0.01psu. The primary pair was selected (except for 2 casts) but the salinity was extremely noisy in high-gradient areas. The noise is mostly bi-polar suggesting flow-rate irregularity.

There are a number of odd features in this data. There are large differences between upcast and downcast data for a few casts. The difference between temperature sensors is more pressure-dependent than expected and varies with descent rate. The difference between conductivity sensors is larger than usual but not pressure-dependent. The salinity differences are very large and the bottle calibration is inconsistent with the results of the post-cruise calibration. Together with the dependence on descent rate this suggests problems with the pump or plumbing. There is also some time-dependence with a suggestion that something changed after cast #7.

There have been secondary temperature problems in other recent cruises with this CTD, but with a different sensor.

Tests run on the Portasal salinity measurements suggest a noise level of ±0.001psu in careful shipboard use. These tests also show that the bottles flushed consistently.

The fluorometer used on the SBE25 (loop) did not perform well, probably due to inadequate cleaning.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Together with the notes from Marie Robert this gave an excellent summary of problems encountered during the cruise. There were a number of equipment problems.

Bottle salinity and dissolved oxygen data were obtained; flags had been added, but no comments.

The cruise summary sheet was completed.

The histories of the conductivity and pressure sensors were obtained.

The calibration constants were checked for all instruments. The date of the calibration for the pressure sensor was wrong. There was an offset of -0.6db in the pressure configuration. During 2004-04, 2004-05, 2004-07 and 2004-10 pressure offsets of +0.2db were found appropriate; in the case of 2004-04 a careful surface pressure test was done. For 2004-22 an offset of +0.4db was need to achieve reasonable results. A test was done using the offset in the original configuration files and for cast #2 upcast data had a surface pressure of about -0.8db with what are clearly in-water conductivity values. An offset of +0.4db will be applied and the results checked later.

A new configuration file 2004-20-ctd.con was prepared with the calibration data and offset changed.

##### Conversion of Raw Data

All data were converted. A few casts were examined and all expected channels are present. The up and downcasts were mostly similar except for cast #40 which had bad downcast primary temperature. For all casts the pairs of temperature sensors are farther apart than usual and the conductivity difference seems quite large, especially late in the cruise. The salinity contains many large spikes.

Rosette files were converted using a start time of -2s and duration of 5s.

Corrections were made to the headers of casts 8, 21, 54 in both ROS and CNV files.

The rosette files were then converted to IOS SHELL files. These were put through CLEAN to add event numbers. Plots were examined on-screen for outliers. Casts #28 and 38 were very noisy around 30db, but examination of the data shows that the CTD had more vertical motion than usual and the gradients are high around 30db, so editing is not appropriate.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### ALIGNCTD

The deck unit was one of the newer ones that advances secondary conductivity channels and there was no dissolved oxygen sensor, so ALIGNCTD was not run.

##### CELLTM

Tests were run on 3 casts with (alpha, 1/beta) set to (0.01,9),(0.02,7), (0.03,7), (0.02,9), (0.0245,9.5) and (0.03,9).

Both (0.02, 7) and (0.02, 9) improved the primary data notably in the top 150m. Below that the effect was poor for 2 casts. This is an unexpected result. For 2004-22 there was a DO sensor mounted, but otherwise the equipment was the same; at that time (0.02, 7) worked well for the primary. More casts were tested and most looked better after using (0.02, 7).

The setting of (0.02,7) improved the secondary conductivity best for 2 casts and was close on the third.

While there is not the obvious difference from cast to cast as seen in the primary, such fine differences are hard to see in the secondary due to the noise level.

A study was made of two casts that seemed worse after CELLTM, #11 and #29. Average descent rate or noisy descent rate does not seem to be a factor. But the casts that seemed worse have very large differences between upcast and downcast data. The upcast data also looks quite noisy, so it may be that there is a problem with the upcast, say with flow to the sensors. The differences between upcast and downcast salinity are slightly larger for the primary sensors. Or it may be that there was a real change in conditions. In either case we cannot expect CELLTM to make the upcast and downcast look similar on a T-S surface, if they are really quite different. The failure of the test on these casts does not imply that the setting is wrong.

CELLTM was run using (0.02, 7) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 10 | 2000 | -0.0004 | +0.0011 | +0.013 | High, moderate noise |
| 20 | 2000 | -0.0002 | +0.0012 | +0.015 | High |
| 30 | 2000 | -0.0002 | +0.0011 | +0.010 | High, very noisy |
| 41 | 2000 | -0.0003 | +0.0011 | +0.013 | High, moderate noise |
| 50 | 2000 | -0.0004 | +0.0011 | +0.013 | High |
| 60 | 2000 | -0.0005 | +0.0011 | +0.014 | High |

All the differences were very noisy.

As has been noted during other recent uses of this equipment, the temperature difference was more pressure-dependent than usual. And the differences in the top 500db are very large. For cast #50 the temperature differences  were -0.01 at 100db, about -0.0015 at 100db and -0.0004 at 2000db. The upcast differences look noisier with some reversals in sign. Since the differences go down with pressure this is presumably a function of gradient and it is probable that one sensor has poorer resolution than usual. But which one? During 2004-21 when the same CTD was used there is a hint that the secondary temperature did not resolve a sharp feature as well as the primary did. However, it was a different secondary sensor.

To investigate further cast #18 was examined in detail. It was found that the secondary temperature was higher than the primary during stops by about 0.0006Cº, lower during downcast motion by about 0.0002Cº when the descent rate was about 1m/s and by about 0.0007Cº when the descent rate was 2m/s. During upward motion there is a lot of variability, but the secondary is generally close to, or less than the primary temperature. If we assume that one of the temperature sensors is behaving properly then the other varies by 0.001Cº between high descent-rate downward motion and non-motion. This would affect salinity by something like 0.001psu. There is no proof that one sensor is more reliable, nor whether the data is better when moving or stopped. However, the behaviour as the CTD slows to a stop at the bottom of cast #18 suggests that the secondary sensor may be the problem. Both temperatures rise slightly, then the primary goes down in an apparent shed wake effect. The secondary stops rising through that event, but doesn’t go down and after the shed wake it continues to increase. So the secondary behaves in a somewhat unexpected way.

Conductivity differences showed little variation with pressure, but were higher than usual.

The salinity differences are unusually large. There is some pressure-dependence, but it is not as striking as in the temperature differences. The secondary salinity is full of fine-scale noise with average excursions of ±0.0015psu at 2000db, while the noise level of the primary is about ±0.0005psu. The secondary conductivity had a noise level of ±0.00015 S/m whereas the primary has noise on the order of ±0.00003.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run. One station name was corrected.

The cruise track was plotted and no problems were noted.

The average surface pressure is 3.5db, which is a little higher than usual, but there is a wide range of values from 0.8 to 7.6db. The cast with the lowest surface reading was examined and values in the top 1db look like near-surface, but in-water values. The cast with a minimum of 7.6db for the downcast, came to about 2db on the upcast and all values look reasonable for their associated pressures. The pressure appears to be close to correct and will not be adjusted further.

A few casts were examined on screen. Casts #31 and 40 have severe problems in the primary sensors during the downcast, as noted in the log. From cast #14 to 60 there is no altimeter mounted. A few casts were checked to ensure the header got the right reading from the altimeter. There was a lot of noise, but the algorithm worked well. The transmissivity and fluorometer traces look reasonable although there are significant differences between downcast and upcast transmissivity in the top 100db for some casts.

No other problems were noted.

##### BOTTLE FILE PREPARATION

The four salinity analysis spreadsheets were edited to remove flags that had been added to many samples; the analyst had applied more rigorous tolerance requirements than usual and agreed to the removal of the flags. “C” flags were left in a few samples for which the drift was continuing after 4 tries, but for which there were 2 values within the usual tolerance. These will be examined after running COMPARE to see if they look ok. The loop samples were removed. The files were converted to individual SAL files. Comments for 2 flagged values were added to the headers of two SAL files.

Dissolved oxygen files (*.add) created by the analyst had a flag channel. The flags are entered in the wrong column and there were no comments. The flags were corrected and comments added based on notes on the rosette logs. The ADD file for cast #7 was originally named #6. This was corrected.

Final extracted chlorophyll and nutrient data were not yet available. However, there was some preliminary CHL data. This was examined and the CTD fluorescence appears to be about 3 times the CHL values. (See 2004-20 chl-fl-comp.xls.)

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Cast #9 had 2 bottles fired but there is a sample number for only Niskin #1. Similarly, for cast #74, all bottles were fired, but only one sampled. The bottles that were not sampled were removed from the SAMAVG file to enable MERGE to work. The SAMAVG file put through CLEAN to fix the headers and the output file was renamed SAMAVG.

The SAL and ADD files were merged with SAMAVG in two steps. (Output: MRG1, MRG2)

DO was plotted against CTD Salinity; the only outlier came from the thermocline at a depth where fluorescence was high – low DO is reasonable there.

NOTE: After the CHE files had been produced the DO data was found to have errors due to a software problem. New files were provided by Sheila Toews which contained the original values and corrected ones. The original values channel was removed (*.ADDREM) and the new one renamed with the standard channel name (*.ADDNEW). MERGE was rerun with the new files and they were used to produce new CHE files.

##### 12. COMPARE

Salinity

COMPARE was run. There is a lot of noise in the comparison for both sensors with some pressure-dependence in both. When a few outliers and all data from cast #7 are excluded, a good fit is found for the differences between the bottles and primary salinity versus pressure (Difference = -8E-07 * Pressure + 0.0008). The primary salinity is higher than the bottles in the top 1200db and lower than the bottles by about 0.0025psu at 4500db. The secondary sensors are high by from 0.0117 to 0.0123; if cast #7 is excluded a good fit is found versus pressure (Difference = -6E-07 * Pressure + 0.0137). For cast #7 the primary salinity is low by about 0.0025psu and the secondary is high by about 0.005; the cast is not very deep and there is only weak pressure dependence. The bottles from before cast #7 were all shallow.  (See 2004-20-sal-comp1.xls.)

Three outliers not previously flagged were identified as needing flags; these were added to the MRG2 files along with comments explaining them. (Samples 34, 84, 148)

The salinity samples flagged earlier were checked and most look suspect, but not obviously bad, so the “c” flags will be left and the comments updated (samples 39, 74, 330, 345). One was a severe outlier so the flag was changed to “d” (sample 40). A few were found to be good in COMPARE and the flag was removed (samples 61, 63).

During cast #43 23 bottles were fired at 2000db. The variations were studied to see how reliable the bottle comparisons are likely to be. There are intrinsic errors in running the analysis, expected to be a little higher at sea due to difficulties in maintaining constant temperature in the analysis room. Also there may be inconsistencies in the flushing of the bottles.

| Analysis of CTD data during the 23 bottle firings at 2000db – cast #43 |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Pressure | T0 | T1 | C0 | C1 | S0 | S1 |
| Avg. | 2002.7 | 1.9535 | 1.9533 | 3.1263 | 3.1273 | 34.5769 | 34.5900 |
| Min. | 2001.0 | 1.9525 | 1.9522 | 3.1262 | 3.1271 | 34.5756 | 34.5869 |
| Max. | 2004.3 | 1.9547 | 1.9548 | 3.1264 | 3.1276 | 34.5781 | 34.5926 |
| Std dev. | 0.7 | 0.0004 | 0.0004 | 0.0000 | 0.0001 | 0.0003 | 0.0010 |
| range | 3.3150 | 0.0022 | 0.0026 | 0.0002 | 0.0005 | 0.0025 | 0.0057 |

The CTD results are averaged over a 5s-window around the firing time for each bottle. The pressure varied through a range of 3.3db; from the downcast gradient we would expect the salinity range to be about 0.0003psu. The two CTD salinity channels varied by 0.0025psu and 0.0057psu during the bottle stop but when averaged on bottle number the range is 0.0004psu and 0.0005psu. The secondary salinity varied a lot more than the primary which is expected because we know there is a lot of noise in the secondary conductivity and there is some question about how well the secondary temperature responded to change. However, the average looks reasonably good presumably because the noise is mostly bi-polar. This result also suggests that the bottles flushed consistently, no one bottle being way out of line. (See Deep bottle study.xls)

The Portasal analysis showed a range of 0.0045psu.

| Avg. | 34.57819 |
| --- | --- |
| Min. | 34.576 |
| Max. | 34.5805 |
| StdDev. | 0.001209 |
| range | 0.0045 |

On average the primary salinity is lower than the bottle salinity by 0.0013psu. Adding that to the primary CTD salinity we get a range of values from 34.5769 to 34.5794psu. The Portasal minimum and maximum values are approximately 0.001psu lower and higher, respectively.  That would suggest a noise level of ±0.001psu in the Portasal readings, at least under the conditions on the Tully during this cruise. It helps set a limit on our expectations from the COMPARE program. The noise in COMPARE is typically much higher than that.

##### Other Comparisons

Previous experience with these sensors – The primary sensor has been used for 5 other cruises since recalibration; the salinity was compared with bottles and was high by 0.0004psu, low by 0.0003, 0.0032, 0.0015 and 0.0025psu in March, April, May, June and July. The secondary sensor was used in July when it was higher than bottles by 0.005psu.

Historic ranges –The temperature was slightly above the maximum of the historic range for cast #10 around 20db and 40-50db and was slightly below the minimum for temperature around 100-200db of cast #24. The temperature was also a little high near the bottom of casts #62 and 63. The latter excursions are likely due to limits in the climatology as these sites may not be represented in the climatology. The higher temperatures near the surface have been observed for other cruises in 2004. There is no evidence of CTD malfunction.

Post cruise recalibration – Post cruise recalibration information indicates that the primary temperature sensor was reading low by about 0.001Cº. This would lead to the salinity reading high by about 0.001psu; these errors vary through the water column becoming larger at depth. If we recalibrate the temperature the resulting salinity will be low by something like 0.002psu which fits with results from other cruises. This would suggest that the bulk of the drift in the temperature calibration developed between July and August.

The secondary temperature sensor had little drift since the previous calibration.

The primary conductivity sensor accounts for a drift in salinity of -0.0055psu if the drift is linear. For the secondary sensors the error in conductivity accounts for salinity being high by about 0.0035psu.

##### 14. SHIFT

Conductivity

The primary sensors have been used often in recent months and a shift of -0.2 records has worked well in every case. So tests were run on casts #4 and #41 using that setting and it was found to work well.

The secondary conductivity sensor has been used only during 2004-22 when it was not shifted since the data was not selected for the archive. For this cruise it will be needed for at least one cast, so tests were run to determine what shift reduces instabilities best without oversmoothing; results were examined in T-S plots after running DELETE. The best results came from an advancement of -0.5 records.

All casts were put through SHIFT using -0.2 records for the primary conductivity and -0.5 records for the secondary. (Output *.SHFC1 and SHFC2).

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was 1 to 1.5s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings about the two casts that contained only upcast data and one cast that had an odd jump in pressure during the upcast.

Cast #10 had an initial section during which the pumps were off. The CTD was returned to the surface and the cast rerun with pumps on. A text editor was used to remove the first 10070 records so that DELETE would select the records with the pumps on. CLEAN was run to reset the header limits.

##### DETAILED EDITING

It is not obvious which sensor pair to pick for the archive. The primary sensors show some pressure dependence in COMPARE which is worrisome, yet they are closer to the bottles. The differences found for this cruise fit the history of the sensors. The secondary sensors are very different from the bottles, are a little pressure-dependent and much higher than was observed in July for 2004-22. There is also some suggestion that there may be a problem in the response of the secondary temperature. So the primary temperature and salinity were selected for further processing, except for casts #31 and 40, for which the secondary will be used.

Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing, and some required heavy editing. There was a lot of noise in the salinity in the top 100db. The average CTD descent rate was kept very high for the deep casts which reduced the incidence of shed wakes.

Casts requiring heavy editing were: 13, 14, 21, 23, 26-38, 42.

Note was made of the editing details in the headers of the relevant files.

##### Calibration study

There are a number of sources of information on the calibration of the sensors and the results are contradictory:

- Post-cruise calibrations (assuming linear drift for 8 months)

primary temperature is low by ~0.001Cº

secondary temperature – no significant change

→T1-T0 ~ +0.001

primary salinity is low by ~0.0056psu

(Note there was no pressure-dependence in the differences between old and new calibrations.)

secondary salinity is high by ~0.006psu

→ S1-S0 ~ +0.012

- Comparison of sensors during downcast

T1-T0 ~ -0.0002 to -0.0007 with some dependence on descent rate

C1-C0 ~ +0.0011 with little variation with time, pressure or descent rate

S1-S0 ~ +0.010 to 0.015 around 2000db, higher above that.

(Exception: Cast 6 –this was a downcast only - +0.012 at 500db, +0.009psu at 1250db.)

- Comparison of sensors during upcast

T1 ~ T0 (lot of variability)

S1-S0 similar to downcast at 2000db and a little less above that (Exception: Cast 7 – this was an upcast only – +0.007psu at 500db, +0.008psu at 1250db)

- Comparison of sensors during bottle stops

T1-T0 ~ +0.0006

S1-S0 ~ 0.013 (except cast #7 which is +0.008psu)

- Study of sensor behaviour as CTD stops

T1 does not seem to respond as well to changes as T0.

- Comparison to bottles

Primary is pressure-dependent, high by +0.001 at 500db, low by 0.0002db at 1000db, low by 0.0026psu at 4500db, except for cast #7 which is -0.002 at 500db and -0.0026 at 1250db.

Secondary is high by about 0.013 if cast #7 is excluded, and slightly pressure-dependent.

The differences between the salinity from the two pairs of sensors are reasonably consistent except for cast #7. But the differences between each of them and the bottles are not consistent with the results of the post-cruise calibration. It looks like each of the salinity channels are higher (by about 0.004psu to 0.007psu) than they should be. What could cause this? For cast #7 the bottle comparison is closer to the post-cruise calibration. The offset for cast #7 looks like the value for other casts at 4000db. That was the first deep cast, so we don’t know if the problem started after that cast, or just disappeared for the one cast – which is possible if it was a plumbing problem.

The temperature channels hold a clue that something odd is going on. The primary temperature is low according to the post-cruise calibration and the differences during the bottle stops are in agreement with this. But during downcast motion the primary is higher than the secondary, the difference being higher when the descent rate is higher. Which is right? Or are both wrong? In order to explain both salinity channels reading too high, we need temperature too low in both channels while in motion. Could there be an electrical or plumbing explanation for this? And whatever happened, did it stop happening during cast #7 or develop after that cast?

The conductivity differences show no time or pressure-dependence so those sensors are probably fine.

The recalibration of the temperature will be done using the results of the post-cruise calibration. There is some doubt if this is appropriate since when the CTD is in motion the relative differences in the sensors changes. However, there is some suggestion that the changes are caused by poor response of the secondary sensor data.

The recalibration of the salinity will be based on the bottle data. Why this is so different from the results of the post-cruise calibration is unknown but pump or plumbing problems are suspected. It is clear that this should be used for the rosette data, but not so obvious for the downcast data. A note will be put in the headers to indicate that the primary and secondary salinity channels may be low by as much as 0.004psu and 0.007psu, respectively. A separate recalibration will be done for the salinity for casts #1 to #7; while there is some suggestion of pressure dependence a simple offset of +0.0025psu for the primary and -0.0050psu for the secondary does a reasonable job. There is no way to determine how casts #1 through 6 should be recalibrated, but in the absence of evidence it was assumed that they are like cast #7.

##### Initial Recalibration

Based on the post-cruise calibration the temperature was recalibrated in the SAM files to raise the values by 0.001Cº. Salinity was then recalculated just to study the effect on salinity. COMPARE was then rerun. As expected increasing the temperature lowers the salinity by about 0.001psu. This method cannot be applied to the downcast data or the effects of editing will be lost.

A test was run on cast #8 using the post-cruise temperature and conductivity calibration co-efficients and salinity was compared at a number of depths to see if the change is pressure-dependent. The salinity change ranged from 0.0081psu to 0.0085psu with no significant pressure dependence. (See 2004-20-0008CAL_STUDY.xls.)

The primary temperature will be recalibrated based on the bottle comparison. The primary salinity will be recalibrated using equation S0(cor) = S0 +8 * e-07* Pressure – 0.0008. The secondary salinity will be recalibrated using equation S1(cor) = S1 +6 * e-07 * Pressure – 0.0137. File 2004-20-recal1.ccf was prepared to apply these calibrations. This was applied to the EDT, MRG1 and SAM files. (Output: COR, MRGCOR1, SAMCOR1)

COMPARE was rerun for salinity and the calibration was found to have been done correctly. (See 2004-20-sal-comp2.xls.)

##### Special Fluorometer Processing

The COR files were clipped to 100db and put through REMOVE and HEADEDIT to produce files FCTD. The FCTD files were written on a CD-ROM for the use of Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts except #31 & 40 Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.REM)

For casts #31 and 40 the following channels were removed: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag channels. (Output: *.REM)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

There were many problems with the CTD on this cruise. The

differences between temperature sensors varied according to

whether the CTD was moving or not and upcasts look quite different

from downcasts. The primary salinity was usually cleaner than

the secondary, but for casts #31 and #40 there were sudden shifts

in primary values.

The post-cruise calibration suggests that the sensors

were reading lower than was shown in the bottle comparison, but

since there is likely a source of error besides sensor drift, the

bottle comparison was used to recalibrate. There remains some

doubt about the CTD salinity which may be low by up to 0.004psu.

The fluorescence and transmissivity data are nominal and

unedited except that some records were removed in editing T and S.

For casts #31 and 40 the comment was adjusted to reflect that the salinity may be low by up to 0.007psu.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S fluorescence and transmissivity profiles

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag for all casts except #31 and 40.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag for casts #31 and 40.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

At this point the final nutrient files became available; they had comments but no flags. The spreadsheet data was edited and converted to individual NUTS files. Editing was done to the comments in the NUTS files. These were then merged with the HDR files, sorted on pressure and named *.CHE.

Plots of sample number versus pressure were made to check for errors and showed that cast #1 contained no samples. This was a test cast with no sampling so the CHE file was deleted.

##### Conversion of SBE25 loop data

Marie Robert produced files combining averaged CTD values from the loop (1 value per second) and times and positions taken from ship’s logging information. The files produced by Marie were converted to IOS files. This was tricky because of variable time formats. CLEAN was used to add various items to the headers. ADD TIME CHANNEL was used to add a record number and then plots were made of T and S versus record number. The only major problems are in file #4 which has bad data around the time of the two computer crashes early in the cruise. CTDEDIT was used to replace the bad data with pad values. The files were bin-averaged in 5s bins and then thinned to 30s steps and the record number was removed. The file names were changed to 2004-20-0001.thn* where * was the number of the original thinned files. They were then joined. A record number channel was then added.

##### Calibration of SBE25 loop data

A SeaBird 25 was used in the loop instead of a thermosalinograph. This was not noted in the log book, but according to the configuration file the sensors were conductivity #2754, temperature #4054 and pressure gauge #290482. There was a fluorometer but no sensor number is entered in the con file. (this was later obtained from Jim Gower – WS713P.)

The calibrations were checked and all were correct.

The sensor history was checked. This CTD has only been used once and only near the surface; at that time the CTD was found to be close to the bottles and the pressure sensor was found to be reasonably accurate as well.

Temperatures were checked for a few casts (10, 32, 36, 44) trying to match the time of a rosette bottle with the SBE25 measurement. The loop temperatures were found to be slightly higher than the SBE911+ measurement, by 0.19 Cº, 0.24 Cº, 0.26Cº and 0.30Cº. Some of the variability would be due to mismatch of the level of rosette and loop intake.

Loop sample salinity data was obtained. The times were taken from the log assuming the samples came at the end of the cast for those cases in which there was a CTD cast. The SBE25 salinity was taken from Marie’s files and variability was estimated by recording the minimum and maximum values in a 6s-window. The SBE25 was found to be low by from 0.011 to 0.042psu, with the lowest values coming from the lowest variability cases. (SBE25-loop.xls)

A spreadsheet was produced containing all salinity data from rosette bottles in the top 6db. The start time of the casts was included in the spreadsheet. This differs from the time of bottle firing by as much as an hour, but finding the exact times is a slow process. The SBE25 salinity did not vary greatly through a cast and any variations were not systematic, so over many casts should not introduce a significant error.

For a first comparison with rosette sampling the spreadsheet was examined to pick out bottles from 3.6 to 4.4db. There were 4 such samples. Once again the variability was estimated. The SBE25 was found to be low by from 0.012 to 0.026psu with values of 0.017psu for the lowest variability cases.

Later more bottles were included, those from casts where the top 6m were well-mixed. Variability was estimated in the same way as for the loop samples. For 16 casts with variability less than 0.007psu, the bottles were low, on average, by 0.0143 and when 6 casts with variability less than 0.016psu were used the average was 0.0142. To ensure that the difference in time between the rosette and the beginning of the cast is not a major problem, the 4 casts checked for temperature were also checked for salinity difference. The SBE25 was found to be lower than bottles by 0.0131, 0.0184, 0.0213 and 0.0152psu, which is reasonably consistent with other results. (SBE25-rosette bot.xls)

Finally, the file 2004-20 loop joe.xls provided by Frank Whitney was used to prepare a comparison with all casts to check for time-dependence. This file contained surface data for each cast. It was used to prepare file 2004-20-loop-calibration.xls containing SBE25 pressure, temperature and salinity with SBE911+ data and bottle data. Plots were prepared and near-shore data was removed since there was huge variation in those regions. For the offshore casts there is only slight suggestion of time-dependence and given the noise level it does not justify a time-dependent recalibration. The average difference is 0.215Cº. For salinity the SBE25 is lower than the SBE911+ by 0.017psu and lower than the bottles by 0.019psu with no indication of time-dependence. These results are consistent with the other comparisons.

The SBE25 temperature is clearly high and the salinity clearly low. Recalibration was done by simple offsets of -0.021Cº and +0.017psu.

The fluorescence values in the loop files are very low at the beginning of the cruise and gradually increase to 5 (assumed to be in volts). The data does not look believable, and a check of fluorescence at 5db from the SBE 911+ shows no such trend. An error in gain cannot explain this result. Presumably there was some deposit building up through the cruise. The loop file was used to compare fluorescence and chlorophyll and suggest that the data from the first portion of the cruise may be believable, but since some doubt exists the channel will not be put in the archive.

REMOVE was used to remove the record # and fluorescence channels. EDIT HEADER was used to change the format of the temperature and salinity so only 2 decimal places are displayed and to add a comment briefly describing how the processing was done. The final file was named 2004-20-0001.tob.

##### Producing final files

Because of the complexity of this processing job, the header check was run on CHE and CTD files to check that the correct calibrations were done, the right channels selected in each case. No problems were found.

A cross-reference listing was produced.

The sensor history was updated.

Particulars

4. Computer crashed at end of cast.

6. Downcast only – could not close bottles.

7. Upcast corresponding to cast 6.

8. Downcast only due to blackout. Wrong station name – should be P5. Fixed

9. Upcast corresponding to cast 8.

10. Initial section run with pumps off. CTD back to surface and full cast then run.

21. Rotated SeaBird headers. Fixed

31. Primary salinity bad from 720 to 830db. Used secondary for archive.

40. Primary Temperature and conductivity bad for downcast 30 to 350db. Used secondary for archive.

54. Wrong station name – should be R9. Fixed.

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2004-20 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 20 August 2004 End: 2 September 2004 |  |  |  |  |  |  |
| Location: Gulf of Alaska |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Whitney F. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |  |
| 2* | SEABIRD | 25 | ? | No | Yes |  |

* - Used in loop only.

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2004-20

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 20/12/03 | Factory “ | 14/10/04 | Factory |
| Conductivity | 1763 | 19/12/03 | “ | 19/11/04 | “ |
| Secondary Temp. | 2106 | 06/05/04 | “ | 18/11/04 | “ |
| Secondary Cond. | 2102 | 07/05/04 | “ | 15/10/04 | “ |
| Transmissometer | 732DR | 24/10/03 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2229 | July 01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/2000 | Factory |  |  |

Make/Model/Serial#:	SEABIRD/25

Cruise ID#:		2004-20

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 19/06/02 | Factory “ |  |  |
| Conductivity | 2754 | 22/05/02 | “ |  |  |
| Fluorometer | WS 713P |  | IOS |  |  |
| Pressure Sensor | 290482 | 13/04/02 | Factory |  |  |