## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected the precision in Salinity:Bottle lost during the addition of HPLC. S.H. |
| ### 19 January 2021 | ### Added HPLC Data. S.H. |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 29-Nov-2004 | Added nutrient data to the rosette files. J.L. |

## PROCESSING NOTES

Cruise: 2004-21

Agency: OSAP

Location: Strait of Georgia / Juan de Fuca Strait

Project: JdeF / SoG

Party Chief: Masson D.

Platform: Vector

Date: June 21, 2004 – June 26, 2004

Processed by: Germaine Gatien

Date of Processing: September 23, 2004 – October 4, 2004

Number of original CTD casts: 73

Number of CTD casts processed: 73

Number of rosette casts: 25

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlab transmissometer (#498DR for casts 34-109), a Seabird Dissolved Oxygen Sensor (#0047), an altimeter (1024), a Seapoint Fluorometer (#2229) with a 10X cable, Biospherical QSP2300 PAR sensor (#4565) and Biospherical Surface PAR sensor QSR-2240 (#16504). The deck unit was a model 911 (#0508) and the logging computer was #FS03. The salinometer was a Portasal model 8410 (#59724). The DO sensor was on the primary pump and the fluorometer was on the secondary pump.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in good order.

The rosette log contains some errors, with some samples drawn that are not noted, and some noted but not drawn. There were several errors in the salinity sample numbers.

Some of the file names were in non-standard format.

The altimetry was unusually clean probably due to the absence of a weight on the rosette for most casts.

There are concerns about the temperature sensors. It is not clear which sensor is faulty. If it is the primary sensor, then there may be errors in salinity on the order of 0.001psu.

The casts in the Strait of Georgia had a very steady descent rate, so the quality of the data from that area is expected to be very high.

Three casts had bad primary salinity below 200db. This may have been caused by jelly fish encounters. The secondary channels were selected for these casts only.

The dissolved oxygen data in the CTD files should be considered:

- ±0.4ml/l from 0db–50db.

- ±0.2ml/l below 50db – 250db

- ±0.1ml/l below 250db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Some of the file names were non-standard – extra 0s in them. These were corrected.

There is a file 2004-21-0068.DAT which is empty and a file 2004-10-0068.DAT which has data in it. The latter appears to be the file for cast #68 of 2004-21. It will be renamed after conversion.

Bottle salinity, titrated chlorophyll and dissolved oxygen data were obtained with flag channels. Flags and comments had been added to the DO and CHL data. There is a flag channel in the salinity data, but no flags were deemed necessary.

Nutrient bottle data were not yet available.

The cruise summary sheet was completed.

The histories of the conductivity, pressure and dissolved oxygen sensors were obtained.

The calibration constants were checked for all instruments. The date of calibration and the calibration constant for the PAR sensor were fixed. The date of the calibration for pressure was fixed.

There was an offset of -0.6db in the pressure configuration. During 2004-04, 2004-05 and 2004-07 pressure offsets of +0.2db were found appropriate; in the case of 2004-04 a careful surface pressure test was done. In processing 2004-22 an offset of +0.4db was found appropriate. A test conversion was done using the configuration file used at sea with the -0.6db pressure offset. When the conductivity went to zero at the end of the cast, the pressure was about -1.25db; however, the pumps were on so the surface was probably a little less, say -1db. This suggests an offset of +0.4db is need. The offset was changed in the configuration file. Fine-tuning of the pressure offset can be done later if needed.

The new configuration file was named 2004-21-ctd.con.

##### Conversion of Raw Data

All data were converted. Cast 2004-10-0068.cnv was renamed 2004-21-0068.cnv after conversion. A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. The altimeter data is cleaner near the bottom than that found in other recent cruises.

Rosette files were converted using a start time of -2s and duration of 5s.

The station name was corrected in cast #38 CNV and ROS files.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

The deck unit was one of the newer ones that advances secondary conductivity channels so ALIGNCTD was not run on conductivity.

The dissolved oxygen channel was aligned by +5s based on tests on a few casts. After this step DO for the up and downcasts are much closer. This can be fine-tuned later if found necessary.

##### CELLTM

Tests were run with (alpha, 1/beta) set to (0.01,7) (0.01,9), (0.02,7), (0.02,9), (0.03,7), (0.03,9) and (0.0245,9.5). For two casts (0.03,9) was the best choice for both channels. For one plot (0.03,9) improved both channels in some areas, but in others all choices made things worse. During 2004-10, when the same sensors were used, the setting of (0.03,9) was found best for both channels.

All casts were put through CELLTM using (0.03,9) for both conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 55 | 360 | -0.0008 | +0.0005 | +0.006 | Fairly high, steady |
| 66 | 350 | -0.0008 | +0.0005 | +0.006 | Fairly high, steady |
| 93 | 415 | -0.0007 | +0.0003 | +0.0035 | High, steady |

All the differences were very noisy. The temperature differences are similar to those during 2004-10; the salinity and conductivity differences are slightly higher, but the maximum pressures are much lower so this is expected. As noted many times with these sensors, the secondary salinity is full of fine-scale noise with average excursions of ±0.002psu at 2000db, while the noise level of the primary is about ±0.0005psu. The secondary conductivity had a noise level of ±0.00015 S/m whereas the primary has noise on the order of ±0.00003. Quick checks on a few other casts show no others with as low a salinity difference as cast #93.

The temperature and salinity differences during bottle stops tend to be slightly lower, typically -0.0003Cº and +0.005psu than when the CTD is moving. The conductivity shows no significant difference. So part of the differences seen between salinity sensors during the downcast must be due something other than conductivity calibration; this limits the accuracy of the data since the bottle calibration will not address that error. Looking at temperature traces closely show lots of small-scale differences but there is no obvious evidence that one trace is better than the other. It was noted during 2004-10 and 2004-20 that the temperature differences were more pressure-dependent than usual. The pressure-dependence may not imply pressure is causing a problem; since the variations are generally lower at higher pressure this may just mean that slightly different response results in large differences in high gradient areas.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files.

CLEAN was run on the rosette files to add an event number and the output was named BOT.

All BOT files were plotted and no significant outliers were found.

##### Checking Headers

The header check and header summary were run. The only errors found were in station names for casts 7, 38 and 87; the name was also wrong in the log book for cast #7.

The cruise track was plotted and no problems noted.

The average surface pressure is 2.9db, a reasonable value for the Vector. A few casts were checked to see if conductivity looked reasonable near the surface. The only file found with near-zero values for pressure was cast #2 which is the one used for estimating the offset in the con file, so not useful in testing that offset. The pressure will not be adjusted further.

The surface checks shows no significant mixed layer, so the surface salinity samples will not be useful for calibration purposes.

A few casts were examined on screen. No problems were noted. The altimetry is unusually good. There were readings for every cast showing that the CTD was within 15db of the bottom for every cast.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheet was reordered by sample number and cast number, and edited to add cast numbers. It was saved as 2004-21sal.csv. That file was converted to individual SAL files.

The dissolved oxygen files (*.add) were created by the analyst and have quality channels and comments for flagged values.

The extracted chlorophyll data were delivered in spreadsheet form by the analyst and had quality channels and comments for flagged values. The channel names were changed to standard form and then  converted to individual CHL files. The comments were transferred to the headers of the relevant files. Notes from the analyst indicate that there are many errors in the rosette sheets concerning chlorophyll and nutrient sampling. In some cases samples were taken but not recorded, and in others samples were recorded but not taken.

The nutrient data was not yet available.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The SAL, ADD and CH: files were merged with SAMAVG in 3 steps. (Output: MRG1, MRG2, MRG)

##### 11. COMPARE

Salinity

An initial run of COMPARE turned up two extreme outliers in salinity. Checks of the salinity analysis sheets and rosette log were made to see if errors had been made. The following conclusions were reached:

- Cast #58- Sample #186 is entered on the salinity analysis sheet but the rosette log shows no salinity sample from that bottle. There was a salinity sample for #188, but that is not given on the salinity analysis sheet, so it is likely that the sample number was just written down wrong either on the sample bottle or at the analysis stage. The dissolved oxygen values look reasonably close to the CTD for the two bottles, so it does not appear likely to be a matter of a misfiring bottle. The salinity result is different from the CTD for the level of sample #186 by 5.6psu and for the level of sample #188 by 0.25psu. The sample number was changed in the SAL file and flagged “c” with an explanation for the switch.

- Cast #58 – Samples 174 and 175 appear to be reversed. The titrated dissolved oxygen samples agree with the CTD, so this does not appear to be a matter of misfiring. The numbers were changed in the SAL file, flagged “c” and a note of explanation was put in the header.

- Cast #69 – Samples number 222 and 232 appear to have been reversed. One is about 4psu higher than the CTD and the other low by about the same amount. The deep sample does not resemble a nearby bottle. When reversed the values are reasonable compared to the CTD and compared to the nearby bottle. The dissolved oxygen data looks reasonably close to the CTD so whatever error occurred seems to affect salinity only. Either the bottles were labelled wrong or were recorded wrong on the analysis sheet. The sample numbers were changed in the SAL file and were flagged “c” and an explanation put in the header.

- Cast #83 – Sample 247 from 300db differed from the CTD by 0.03psu, but there does not appear to be any error in the labelling. It looks like it may have been collected from Niskin #1 rather than #2 as recorded in the rosette log. It was later decided to flag this salinity sample as “d”.

All changes to the SAL files were approved by the chief scientist, Diane Masson.

COMPARE was rerun after these corrections. There is a lot of scatter in the comparison as is expected given the shallow sampling and the active mixing in the region. When all data is included below 300db there is evidence of similar time dependence for both sensors. This is probably a geographic dependence rather than due to sensor problems. The primary salinity is low by about 0.003 and the secondary high by about 0002 or 0.0025 depending on how many outliers are rejected. Looking at just the deepest bottles for each cast (with the exception of cast #61 for which the 2nd deepest was selected because the deepest was an outlier) a fairly flat distribution with pressure and time is seen. From that comparison the primary salinity is found to be low by about 0.0014psu and the secondary high by 0.005psu. When only the 5 bottles from below 320db are used the primary is seen to be low by 0.0014psu and the secondary high by 0.0035psu. (See 2004-21-sal-comp1.xls)

Dissolved Oxygen

An initial run of COMPARE turned up some problems in the ADD files. The name of the flag channel changed after cast #55. This was fixed using HEADEDIT and Merge was rerun with the newly named ADD1 files.

COMPARE was run using titrated dissolved oxygen data and plots were prepared of differences between CTD and titrated values versus pressure, DOX and cast number. The best fit was versus DOX value. Some outliers were identified and removed from the comparison. Using those values the fit is:

DOX (Titrated) = 1.0437 * DOX (CTD) + 0.070

The calibrations from other recent cruises using this equipment were:

DOX (Titrated) = 1.0461 * DOX (CTD) + 0.1206 (2004-10 Line P June)

DOX (Titrated) = 1.0336 * DOX (CTD) + 0.0751 (2004-07 April, near-shore)

DOX (Titrated) = 1.0772 * DOX (CTD) + 0.0331 (2004-04 March, near-shore)

The results are reasonably close to the April near-shore results. There has been some suggestion of a gradual drift in this instrument, but these results suggest otherwise. The differences from cruise to cruise may be mostly based on the range of DO sampled. (See 2004-21-dox-comp1.xls)

The above fit will later be used to recalibrate the CTD DOX channel. The only severe outliers have already been flagged.

Extracted Chlorophyll

Spreadsheet 2004-21chlarc2.xls was obtained from Melanie Quenneville. A few changes were made to headers and it was saved as 2004-21chl.csv. This file was converted to CHL files. Comments were transferred from the spreadsheet to the relevant CHL files.

A comparison of fluorescence and extracted chlorophyll shows a lot of scatter with an average value of Fl = 0.78 * CHL. A study was made of the ratio as a function of time of day. The points with the highest ratio occur mostly between 10am and 6pm, but there are many values with lower ratio within those bounds. So there is no clear pattern. Data was also examined at different levels. The ratio goes up steadily with pressure from 0.67 near the surface to 1.03 at 40db. Only the data at 6 and 10db shows any hint of time-dependence and it is not very convincing. (See 2004-21-fluor-chl.xls.)

##### 12. SHIFT

Conductivity

Tests were run on the primary conductivity channels to determine what shift reduces instabilities best without oversmoothing; results were examined in T-S plots after running DELETE. The best results came from an advancement of -0.2 records. That would mean a net advancement of about -0.065s, since the channel was already advanced by +0.073s earlier. This is the same shift as was used for 2004-10 and 2004-22.

All casts were put through SHIFT using -0.2 records for the primary conductivity. (Output *.SHFC)

No tests were done initially on the alignment of the secondary channel, since it was not expected that it would be selected for the archive. However, three casts were found to have poor primary salinity. A SHIFT of -0.2 records was applied during 2004-09 when the same sensors were used.  So that was applied to the three special casts only and the results checked. SHIFT did reduce the noise in the salinity slightly.

Dissolved Oxygen

A shift equivalent to +120s had been applied earlier using ALIGNCTD. Tests were run using advancements of -20 records to +20 records to determine if further shifting the data would improve the data. The choice of +120 records (~5s) was found best, so no further shift was applied.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was 1 to 2s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing since the secondary sensors have a lot of fine-scale noise.

Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used. It was decided in the courses of editing to use the secondary sensors for casts #67, 75, and 91.

The following casts required no editing: 29, 51, 55, 59

The following casts required heavy editing: 5-10, 12, 14-21, 24, 25

All other casts required only lightly editing.

The descent rate was extremely noisy for most of the casts in Juan de Fuca Strait, but was unusually smooth for the casts in the Strait of Georgia.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensor has produced salinity high by 0.0004psu in March and low by 0.0003 in April, low by from 0.001 to 0.003 in May and low by 0.0015 in early June. In July it was low by from 0.002 to 0.0025psu. The secondary sensors gave salinity high by 0.0017, 0.0025, 0.0016 and 0.0018psu during cruises from March to June of 2004.

Historic ranges – All data fell within the local climatology with the exception of cast #14 which had high temperature from 60 to 100db and low salinity from 60 to 120db. This is an area of great variability and is not considered evidence of instrument malfunction.

##### Initial Recalibration

The bottle comparisons for this cruise are expected to be quite scattered because there are no really deep samples and the area is one of active mixing. COMPARE indicates that the primary salinity is low by from 0.0001psu to 0.0033psu with -0.0014 for the deepest bottles. The sensor history suggests a drift to lower values. Recalibration by adding 0.0015psu looks like a reasonable choice for this cruise. For the secondary sensors the history shows less variation. The estimates from COMPARE were from +0.002 to +0.005psu, with +0.0035 from the deepest bottles. These sensors were not used for 2004-22. Recalibration by subtracting 0.003psu is a reasonable choice.

File 2004-21-recal1.ccf was prepared to apply an offset of +0.0015 to the primary salinity, -0.003 to the secondary and to recalibrate the dissolved oxygen using the results of section 11; it was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for salinity and DOX and the calibration was found to have been done correctly. (See 2004-21-dox-comp2.xls and 2004-21-sal-comp2.xls.)

##### Final Dissolved Oxygen comparison

Another run of COMPARE was made using metre-averaged, thinned downcast files. When outliers were excluded the differences versus pressure show an offset of about 0.011 ml/l with the CTD reading a little high and no significant pressure-dependence. When the differences are plotted against DOX values the results are similar.

The offset does not seem likely to be due to an imperfect alignment since a small error would cause little change at depth. It is likely due to the response time being a little slow, so the sensor gives slightly higher values than expected on the downcast. Only the downcast files are affected by this, the bottle files do not require recalibration because the sensor had time to equilibrate. During 2004-04, -05, -07 and -10 when the same instrument was used an offset of -0.03ml/l was found appropriate.

The COR1 files and the thinned downcast files were recalibrated by subtracting 0.011ml/l using file 2004-21-recal2.ccf. (Output: COR2)

COMPARE was rerun using the recalibrated thinned files and the average difference was +0.0003ml/l. (See 2004-21-DOX-comp4.xls)

##### Special Fluorometer Processing

The COR2 files were clipped to 100db and put through REMOVE and HEADEDIT to produce files FCTD. The FCTD files were written on a CD-ROM for the use of Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	 Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts except #67, 75 and #91. (Output: *.REM)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from casts #67, 75 and 91. (Output: *.REM)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and unedited

except that some records were removed in editing T and S.

The DO data in the CTD files should be considered:

•	±0.4ml/l from 0db–50db.

•	±0.2ml/l below 50db – 250db

•	±0.1ml/l below 250db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T and S profiles. Separate profile plots were prepared with PAR, DOX, FL and Transmissivity versus pressure.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure. All casts except #67 were put through REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag. For cast #67 the primary temperature and salinity were removed instead of the secondary.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### 24. Check of altimeter readings in headers

A list was made of the altimeter reading in the headers of these files to ensure the information is reasonable. All the data looked good. There was a lot less noise than usual in the altimetry which is believed to be due to the absence of a rosette weight on most casts.

Particulars

30. Jellyfish

38. Error in station name.

53. Jelly Fish

66. Jelly Fish

67. Jelly Fish – bad primary salinity.

75. Bad primary salinity

77. The dissolved oxygen file, ADD, originally named as #76 was changed to #77.

83. Noise in transmissivity around 250db.

91. Bad primary salinity from 290db to bottom.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-21

Dates: Start: June 21, 2001 End: June 26, 2004

Location: Juan de Fuca Strait / Strait of Georgia

Vessel: Vector

Party Chief: Masson D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2004-21

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 20/12/03 | Factory “ |  |  |
| Conductivity | 1763 | 19/12/03 | “ |  |  |
| Secondary Temp. | 2095 | 05/09/03 | “ |  |  |
| Secondary Cond. | 1764 | 05/09/03 | “ |  |  |
| Transmissometer | 498DR | 07/04/04 | IOS |  |  |
| Dissolved Oxygen | 0047 | 13/01/04 | Factory |  |  |
| Altimeter | 1024 |  |  |  |  |
| PAR | 4656 | 19/12/03 | ? |  |  |
| Surface PAR | 16504 | 2 Jan 04 | ? |  |  |
| Fluorometer | 2229 | July 01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |