## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2004-22

Agency: OSAP

Location: WCVI/Effingham/Juan de Fuca Strait

Project: Seabreeze/Effingham

Party Chief: Thomson R.

Platform: John P. Tully

Date: July 5, 2004 – July 16, 2004

Processed by: Germaine Gatien

Date of Processing: August 19, 2004 – August 31, 2004

Number of original CTD casts: 97

Number of CTD casts processed: 96 (1 cast has no pumped channels)

Number of rosette casts: 47 (2 casts have no pumped channels)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlab transmissometer (#498DR), a Seabird Dissolved Oxygen Sensor (#0047), Altimeter OA-916D (#1024) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0508) and the logging computer was #FS03. The salinometer was a Portasal model 8410 (#58879).

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in good order.

There is more small-scale noise than usual in both salinity channels even when the descent rate is quite steady. This suggests flow-rate irregularity, but may be due to the fact that waters were sampled with an unusual combination of large temperature gradients and small salinity gradients.  The noise is generally two-sided and the error will be minimized by metre-averaging.

There is also some very fine-scale noise in the secondary salinity. This has been observed from other sensors mounted on the secondary pump. It is two-sided noise so not a major problem.

The salinity near the surface in Effingham Inlet is frequently very high, gradually returning to believable values by about 30db. Given that the temperature data is not obviously affected, this is likely to be due to something in the plumbing getting clogged and clearing when the pressure is high enough. The problem sometimes recurs on the upcast near the surface. This suggests biological fouling. For 6 casts the upcast sections were selected for the archive and for those, salinity should be considered ±0.05psu in the top 15m and ±0.01psu below that.

The pumps were turned off prematurely for two casts. A few seconds should be allowed after the firing of the final bottle before turning off the pumps. The pumps were not turned on at all for two other casts. In one case this was noted at sea and the cast was rerun, so the first cast will not be archived. For the other case it was not repeated, so the pumped channels were removed in the file for the archive.

The dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l from 0db–150db.

- ±0.1ml/l below 150db – 300db

- ±0.05ml/l from 300 – 1500db

- unreliable below 1500db (that channel was removed from the final CTD files for P>1500)

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Bottle salinity and dissolved oxygen data were obtained; flags and comments had not been added. There was no chlorophyll or nutrient sampling.

The cruise summary sheet was completed.

The histories of the conductivity, pressure and dissolved oxygen sensors were obtained.

The calibration constants were checked for all instruments. The calibration for the secondary temperature was incorrect for the first cast, but was fixed after that cast. The dates of the calibrations for the pressure sensor and transmissometer were wrong. There was an offset of -0.6db in the pressure configuration for the early casts. From cast #24 onwards an offset of +0.4db was applied based on a pressure check on deck. During 2004-04, 2004-05 and 2004-07 pressure offsets of +0.2db were found appropriate; in the case of 2004-04 a careful surface pressure test was done. For 2004-10 an offset of +0.2db was also used, though it was considered that it might be a little low. An offset of +0.4db will be applied and the results checked later.

Calibration file 2004-22-CTD.con was prepared with corrections to the calibration dates and a pressure offset of +0.4db

##### Conversion of Raw Data

All data were converted. A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. There were a lot of spikes in cast #1. The altimeter data is very noisy, though a useful signal is apparent, at least visually, amidst the noise.

Rosette files were converted using a start time of -2s and duration of 5s.

The rosette files were then converted to IOS SHELL files. The rosette file for cast #36 could not be converted to IOS SHELL file until it was recreated without an altimeter channel (The altimetry data in the CTD file for this cast was fine, so the information from the header was copied to the header of the bottle file.) There were warning messages from the conversion of casts #42 and 58; the pumps were not turned on for those two files. There is a note in the log about #42, but not #58.

CLEAN was run to add the event number to the rosette files. Warnings were issued for casts 42, 58, 62 and 68 because of problems in the pump status channel. Two of these casts (42, 58) were run with the pumps off and the data from the bottle firing may be needed later for comparison with bottles. Two casts (62, 68) just had the pumps turned off prematurely at the end of the last bottle firing. Since it only means the loss of 4 and 5 records respectively from the last bottle, this is not a problem. Casts #42 and #58 were reconverted without the pump status channel.

All BOT files were examined to see if there were spikes in the temperature or salinity data that should be edited. Cast #38, bottle #4 was edited lightly. The output file was copied to *.BOT. Other noise was mostly near the surface and not editable.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

The deck unit was one of the newer ones that advances secondary conductivity channels so ALIGNCTD was not run on conductivity.

The dissolved oxygen channel was aligned by +5s based on tests on a few casts. After this step DO for the up and downcasts are much closer. This can be fine-tuned later if found necessary.

##### CELLTM

Tests were run with (alpha, 1/beta) set to (0.01,9),(0.02,7), (0.03,7), (0.02,9), (0.0245,9.5) and (0.03,9). Most choices improved the data, but the optimal setting varied from cast to cast and even within a cast. A setting of (0.02,7) was best overall for the primary and (0.02,9) for the secondary channels.

All casts were put through CELLTM using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2 | 1400 | -0.0004 | +0.0004 | +0.005 | High |
| 5 | 1400 | -0.0007 | +0.0004 | +0.007 | High |
| 6 | 1400 | -0.0004 | +0.0004 | +0.006 | High |
| 9 | 1400 | -0.0012 | +0.0005 | +0.0075 | High |
| 12 | 1400 | -0.0014 | +0.0005 | +0.008 | High |
| 18 | 1400 | -0.0012 | +0.0007 | +0.0095 | High |
| 22 | 1400 | -0.0010 | +0.0006 | +0.0085 | High |
| 71 | 1400 | -0.0014 | +0.0007 | +0.010 | High (slowed ~ 1400) |

All the differences were very noisy. The temperature differences were more pressure-dependent than usual. Conductivity showed little variation with pressure but there does appear to be some time-dependence in the early part of the cruise. The salinity differences were pressure-dependent and increased through the early part of the cruise. The secondary salinity is full of fine-scale noise with average excursions of ±0.002psu at 2000db, while the noise level of the primary is about ±0.0005psu. This is caused by fine-scale noise in the conductivity, not the temperature. The secondary conductivity sensor had been changed since the previous use, but the noise looks the same as from other recent cruises. This suggests that the problem is in the plumbing, not the sensor. The fluorometer was on the secondary pump; could this be causing a problem?

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

Cast #5 could not be converted. Attempts were made to convert versions of the file that had been edited in a variety of ways. Success came after the altimeter channel was stripped from the CNV file. The signal in that channel was very strange with low values at mid-depths and an extremely high noise level. Judging by the log book entries for maximum depth and maximum sampling, it does not appear that the CTD was ever closer than 50m from the bottom.
Casts #57 and 89 had rotated Sea-Bird headers which had to be corrected to get positions in the headers.

##### Checking Headers

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The header check and header summary were run. Errors were found and corrected in the station names for casts #68 and #80.

The cruise track was plotted and no problems noted.

The average surface pressure is 1.8db, which is lower than usual for the Tully, but there was a lot of scatter. When casts run by Doug Anderson are examined separately the average is higher, around 2.2db. For those casts that have a low surface reading the pumps were not turned on until the pressure was closer to the usual. A few casts were checked to see if conductivity looked reasonable near the surface. For cast #32 when pressure was 0.2db at the end of the cast, the conductivity looked like “in-water” values but low enough to be near the surface. The pressure appears to be close to correct and will not be adjusted further.

A few casts were examined on screen. No problems were noted.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheet was in the original *.txt format. It was put through the Salinometer Analysis Averaging routine. No flag were needed (based on remarks on the analysis sheets and discussion with Doug Anderson.) The file was reordered on cast # and saved as 2004-22sal.csv. That file was converted to individual SAL files.

The dissolved oxygen files (*.OXY) were created by the analyst but do not have quality channels and comments for flagged values. Channel Flag:Oxygen:Dissolved was added with the output named *.ADD. Flags and comments were entered for 3 samples based on the notes on the rosette log sheets and discussion with Doug Anderson.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Cast #49 had 10 bottles but there are entries in the rosette log for only 9. It was assumed that the last bottle was not sampled, but this may not be correct. After COMPARE it is hoped that this will be clear.

The SAL and ADD files were merged with SAMAVG in two steps. (Output: MRG1, MRG)

##### 11. COMPARE

Salinity

COMPARE was run. There is a lot of noise in the comparison for both sensors with suggestions of pressure-dependence and time-dependence only for the secondary. The primary salinity is low by about 0.0025psu and the secondary is high by an average of 0.005psu (See 2004-22-sal-comp1.xls)

There were no extreme outliers.

Dissolved Oxygen

COMPARE was run using titrated dissolved oxygen data and plots were prepared of differences between CTD and titrated values versus pressure, DOX and cast number. The best fit was versus DOX value. 23 outliers were identified and the relevant files examined for errors. Errors were found in the addsamp file for one cast. And 8 samples were from casts run with the CTD pumps off. COMPARE was rerun after corrections and without the two casts with pumps off. A few outliers were then removed from the comparison, including values from 1500db downwards since the sensor is not expected to give good values at those pressures. Using those values the fit is:

DOX (Titrated) = 1.0785 * DOX (CTD) - 0.0309

The sensor was last used for 2004-10 in June 2004 when the fit was:

DOX (Titrated) = 1.0772 * DOX (CTD) + 0.0331

So while the slope is close to that in June, the offset is quite different. This may be because this cruise sampled an extremely large range of values, many of which were very low.

The calibration from other recent cruises using this equipment were:

DOX (Titrated) = 1.0336 * DOX (CTD) + 0.0751 (April, near-shore)

DOX (Titrated) = 1.0526 * DOX (CTD) + 0.0291 (February, Line P)

The slope appears to be drifting fairly gradually.

The fit against cast number is remarkably flat given the different regions sampled.

The results for cast #49 suggest that the sample numbers were assigned correctly in the previous step.

(See 2004-22-dox-comp1.xls)

The above fit will later be used to recalibrate the CTD DOX channel.

All remaining outliers were investigated. Three had already been flagged in the ADD files and one was from below 1500db. Ten others were associated with noisy CTD records.

##### 12. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was 1 to 1.5s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Dissolved Oxygen

A shift equivalent to +120s had been applied earlier using ALIGNCTD. Tests were run using advancements of -20 records to +20 records to determine if further shifting the data would improve the data. The best choice varied from cast to cast and from one depth to another. Near the surface no further shift looked best, but from about 200db downwards +20 records looked better. Overall the choice of +20 records for a net shift of +140 records (~6s) was found best. A shift of +20 records was applied to all casts.

Conductivity

Tests were run on the primary conductivity channels to determine what shift reduces instabilities best without oversmoothing; downcast results were examined in T-S plots. The best results came from an advancement of -0.2 records. That would mean a net advancement of about -0.065s, since the channel was already advanced by +0.073s earlier. The same result was found for 2004-10 using the same sensor.

All casts were put through SHIFT using -0.2 records for the primary conductivity. (Output *.SHFC)

No shift was applied to the secondary channel since it is unlikely to be used for the archive and the noisiness of the data would make it difficult to detect a difference.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There was a warning for cast #53. Spikes in salinity led to almost all data being removed. The problem arose in the three runs of SHIFT when entries of ********* were written in the salinity column. The three SHIFT steps were rerun; after each step the ********* entries were replaced with pad values -99.0000. Then DELETE was rerun without problem.

##### DETAILED EDITING

The primary sensors were chosen for further processing for all but one cast since the fit against bottles showed less pressure-dependence than for the secondary. Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used. Detailed observations follow:

- Cast #1 was extremely noisy with skips in pressure and noise in salinity that look like evidence of flow-rate irregularity. Other casts had similar salinity noise but not the pressure problem. The secondary sensors had even worse salinity noise. A test was made on cast #14 to see if varying the shift in the primary conductivity could reduce this noise. No setting was found that worked better. The noise is most noticeable in areas of large temperature gradient and small salinity gradient. In many cases metre-averaging will remove the problem and the errors are relatively small. Editing of noisy salinity was done where that was not the case and it was obvious how to edit.

- Cast #16 had high and noisy temperature in the bottom 40m; the upcast looks similar so this is presumed to be real.

- There were a number of casts in Effingham Inlet with normal temperature and high salinity near the surface (>34psu) gradually settling down to believable salinity values at a pressure of about 30db. The secondary sensors are not affected as much, though there are odd features there too. Usually the upcasts are ok, but there are some problems in the top 10-15db for others. This fits the behaviour expected when the bleed-valve of the CTD becomes clogged, but there is no bleed-valve for rosette casts. So it could be some other plumbing problem. Not all casts were affected in this way. It seems likely this is due to biological fouling of some sort. The transmissivity was very low near the surface, frequently less than 10%/m. Casts with such problems in the downcast are #35, 41, 44, 48 and 53. Each was studied in detail to decide whether the downcast or upcast, primary or secondary sensors would be useful. The upcast primary data was found to be best for all these casts, but since most were rosette casts they needed heavy editing to remove the shed wakes caused by the many stops.

- The DO comparisons show large differences between titrated values and bottles near the surface for casts 41, 44, 45, 47 and 48, which may well be due to the same problems. There was no salinity sampling in Effingham Inlet.

- Cast #93 had bad primary salinity data from 127-140db during the downcast. It was decided to use the secondary salinity for this cast only.

- All casts required some editing, and the following casts required heavy editing: 2, 4-6, 11, 12, 29, 41, 44, 48, 53.

- Cast #42 was not edited; the pumps were not on and the cast was rerun as cast #43.

- Cast #58 was edited for temperature only since the pumps were not on and salinity will have to be stripped later.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensor has produced salinity high by 0.0004, low by 0.0003 and low by 0.001 in March, April and June of 2004, but the March cruise had few bottles and the June cruise had a lot of noise in the comparison. This was the first use of the secondary sensors since they were last calibrated.

Historic ranges –All casts were plotted with historic ranges superimposed. The temperatures were at or above the historic maxima from about 2000db to 2400db in the Endeavour Ridge area. This is presumed to be due to plumes and not symptomatic of sensor problems. There were also temperatures near the maxima around 1000db for the more southerly casts in that area. Temperature was low and salinity high from 30 to 40db for cast 97 in Juan de Fuca Strait. None of these observations suggest problems with the sensors – they are likely due to real changes in the ocean or sampling in slightly different areas from those in the local climatology.

Repeat Casts – Casts #1 and #11 were close together, but more than a day apart. The differences along γ-lines at about 1500db are about 0.003Cº and 0.0005psu. At around 2200db there are more significant differences, especially in temperature, but this is the depth of the plume so this is expected.

##### Initial Recalibration

COMPARE indicates that the primary salinity is low by from 0.0025psu and the secondary high by about 0.005psu, but there is a lot of noise in the comparison. The history of the primary sensors suggests a slow drift to lower values. Recalibration by adding 0.002psu looks like a reasonable choice.

File 2004-22-recal1.ccf was prepared to apply an offset of +0.002 to the primary salinity and -0.005psu to the secondary salinity (for cast #93 only) and to recalibrate the dissolved oxygen using the results of section 11 (1.0785 * DOX (CTD) - 0.0309); it was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for salinity and DOX and the calibration was found to have been done correctly. (See 2004-22-dox-comp2.xls and 2004-22-sal-comp2.xls.)

##### Final Dissolved Oxygen comparison

Another run of COMPARE was made using metre-averaged, thinned downcast files. The differences versus pressure show the downcast CTD values low by an average of 0.0009ml/l with no significant pressure-dependence nor DOX-dependence. Other recent cruises using this equipment had an offset of about 0.03 ml/l with the downcast CTD reading high. The different result from usual may be linked to the calibration offset being quite different. No further recalibration will be done. (See 2004-22-DOX-comp3.xls)

##### Special Fluorometer Processing

The COR files were clipped to 100db and put through REMOVE and HEADEDIT to produce files FCTD. The FCTD files were written on a CD-ROM for the use of Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from cast #93 only. (Output: *.REM)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all other casts. (Output: *.REM)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and unedited

except that some records were removed in editing T and S.

The DO data in the CTD files should be considered:

•	±0.2ml/l from 0db–150db.

•	±0.1ml/l below 150db – 300db

•	±0.05ml/l from 300 – 1500db

Additional comments were added to casts 35 through 56 to indicate that there were problems with the pumped channels in the top 30m and the data should be used with caution. For those casts for which the upcast was used a comment was put in that the salinity should be considered ±0.005 at depth, and ±0.01 in the top 15m.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T and S profiles. Separate profile plots were prepared with Temperature, DOX and FL versus pressure. Transmissivity was included in the page plots.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag from all files. (Note cast #93 was not a rosette cast.)

For casts #42 and 58 the pumped channels (Oxygen:Dissolved:SBE, Fluorescence and Salinity:T0:C0) were also removed since the pumps were not turned on.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE. A separate comment was entered into the headers of casts from Effingham Inlet to warn users that the pumped channels near the surface may have been affected by pump problems.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### 25. Check of altimeter readings in headers

A list was made of the altimeter reading in the headers of these files to ensure the information is reasonable. No obvious errors were found. A few casts were checked in detail and the header info was correct.

Particulars

- Error in calibration constants for secondary temperature. Even with the error fixed, the secondary temperature and salinity are bad. Many spikes, some pressure skips.

36. Could not be converted until altimetry channel dropped from conversion.

42. Pumps off, but bottles fired. Bottle file stripped of pumped channels. Downcast file was deleted since #43 was a repeat at this site with pumps on.

43. Pumps on, no bottles.

58. Pumps not turned on. Bottles fired. Pumped channels deleted from both rosette and CTD files.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-22

Dates: Start: July 5, 2001 End: July 16, 2004

Location: WCVI/Effingham/Juan de Fuca Strait

Vessel: John P. Tully

Party Chief: Thomson R.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2004-22

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 20/12/03 | Factory “ |  |  |
| Conductivity | 1763 | 19/12/03 | “ |  |  |
| Secondary Temp. | 2106 | 06/05/04 | “ |  |  |
| Secondary Cond. | 2102 | 07/05/04 | “ |  |  |
| Transmissometer | 498DR | 07/04/04 | IOS |  |  |
| Dissolved Oxygen | 0047 | 13/01/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2229 | July 01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |