## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected the Salinity:Bottle precision that was lost during the addition of HPLC. S.H. |
| 19 January 2021 | Added HPLC Data. S.H. |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2004-26

Agency: OSAP

Location: Strait of Georgia / Juan de Fuca Strait (maps on last 2 pages)

Project: JdeF / SoG

Party Chief: Masson D.

Platform: VECTOR

Date: 13 September 2004 – 17 September 2004

Processed by: Germaine Gatien

Date of Processing: 1 February 2005 – 12 February 2005

Number of original CTD casts: 70

Number of casts processed: 70

Number of rosette casts: 23

Number of rosette casts processed: 23

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlabs CSTAR transmissometer (#723DR), a PAR sensor (#4656), a Surface PAR (#16504) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit is unknown. The salinometer was a Portasal model 8410 (#59724).

# SUMMARY OF QUALITY AND CONCERNS

The log book contained no information about the equipment used.

Some of the Rosette log sheets were missing. Photocopies were obtained from Janet Barwell-Clarke, but the quality of the copies is poor so some information is missing. It is recommended by Janet that more consistent pen/pencil types be used on the sheets. When pencils are used, a soft lead appears to give the highest quality photocopies.

Two salinity samples noted on the rosette log are missing from the analysis sheets and could not be found. There is no note of explanation in either the Daily Log Book or the rosette log sheets.

There were flags, but no comments in the salinity and dissolved oxygen bottle analysis data; it was clear from the rosette log sheets and the salinity analysis sheets what the problems were. Some surface salinity samples were gathered but not noted on the rosette log sheets. In some cases entries were erased from the rosette log sheets, but the samples were gathered.

The transmissivity and fluorescence data are unedited, except where records were removed in the editing of temperature and salinity.

Comparisons of CTD and bottle salinity were found to be unreliable near the bottom; the most likely cause is the effect of sediments in near-bottom sampling. For this region bottles fired close to the bottom should not be selected for calibration purposes.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

The bottle chlorophyll, salinity, nutrients and titrated dissolved oxygen data were obtained. There were flags in all files, and comments in the chlorophyll and nutrient data.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked.

The only errors were in the pressure configuration – the date was wrong and an inappropriate pressure offset had been entered. The offset was changed from -0.6db to +0.4db. The latter value has been found appropriate in recent use of this CTD. The file was saved as 2004-26-ctd.con.

There is also post-cruise calibration information available, so a second configuration file was prepared, 2004-26-ctd-post.con. I will refer to data from the two calibrations as pre-cals and post-cals. Tests were run on a few rosette casts to see which calibrations looked closer to the bottle results. Unfortunately, this area is not one where bottle comparisons are terribly helpful since most casts are shallow and have large salinity gradients.

Three casts were checked that had bottles near the bottom. In every case the sample from Niskin #2 looked much closer to the data that derived from the post-cals. However, the data from Niskin #1 is lower than both sets, but closer to the pre-cals. It often seems that the bottom bottle provides data that is low compared to all other bottles; it is not known if this is due to a bad Niskin or if bottle flushing is poor at the bottom. One thing that was clear is that the post-cals had smaller differences between salinity channels than the pre-cals, which suggests that the post-cruise calibrations are better. The primary salinity was higher by 0.0086psu after recalibration and the secondary went up by 0.0017psu. The differences were about 0.006psu with pre-cals and <0.001psu with post-cals. The primary channel appears to have a much lower noise level and is likely to be chosen for the archive.

The history of the conductivity and pressure sensors was found.

##### Conversion of Raw Data

The post-cruise calibrations were selected; all data was converted using 2004-26-ctd-post.con.

A few casts were examined and all expected channels are present. The up and downcasts are reasonably similar and the pairs of sensors are close during both upcast and downcast. Cast #11 had bad secondary conductivity for the upcast. The fluorescence, transmissivity, PAR and surface PAR look reasonable.

Rosette files were converted using a start time of -2s and duration of 5s, using 2004-26-ctd-post.con.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.

##### ALIGNCTD

It is unknown what deck unit was used, but in recent cruises all of them seem to advance both conductivity channels. If this is incorrect the problem will be found and corrected when SHIFT is run.

##### CELLTM

A few casts were identified as being sufficiently deep and having a quiet descent rate without stops for bottles. Tests were run on these casts with (alpha, 1/beta) set to (0.1, 7), (0.01, 9), (0.02, 9), (0.03, 9), (0.02, 7) and (0.03, 7) and (0.0245,0.5). The best results varied from feature to feature with only slightly differences between (0.02, 7) and (0.0245, 9.5). CELLTM was run using (0.02, 7) on both C0 and C1.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 12 | 300 | ~0 | -0.00008 | -0.0006 | Moderate, moderate |
| 27 | 300 | ~0 | -0.0001 | -0.001 | Very steady/mod. |
| 48 | 300 | ~0 | -0.00002 | -0.001 | Very steady/mod. |

These are very low differences and show that the choice of post-cruise calibrations worked well.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers to the headers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files. CLEAN was run to add event numbers and the output files were named BOT.

All BOT files were plotted and some noise was noted in #49, 55, 62 and 71. CTDEDIT was used to edit casts #49, 62 and 71. The output ED1 was then copied to BOT.

##### Checking Headers

The cruise track was plotted and no problems noted.

A header summary, header check and cross-reference listing were produced. There were errors in 3 station names. These were corrected in the headers.

The average surface pressure is 2.78db which is high for the Vector. A few files were examined and values were found with pressure at the end of the upcast at 0.3db and in-water conductivity. Moreover, two files had a starting pressure of about 1db. So the average surface pressure appears to reflect a deeper than usual start to casts, not an error in the pressure offset.

##### 11. SHIFT

Conductivity

Tests were run on several casts with various shifts of conductivity and the best results overall were with an advancement of -0.2 records to the primary. For the secondary the best choice varied from feature to feature from -0.2 to -0.5 records with -0.4 records looking best overall.

All casts were put through SHIFT using settings of -0.2 and -0.4 records for the primary and secondary conductivity, respectively.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between the two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this cruise the results varied from +0.9s to +1.8s. A shift of 1s (+24 records) has been used for most cruises in the past and was applied to this data.

##### 12. BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. The addsamp file was edited to add sample numbers taken from the rosette sheets.

The salinity analysis spreadsheet lacked entries for event numbers; this is required in order to serve as a break point in the conversion to individual files and for naming the output files. Event numbers were added and the file converted to SAL files. The samples from cast #1 are missing and there is no note of explanation in the log or rosette sheets. There was only 1 flagged value and no comment to go with it. The salinity analysis sheet shows that 4 readings were needed so the comment “required 4 readings to get stable reading” was added. There was some surface sampling that is not recorded on the rosette log sheets. In one case the salinity column was checked for two surface samples and then erased, but samples were gathered and analyzed.

The ADD files had flags but no comments and the spacing was wrong in the files with flags. The comments were transferred from the rosette sheets and the spacing was fixed. All the flags were “d”. Based on the standards agreed to in October 2004, most were changed to “c” unless clearly outliers. During cast #7 sample #29 was misnamed #28 and because of the way the oxygen program works the sample for #28 had to be named something else; the analyst chose #928. The ADD file was edited to correct the sample numbers. The rosette log record made this very clear. For cast #12 two records in the ADD file duplicate other lines, but with OXY values of 0. It is odd, but easily fixed. The bad data lines were removed. In that file and in a few others, some lines were out of order; they were re-ordered according to sample number.

File 2004-26chlarc.xls was obtained from the analyst. Non-standard channel names were corrected. The corrected file was saved as 2004-26chlarc.csv. This was converted to CHL files. The general comment from the header of the spreadsheet was added to the general comments to be added at the end of processing, 2004-26-bottle-header.txt.

File 2004-26nuts.xls was obtained from the analyst. The file was saved as a csv file and then converted to individual NUTS files. These were sorted on sample number and called NUT1.

The sample numbers were added to the BOT files to create SAM files, which were averaged to create SAMAVG files. They were then merged with the SAL, ADD, CHL and NUT1 files in 4 steps. (Output: MRG1, MRG2, MRG3 and MRG)

##### 13. COMPARE

Salinity

COMPARE was run. Because there was a post-cruise calibration and because the pairs of sensors were in good agreement when the post-cruise calibrations were used, we expect that the CTD and bottles should agree quite well.

However, there is a lot of scatter and many outliers, some severe. When pressures less than 150db and differences greater than 0.01psu are excluded the primary salinity was found to be high by 0.0017psu. When Niskin #1 was excluded as well, the primary salinity was found to be high by 0.0007psu. Looking at only the bottom bottles it is obvious that many compare badly with the CTD. Investigations show that many of the poor comparisons are in turbid waters very close to the bottom. It is not known if the problem is with the bottles or the CTD, but in either case the samples are not useful for calibration. One of the bottles about 25db above bottom also compared badly, and close examination suggests a turbid shed wake occurred shortly before firing.

For the secondary salinity if we exclude bottom bottles, bottles from above 150db and differences >0.025psu, the CTD was found to be low by 0.0004psu. (See 2004-26-sal-comp1.xls.)

There were 9 outliers excluded from the comparison. These were investigated to see if flags should be assigned to the bottle values. Four were in areas of large salinity gradients, so minor flushing problems can explain the differences. Noisy CTD data accounted for another outlier. The following samples were studied further:

- For cast #15, sample #61 was a severe outlier in COMPARE . The value of 27.94 at 151db is completely unbelievable. It was flagged “d”.

- For cast #20, sample #73 at 351db was an outlier in COMPARE, differing from the CTD by 0.09psu. It was noted in the log book that the CTD hit bottom so there may be mud in the cell. The salinity was very noisy at the bottom, but had pretty much settled down by the time Niskin #1 was fired, and the comparison there looks ok. Just before the 2nd bottle was fired a shed wake arrived with low transmissivity water. The transmissivity was very noisy at the time of firing of Niskin #2. The presence of mud could have affected either the CTD and/or the bottle. The sample was flagged “c”.

- For cast #20, sample at 1.9m was an outlier, but given that it was very near the surface it was not flagged as it could well be more reliable than the CTD value.

- The surface sample from cast #36 was flagged “c” by the analyst. No change was made to this. The point is an outlier, but it is from near the surface.

- For cast #46, sample #179 at 188db was an outlier in COMPARE, differing from the CTD by 0.12db. The CTD data was noisy at that level, but the range of salinity values during the bottle stop was only about 0.02psu, so could explain only 10% of the difference between bottle and CTD. The sample was flagged “c”.

- For cast #58, the bottle at 111db is out of line with the bottle above it and is an outlier in COMPARE. Flagged “c”.

Fluorescence versus titrated Chlorophyll

COMPARE was run and the CSV file used to plot chlorophyll versus fluorescence. There was a lot of scatter, but nothing that looked like an outlier. The maximum chlorophyll was about 3.3ug/l. When plotted together and the dark value removed from the fluorescence, the slope was about 0.8, with chlorophyll about 20% higher than FL, but when FL is zero, CHL is about 0.2 which does not seem reasonable. When plotted in separate groups for FL < or > 1, the slope for each group was close 1 with an offset of 0.12 for low fluorescence and -0.3 for high fluorescence. If all data are plotted together and the offset forced to be 0, then the slope is 0.96.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings referred to 1 upcast record ~5db in cast #26.

All the DEL files were copied to *.EDT.

##### DETAILED EDITING

The primary sensors will be chosen because the salinity is less noisy than that from the secondary.

Page plots were produced. These were used to guide the editing.

The average descent rate was quite low during this cruise, typically 0.5 to 0.7m/s. The descent rate was generally steady. However, near the mouth of Juan de Fuca Strait the descent rate was extremely noisy with frequent reversals, leading to considerable corruption by shed wakes. There were small spikes in salinity, but it was not a major problem.

The following casts required heavy editing: 64-69

The following casts required no editing: 22, 41, 51,

All other casts required light editing.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – Sensor #1763 has been used frequently and was found to be low by about 0.0035psu in late August. There was a post-cruise calibration that indicated it was low by about 0.004 in October. Sensor #1766 had not been used since the previous calibration.

Historic ranges – Most of the casts in the Strait of Georgia had temperatures above the historic maxima; this is believed to be real and continued to be true in December 2004 when the area was revisited.

##### Recalibration

No recalibration was applied.

##### Special Fluorometer Processing

The EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD.

The SAM files were put through CLEAN, REMOVE, HEADEDIT and saved as BOF.

The FCTD and BOF files were saved to a CD-ROM for Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the full EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

(Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared.

Separate profile plots were prepared of the top 100db with Transmissivity, PAR, Temperature and Fluorescence versus pressure.

##### FINAL CTD files steps

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and units and to add the following comment using file 2004-26hdr.txt:

Transmissivity: The data are unedited except where records were

removed in editing temperature and salinity.

Fluorescence:URU:Seapoint - The data are unedited except as above.

The Standards Check routine was run and HEADER EDIT adjusted and rerun until no further problems were found.

The final files were named CTD.

A cross-reference listing was produced.

The sensor history was updated.

##### Final Bottle Files

The MRG files were put through CLEAN to remove the SeaBird.

SORT was used to put the data in order of increasing pressure.

REMOVE was used to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate  and Flag.

HEADER EDIT was run to fix channel names and formats and to add a standard comment including an explanation of the quality flags. The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

Particulars

2. Hit bottom

7. Error in sample numbers in oxygen titration. Sample named #29 in ADD file should be #28. Sample named #928 should be #28.

20. Hit bottom; small kink in cable

21. The sample numbers written beside this event should be beside event #20. This was not a CTD cast.

25. CTD cable re-terminated

27. Problems with CTD cable

35. Station name may be wrong in DAT file.

36. Extra label was made. Samples 150 and 151 are both at 300m.

46. Strong current, rain

58. Sample #227, Niskin #1 – Note says “1st BOT on File “cast 55”. DO value is on rosette sheet,

but not in ADD file. Value was added to MRG file.

65. Station name wrong in file.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-26

Dates: Start: September 13, 2004 End: September 17, 2004

Location: JdeF/SoG Vessel: Vector

Party Chief: Masson D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585               Cruise ID#:		2004-26

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 06/05/04 | Factory | 18/11/04 | Factory |
| Conductivity | 1763 | 19/12/03 | “ | 19/11/04 | “ |
| Secondary Temp. | 2038 | 17/05/03 | “ | 16\10\04 | “ |
| Secondary Cond. | 1766 | 15/05/03 | “ | 15\10\04 | “ |
| Transmissometer | 723DR | 24/10/03 | IOS |  |  |
| Fluorometer | 2229 |  |  |  |  |
| PAR | 4656 | 11/02/03 |  |  |  |
| Pressure Sensor | 77511 | 30/09/00 | Factory |  |  |