## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 22 Dec. 2004 | Recalibrated CTD salinity based on post-cruise calibration; see note below. |

## PROCESSING NOTES

Cruise: 2004-28

Agency: OSAP

Location: Hecate Strait / Queen Charlotte Sound / WCVI

Project: Northern Herring

Party Chief: Therriault T.

Platform: W. E. RICKER

Date: August 3, 2004 – August 15, 2004

Processed by: Germaine Gatien

Date of Processing: 16 September 2004 – 23 September 2004

Number of original CTD casts: 34

Number of casts processed: 34

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit was a SeaBird model 11 (S/N not recorded.)

# SUMMARY OF QUALITY AND CONCERNS

- This cruise was a continuation of cruise 2004-19. The same methods were used for this cruise.

- The CTD Daily Log Book contained no information about the equipment or scientific crew.

- There were many errors in the times and dates in the log book; many seem to be due to errors in calculating UTC time because of failure to express afternoon times in 24-hour clock format.

- The transmissivity data has very different upcast and downcast sections for many casts. The downcasts look reasonable, but should be used with caution.

- Fluorescence data looks poor near the surface in the upcasts. The downcasts seem ok.

- Cast #87 had pressure spikes and jumps; data below 55db was removed since the pressures may be wrong.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. There were no salinity samples.

The configuration file from 2004-19 was copied and renamed 2004-28ctd.con.

##### Conversion of Raw Data

The raw data were converted using configuration file 2004-28CTD.con.

The first cast, #43, was corrupted, with no useful data in it after the initial conversion. There was a note in the log that there might be an error in this cast. The DAT file was put through a routine that finds and repairs gaps and the output was named 2004-28-9943.DAT; conversion then produced reasonable data. The converted file was renamed 2004-28-0043.CNV.  All expected channels were present.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5	Points per block = 50

##### ALIGNCTD

ALIGNCTD was used to advance the secondary conductivity by 0.073s since it is likely that this deck unit was one of the older models. Fine-tuning of the alignment will be done later using SHIFT.

##### CELLTM

The results of 2004-19 were used and this step was skipped.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The differences between the pairs of channels looks similar to the results for 2004-19, on average, but there are sections that look very odd. For example in cast #87 the differences between the salinity channels is about 0.001; at about 130db of the downcast the differences suddenly became much lower, then gradually moved back to the 0.001 value. On the upcast the differences are much larger.

The transmissivity sometimes looks ok, but for other casts there are hugh differences between upcast and downcast traces. Often the shape is more or less the same, but with a varying and often large offset.

The fluorescence looks like that of 2004-19 with the upcast values sometimes significantly lower than the downcast near the surface.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

Cast #43 could not be converted due to the absence of positions and time in the SeaBird headers. These were added to the CNV file using information from the log and then conversion was possible.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header summary and a header check were produced. The following problems were found:

- The header for cast #57 had scrambled times/positions; these were fixed in the CNV file and it was reconverted and CLEAN rerun.

- There are no station names in the files or log book.

- The times in the headers frequently differ from times in the Daily Log Book by 12 hours. It is assumed the error is in the log book.

- The dates are wrong for some casts in the log book. It is clear that the headers are correct.

- The header and log book times for cast #43 differed by 32 minutes. Since the headers were incomplete for that cast it will be assumed that the log book is correct. The time was corrected in the headers.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 1.5db, with a minimum of 0.9db; as found for 2004-19 these values are low for the Ricker. An examination of one data file with a minimum of 0.9db shows that the CTD really was close to the surface with very low conductivity values. The pumps were not turned on until the pressure was about 3db. The unusual surface values are clearly due to the method used to deploy the CTD which differs from that used by Hugh Maclean when he samples from the Ricker. The pressure calibration is not a problem.

T0, T1, S0 and S1 were plotted for all casts. The secondary salinity looks poor, especially near the surface of the upcast when the values fall well below the historic. The upcast secondary temperature is often very different from the primary and both downcast traces. The fluorescence looks reasonable on the downcast, but for some casts the upcast values in the top 30m look much too low and are sometimes extremely spiky. The transmissivity frequently looks very different between upcast and downcast, but the downcasts are probably ok.

##### SHIFT

Fluorescence

The results of cruise 2004-19 were used. Cast #71 was checked before and after to ensure the alignment was reasonable. A shift of +24 records has been used for most cruises in the past and seems reasonable for this one, so it was applied to all casts.

Conductivity

The results of cruise 2004-19 were chosen for this cast. Cast #71 was checked and spikiness was reduced

All data were put through SHIFT using -1.2 for the primary. No shift was applied to the secondary data since it is unlikely to be archived.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Salinity: 5

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maxiumum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The only warning referred to cast #109. The pressure jumps from 61.96db to 72.04db from one scan to the next. The other channels change very slightly as expected, so this does not appear to be a matter of missing scans. There was a stop around 50db and then erratic pressure. The data above 50db looks ok, so removing the bottom of this cast seems the best approach. There is a note in the log book about a problem occurring at about 55m. An attempt was made to repair the file, but no gaps were found. A carriage return line feed was found and removed, but this made the data worse.

##### DETAILED EDITING

The secondary sensors did not perform well. The primary sensors have some noise but are reasonable.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case.

Something was obviously wrong with cast #57. Returning to the original files it was found that the pressure started at about -18db, with reasonable looking temperature and salinity. An attempt was made to repair the DAT file, and a gap was found and removed. Steps 3 through 13 were repeated and the data looks fine.

For Cast #109 the data was removed from 49db to the bottom because of the problems in the pressure channel. The temperature and salinity data looks good, but it seems doubtful that it is always from the pressure given. The skips in pressure do not correspond to comparable jumps in other variables.

The descent rate was generally fairly quiet and high minimizing the corruption by shed wakes, so only light editing was required. All casts required some editing.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that do not require editing.

##### Other comparisons

Previous experience with these sensors – This is discussed in detail in the report for 2004-19. The history of the primary sensors shows a gradual shift to lower values with the most recent result being that it is probably low by about -0.001psu. However, there is sparse sampling and a lot of scatter in this history.

Historic ranges –There were many excursions from the historic ranges in near-surface temperature in Queen Charlotte Sound. Similar observations have been made by a number of cruises in 2004. There is no indication of instrumental error. There were two casts in Dixon Entrance with salinity low near the surface in one case and from 20-35m in another. Similar observations were made during 2004-17 a month earlier, though the low salinity was seen a little further to the north.

Repeatability – There were no repeat casts, but casts #47 and #109 were quite close in space as were #105 and #107. The former are 11 days apart and look quite different with #109 being much warmer at the surface. The latter two casts are reasonably similar above and below the thermocline, but they are not close enough in space to tell us anything about the CTD performance.

##### Recalibration

The primary salinity is believed to be slightly lower than the bottles, probably within 0.001psu. No recalibration was applied.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the bin-averaged, edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and a header entry and to add the following comments:

Transmissivity and Fluorescence – The data are nominal and unedited, except that some records were removed in editing T and S.

The transmissometer data showed unusually large differences between upcast and downcasts during this cruise, but the downcasts generally seemed reasonable. The data should be used with caution

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-28

Dates: Start: July 20, 2004 End: August 2, 2004

Location: West Coast Vancouver Island

Vessel: W.E. Ricker

Party Chief: Therriault T.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2004-28

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory |  |  |
| Conductivity | 2424 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2374 | 20/06/02 | “ |  |  |
| Secondary Cond. | 2399 | 16/04/02 | “ |  |  |
| Transmissometer | 197 | 16/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |

Dec. 22, 2004: All CTD files were recalibrated using file 2004-28-recal2.ccf to add 0.0138psu to the primary salinity based on post-cruise calibration of the conductivity sensor. It was assumed that the drift was linear with time. G. Gatien