## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2004-30

Agency: OSAP

Location: Strait of Georgia

Project: COPRA / Desolation Sound / Euphausiid Acoustics

Party Chief: Romaine S.

Platform: Vector

Date: October 18, 2004 – October 25, 2004

Processed by: Germaine Gatien

Date of Processing: 7 February 2004 – 21 February 2005

Number of original CTD casts: 	2

Number of casts processed: 2

Number of rosette casts: 2

Number of rosette casts processed: 1 (no sampling done for 1 cast)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea transmissometer (#498DR), and Altimeter (#1024); a Dissolved Oxygen Sensor (#0047) wa4s on the primary pump and a Seapoint Fluorometer (#2356) was on the secondary pump with a 10X cable. The deck unit was a model 911 (#0508).

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order except that the second event was crossed out. There were problems with that cast, but there is data.

The pressure offset used at sea was -0.7db whereas at least +1db is appropriate.

There is insufficient information to give an estimate of the errors in data from the dissolved oxygen sensor. This data was recalibrated based on only 4 bottles.

The transmissivity, dissolved oxygen and fluorescence are unedited, except where records were removed in the editing of temperature and salinity.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book and rosette sheets were obtained. There is a salinity analysis sheet and dissolved oxygen values are entered on the rosette log, but there are no files with any of this data.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The date of the pressure calibration was wrong and the pressure offset of -0.7db seems unlikely to be correct. Recent uses of this pressure sensor suggest that it is reading low by from 0db to 1.5db. The offset was changed to +1db based on the most recent use during 2004-29; the resulting file was named 2004-30-CTD.con.

##### Conversion of Raw Data

The raw files were converted using the configuration file 2004-30-ctd.con.

Rosette files were converted using a start time of -2s and duration of 5s.

The converted files were plotted and the temperatures are close during the downcast but not during the upcast for the first cast. The conductivity channels are quite far apart in the high gradient region even during the downcast. The second cast has huge spikes in the secondary temperature and conductivity. The other channels seem ok.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only. Parameters used were: 	Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### CELLTM

CELLTM was run using (0.02, 9) on the primary channels and (0.0245, 9.5) on the secondary based on the results of 2004-29.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity.

##### Test Plots and Channel Check

The two casts are too shallow for these checks to be useful.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number and to fix the pad numbers in the dissolved oxygen channel. A station name was added to the first cast based on a log entry.

The secondary salinity is bad for cast #2. Since the fluorescence looks ok, this is probably not a pump problem. The transmissivity goes very low at the bottom of cast #2.

The rosette files were converted to IOS files and the station name added to cast #1. The extension was changed to BOT.

The BOT files were plotted and no editing was found necessary.

##### Checking Headers

A header summary and header check were done and no errors found. The cruise track was plotted.

The average surface pressure is 2.0db which is reasonable.

The altimeter entries in the headers were checked and match the original data.

##### BOTTLE FILE PREPARATION

There were 2 salinity bottle samples and the results were on an analysis sheet, but not in a CSV file. A file was prepared. A flag “c” was applied to 1 sample that required many readings to reach stable results.

There were 5 titrated DO values recorded on the rosette sheet, but once again no spreadsheet. A spreadsheet was prepared. As there are no comments on the rosette sheet, no flags were entered.

The two spreadsheets were then converted to SAL and DOX files.

The BOT files were averaged to enable an addsamp file to be created. The addsamp file was edited to add sample numbers taken from the rosette sheets. The sample numbers were added to the BOT files to create SAM files, which were averaged to create SAMAVG files. The SAMAVG files were then merged with the SAL and DOX files in two steps.  (Output: MRG1 and MRG)

##### 11. COMPARE

Salinity

There were only 2 bottle salinity samples available, both from cast #1 and in shallow water. The differences from the CTD are large for both sensors. During 2004-08 the CTD was found to be low by about 0.0006 and during 2004-29 by 0.003psu. But neither was considered to be a good comparison due to there being a lot of scatter and problems with the CTD. The post-cruise calibration shows a drift upwards of only 0.00015C º in the primary temperature since the previous calibration in February 2005, and no drift in the primary conductivity. This would lead to no significant drift in salinity. No recalibration will be applied. (See 2004-30-comp.xls.)

Dissolved Oxygen

During 2004-24 when this sensor was last used, the following fit was found:

DOX(Titrated) = 1.078 * DOX(CTD) + 0.023

Plotting the data from 2004-30 we find a fit of

DOX(Titrated) = 0.9872 * DOX(CTD) + 0.8822

That is very different from all other results and based on only 5 shallow samples. We expect some differences since the range of DOX and pressure values in the fits vary, and in this case the ranges of both are quite narrow. When the sensor went back to SeaBird in February 2005 they found a hole in the membrane and were unable to provide a calibration. It is possible that the damage occurred before this cruise which might explain the big difference from 2004-24. However, it seems likely that if the damage had occurred before these observations, the values would be much further from bottles than they are.

Dario Stucchi considers that the bottle sample from the bottom of the cast should be flagged. He does not expect a rise in dissolved oxygen at this point and the CTD shows no sign of DO increasing. The sample was flagged “c”. The fit based on just 4 bottles is:

DOX(Titrated) = 1.065 * DOX(CTD) +0.516

The slope is similar to that from other cruises, although the offset is larger than usual.

This equation will be applied to the DO sensor data. A note will be added to the headers indicating that there is no way of estimating the error in the DO data.

##### 12. SHIFT

Since there was only one full cast (#2 has not upcast) the results of other cruises were applied and the results checked to ensure they were reasonable.

Conductivity

During 2004-11 and 2004-08 when the same CTD was used the primary conductivity was shifted by -0.2 records. During 2004-29 when there was no DO sensor -0.2 records was also found to be the best setting.

All casts were put through SHIFT to apply the -0.2 record shift to the primary conductivity only. There is no significant change, with no sign of misalignment before or after.

Fluorescence

The fluorescence channel was advanced by 24 records for all casts since this value is almost always the best choice. It worked well on cast #1.

Dissolved Oxygen

The last time this sensor was used was in August during 2004-24 when a shift of +120 records was applied. The same shift was applied and worked reasonably well on cast #1.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing. Page plots were produced using (T0,S0). These were used to guide the editing. Both casts needed records removed from the top and near or at the bottom; no further editing was required.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The sensors used in the primary position have been used for many cruises since their previous calibration, including most recently, 2004-08, -11 and -29. (They were in the secondary position for most of the cruises.) They were not selected for archiving except for 2004-29 when no recalibration was applied. The secondary sensors have been used twice since the last calibration, for 2004-29 and 2004-37. In neither case were they selected for archiving.

Historic ranges – There is no climatology available for these casts.

Repeat casts – The two casts are very close together. The differences at the bottom along gamma lines are on the order of 0.02Cº and 0.003psu. So the repeatability of the instrument looks reasonable.

##### Initial Recalibration

No recalibration will be applied to the salinity.

The oxygen was recalibrated using the equation 2004-30-recal.ccf:

DOX(Titrated) = 1.065 * DOX(CTD) +0.516

A check was made after recalibration to ensure it was applied properly and it was.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

(Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared using the edited data and displaying Temperature and profiles. Separate profile plots were produced with temperature, transmissivity, fluorescence and PAR.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and units and to add the following comment using file 2004-30-hdr.txt:

******************************************************************

There is no way of estimating the error in the CTD Dissolved

Oxygen data because there are only a few bottle samples

available. Recalibration was based on 5 bottles from 1 cast.

The DO sensor was found to have a hole in the membrane in February 2005.

******************************************************************

The dissolved oxygen, transmissivity, PAR and fluorescence data are unedited except where records were removed in editing temperature and salinity.

The Standards Check routine was run and HEADER EDIT adjusted and rerun until no further problems were found.

The final files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag.

HEADER EDIT was run to fix channel names and formats and to add a standard comment about quality flags. The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE. Since there are no samples for cast #2, that file was deleted.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

2. CTD system failed at bottom of cast.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-30

Dates: Start: September 13, 2004 End: September 23, 2004

Location: WCVI

Vessel: Vector

Party Chief: Romaine S.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550   Cruise ID#:	2004-30

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 22/04/03 | Factory | 01/02/04 | factory |
| Conductivity | 1729 | 24/04/03 | “ | 03/02/05 | Factory |
| Secondary Temp. | 2038 | 19/08/04 | “ |  |  |
| Secondary Cond. | 2173 | 19/08/04 | “ |  |  |
| Transmissometer | 498DR | 05/08/03 | IOS |  |  |
| Dissolved Oxygen | 0047 | 13 Jan 04 | factory |  |  |
| PAR | 4615 | 15 Dec 2000 | ? |  |  |
| Altimeter | 1024 | 12-May-02 |  |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 75636 | 06/04/99 | Factory |  |  |