## File List
Number of Files= |  | 20
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2004-37-0001 | 2004-37-0001 | 11 | 14 | 2004-37-0001 & 2004-37-0001
2 | 2004-37-0002 | 2004-37-0002 | 16 | 29 | 2004-37-0002 & 2004-37-0002
3 | 2004-37-0006 | 2004-37-0006 | 31 | 45 | 2004-37-0006 & 2004-37-0006
4 | 2004-37-0009 | 2004-37-0009 | 47 | 60 | 2004-37-0009 & 2004-37-0009
5 | 2004-37-0012 | 2004-37-0012 | 62 | 74 | 2004-37-0012 & 2004-37-0012
6 | 2004-37-0015 | 2004-37-0015 | 76 | 87 | 2004-37-0015 & 2004-37-0015
7 | 2004-37-0018 | 2004-37-0018 | 89 | 97 | 2004-37-0018 & 2004-37-0018
8 | 2004-37-0023 | 2004-37-0023 | 99 | 112 | 2004-37-0023 & 2004-37-0023
9 | 2004-37-0032 | 2004-37-0032 | 114 | 125 | 2004-37-0032 & 2004-37-0032
10 | 2004-37-0034 | 2004-37-0034 | 127 | 142 | 2004-37-0034 & 2004-37-0034
11 | 2004-37-0039 | 2004-37-0039 | 144 | 159 | 2004-37-0039 & 2004-37-0039
12 | 2004-37-0042 | 2004-37-0042 | 161 | 176 | 2004-37-0042 & 2004-37-0042
13 | 2004-37-0044 | 2004-37-0044 | 178 | 192 | 2004-37-0044 & 2004-37-0044
14 | 2004-37-0047 | 2004-37-0047 | 194 | 209 | 2004-37-0047 & 2004-37-0047
15 | 2004-37-0049 | 2004-37-0049 | 211 | 223 | 2004-37-0049 & 2004-37-0049
16 | 2004-37-0052 | 2004-37-0052 | 225 | 236 | 2004-37-0052 & 2004-37-0052
17 | 2004-37-0055 | 2004-37-0055 | 238 | 254 | 2004-37-0055 & 2004-37-0055
18 | 2004-37-0057 | 2004-37-0057 | 256 | 271 | 2004-37-0057 & 2004-37-0057
19 | 2004-37-0059 | 2004-37-0059 | 273 | 283 | 2004-37-0059 & 2004-37-0059
20 | 2004-37-0061 | 2004-37-0061 | 285 | 298 | 2004-37-0061 & 2004-37-0061

## (DOX-CTD - DOX-BOT)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.46413 |  | X | CALC Y
m= | 0.0026539 | 0.00017517 |  | Std Dev of Resid |  | 0.25469 |  | 1.5612 | -1.3622
b= | -1.3663 | 0.02262 |  | Resid Sum of Squares |  | 17.1901 |  | 372.67 | -0.37732
# of Values Extracted = |  |  | 269
# of Values Used in Fit = |  |  | 267
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 110.209 | 0.14646 | 0.0023558 | -0.59954 | -1.0739 | -0.47431 |  | 100.191 | -0.053025
1 | YES | 100.191 | 0.16698 | 0.00027273 | -0.053025 | -1.1004 | -1.0474 |  | 75.0704 | -0.58691
1 | YES | 75.0704 | 2.6501 | 0.0053135 | -0.58691 | -1.1671 | -0.5802 |  | 1.5835 | -1.2407
1 | YES | 1.5835 | 3.6213 | 0.073599 | -1.2407 | -1.3621 | -0.12139 |  | 231.111 | -0.95754
2 | YES | 231.111 | 2.5645 | 0.0012848 | -0.95754 | -0.753 | 0.20454 |  | 200.09 | -1.0462
2 | YES | 200.09 | 2.6258 | 0.0028774 | -1.0462 | -0.83532 | 0.21091 |  | 175.687 | -1.1205
2 | YES | 175.687 | 2.6535 | 0.001119 | -1.1205 | -0.90008 | 0.22042 |  | 150.335 | -1.3028
2 | YES | 150.335 | 2.8042 | 0.0040114 | -1.3028 | -0.96736 | 0.33542 |  | 125.097 | -1.1616
2 | YES | 125.097 | 2.8934 | 0.0035549 | -1.1616 | -1.0343 | 0.12728 |  | 99.9256 | -1.2356
2 | YES | 99.9256 | 3.0364 | 0.0029657 | -1.2356 | -1.1011 | 0.13451 |  | 74.4687 | -1.2787
2 | YES | 74.4687 | 3.3423 | 0.0074156 | -1.2787 | -1.1687 | 0.11001 |  | 49.1603 | -1.3215
2 | YES | 49.1603 | 3.4055 | 0.0020856 | -1.3215 | -1.2359 | 0.085582 |  | 40.6798 | -1.3115
2 | YES | 40.6798 | 3.4705 | 0.0019233 | -1.3115 | -1.2584 | 0.053176 |  | 29.7398 | -1.2646
2 | YES | 29.7398 | 3.5254 | 0.0016215 | -1.2646 | -1.2874 | -0.022827 |  | 19.6317 | -1.3259
2 | YES | 19.6317 | 3.5421 | 0.00039364 | -1.3259 | -1.3142 | 0.011707 |  | 9.6901 | -1.2771
2 | YES | 9.6901 | 3.5499 | 0.0014419 | -1.2771 | -1.3406 | -0.063546 |  | 4.9828 | -1.2555
2 | YES | 4.9828 | 3.5604 | 0.001072 | -1.2555 | -1.3531 | -0.097559 |  | 1.6667 | -1.3018
2 | YES | 1.6667 | 3.5762 | 0.0024279 | -1.3018 | -1.3619 | -0.060079 |  | 244.526 | -0.58065
3 | YES | 244.526 | 1.4684 | 0.001215 | -0.58065 | -0.7174 | -0.13675 |  | 244.611 | -0.81563
3 | YES | 244.611 | 1.4684 | 0.0012711 | -0.81563 | -0.71717 | 0.098455 |  | 203.398 | -0.60664
3 | YES | 203.398 | 1.4764 | 0.0019011 | -0.60664 | -0.82654 | -0.2199 |  | 175.87 | -0.68709
3 | YES | 175.87 | 1.6429 | 0.0012649 | -0.68709 | -0.8996 | -0.21251 |  | 149.282 | -1.0215
3 | YES | 149.282 | 1.7275 | 0.0018305 | -1.0215 | -0.97016 | 0.051381 |  | 125.206 | -0.81862
3 | YES | 125.206 | 2.1854 | 0.0023953 | -0.81862 | -1.0341 | -0.21543 |  | 99.6255 | -0.98561
3 | YES | 99.6255 | 2.3844 | 0.0015021 | -0.98561 | -1.1019 | -0.11633 |  | 75.1051 | -1.0281
3 | YES | 75.1051 | 2.5879 | 0.0012538 | -1.0281 | -1.167 | -0.13887 |  | 50.4953 | -1.2957
3 | YES | 50.4953 | 3.2923 | 0.0013106 | -1.2957 | -1.2323 | 0.063335 |  | 41.1572 | -1.5615
3 | YES | 41.1572 | 3.7275 | 0.0018261 | -1.5615 | -1.2571 | 0.30442 |  | 29.3588 | -1.709
3 | YES | 29.3588 | 4.41 | 0.0036958 | -1.709 | -1.2884 | 0.42059 |  | 20.5425 | -1.6802
3 | YES | 20.5425 | 4.5408 | 0.0016749 | -1.6802 | -1.3118 | 0.36837 |  | 10.4123 | -1.6583
3 | YES | 10.4123 | 4.5747 | 0.0010919 | -1.6583 | -1.3387 | 0.31956 |  | 5.0724 | -1.6551
3 | YES | 5.0724 | 4.5799 | 0.0015894 | -1.6551 | -1.3529 | 0.30221 |  | 1.8669 | -1.7904
3 | YES | 1.8669 | 4.5846 | 0.0023584 | -1.7904 | -1.3614 | 0.42903 |  | 221.623 | -0.54429
4 | YES | 221.623 | 1.4957 | 0.001313 | -0.54429 | -0.77818 | -0.23389 |  | 201.791 | -0.55505
4 | YES | 201.791 | 1.4889 | 0.00038409 | -0.55505 | -0.83081 | -0.27576 |  | 172.797 | -0.65709
4 | YES | 172.797 | 1.5309 | 0.0023629 | -0.65709 | -0.90775 | -0.25066 |  | 150.467 | -0.72293
4 | YES | 150.467 | 1.7571 | 0.0038257 | -0.72293 | -0.96701 | -0.24409 |  | 124.65 | -1.0374
4 | YES | 124.65 | 1.9126 | 0.0019911 | -1.0374 | -1.0355 | 0.0018307 |  | 99.9203 | -0.90424
4 | YES | 99.9203 | 2.1878 | 0.024804 | -0.90424 | -1.1012 | -0.19692 |  | 75.1964 | -1.091
4 | YES | 75.1964 | 2.68 | 0.01257 | -1.091 | -1.1668 | -0.075772 |  | 50.3963 | -1.1775
4 | YES | 50.3963 | 2.9895 | 0.005905 | -1.1775 | -1.2326 | -0.055108 |  | 40.4829 | -1.1277
4 | YES | 40.4829 | 2.9873 | 0.00256 | -1.1277 | -1.2589 | -0.13118 |  | 29.9081 | -1.1636
4 | YES | 29.9081 | 3.2054 | 0.021531 | -1.1636 | -1.287 | -0.12338 |  | 20.177 | -1.3863
4 | YES | 20.177 | 3.6917 | 0.016997 | -1.3863 | -1.3128 | 0.073474 |  | 9.7949 | -1.8326
4 | YES | 9.7949 | 4.1154 | 0.032814 | -1.8326 | -1.3403 | 0.49225 |  | 5.2267 | -1.6496
4 | YES | 5.2267 | 4.4224 | 0.010058 | -1.6496 | -1.3525 | 0.29718 |  | 2.1258 | -1.3901
4 | YES | 2.1258 | 4.4949 | 0.00088504 | -1.3901 | -1.3607 | 0.029399 |  | 184.313 | -0.64399
5 | YES | 184.313 | 1.551 | 0.0014171 | -0.64399 | -0.87719 | -0.2332 |  | 175.011 | -0.51441
5 | YES | 175.011 | 1.5436 | 0.0014294 | -0.51441 | -0.90188 | -0.38747 |  | 149.623 | -0.50614
5 | YES | 149.623 | 1.5529 | 0.002364 | -0.50614 | -0.96925 | -0.46311 |  | 124.73 | -0.59293
5 | YES | 124.73 | 1.6841 | 0.0020007 | -0.59293 | -1.0353 | -0.44239 |  | 99.845 | -0.67674
5 | YES | 99.845 | 1.9403 | 0.0018642 | -0.67674 | -1.1014 | -0.42461 |  | 74.951 | -0.91923
5 | YES | 74.951 | 2.3818 | 0.0014246 | -0.91923 | -1.1674 | -0.24819 |  | 50.2894 | -0.93303
5 | YES | 50.2894 | 2.845 | 0.0020933 | -0.93303 | -1.2329 | -0.29984 |  | 40.0385 | -1.1514
5 | YES | 40.0385 | 3.1446 | 0.0018432 | -1.1514 | -1.2601 | -0.10865 |  | 30.0153 | -1.2485
5 | YES | 30.0153 | 3.6185 | 0.0021177 | -1.2485 | -1.2867 | -0.038206 |  | 19.4498 | -1.3606
5 | YES | 19.4498 | 4.3044 | 0.0035562 | -1.3606 | -1.3147 | 0.045855 |  | 9.701 | -1.7584
5 | YES | 9.701 | 4.3526 | 0.002138 | -1.7584 | -1.3406 | 0.41779 |  | 4.9995 | -1.6149
5 | YES | 4.9995 | 4.3721 | 0.0018055 | -1.6149 | -1.3531 | 0.26182 |  | 1.5612 | -1.6004
5 | YES | 1.5612 | 4.3726 | 0.0019213 | -1.6004 | -1.3622 | 0.23821 |  | 171.173 | -0.56283
6 | YES | 171.173 | 1.6122 | 0.001202 | -0.56283 | -0.91206 | -0.34924 |  | 149.783 | -0.57172
6 | YES | 149.783 | 1.6273 | 0.0015823 | -0.57172 | -0.96883 | -0.39711 |  | 124.912 | -0.59684
6 | YES | 124.912 | 1.6392 | 0.0018545 | -0.59684 | -1.0348 | -0.438 |  | 98.8446 | -0.8592
6 | YES | 98.8446 | 1.8678 | 0.0042556 | -0.8592 | -1.104 | -0.24482 |  | 74.2203 | -0.8727
6 | YES | 74.2203 | 2.0953 | 0.0018423 | -0.8727 | -1.1694 | -0.29666 |  | 49.47 | -0.92179
6 | YES | 49.47 | 2.4122 | 0.0020652 | -0.92179 | -1.235 | -0.31325 |  | 39.8415 | -1.0238
6 | YES | 39.8415 | 2.5562 | 0.0072749 | -1.0238 | -1.2606 | -0.23682 |  | 32.8962 | -1.3838
6 | YES | 32.8962 | 2.9242 | 0.0096326 | -1.3838 | -1.279 | 0.10476 |  | 22.9881 | -1.2837
6 | YES | 22.9881 | 3.1593 | 0.0019928 | -1.2837 | -1.3053 | -0.021575 |  | 13.0632 | -1.3501
6 | YES | 13.0632 | 3.2449 | 0.0019118 | -1.3501 | -1.3317 | 0.018395 |  | 4.2473 | -1.4829
6 | YES | 4.2473 | 3.3821 | 0.0014905 | -1.4829 | -1.3551 | 0.1278 |  | 1.983 | -1.2927
6 | YES | 1.983 | 3.3723 | 0.0024434 | -1.2927 | -1.3611 | -0.06839 |  | 77.2534 | -0.862
7 | YES | 77.2534 | 2.323 | 0 | -0.862 | -1.1613 | -0.29931 |  | 75.1932 | -0.81963
7 | YES | 75.1932 | 2.4244 | 0.01458 | -0.81963 | -1.1668 | -0.34715 |  | 49.5689 | -1.0201
7 | YES | 49.5689 | 2.7389 | 0.016676 | -1.0201 | -1.2348 | -0.21466 |  | 39.9511 | -1.1619
7 | YES | 39.9511 | 2.9771 | 0.0073387 | -1.1619 | -1.2603 | -0.098428 |  | 30.3125 | -1.2746
7 | YES | 30.3125 | 3.2204 | 0.017383 | -1.2746 | -1.2859 | -0.011257 |  | 20.0896 | -1.2694
7 | YES | 20.0896 | 3.4166 | 0.0022552 | -1.2694 | -1.313 | -0.043598 |  | 10.9783 | -1.247
7 | YES | 10.9783 | 3.4661 | 0.00075419 | -1.247 | -1.3372 | -0.090248 |  | 2.8781 | -1.3381
7 | YES | 2.8781 | 3.5019 | 0.0094609 | -1.3381 | -1.3587 | -0.020614 |  | 2.8781 | -1.2791
7 | YES | 2.8781 | 3.5019 | 0.0094609 | -1.2791 | -1.3587 | -0.079614 |  | 210.266 | -1.1024
8 | YES | 210.266 | 2.7516 | 0.0015037 | -1.1024 | -0.80832 | 0.29409 |  | 200.335 | -1.0105
8 | YES | 200.335 | 2.7605 | 0.0011837 | -1.0105 | -0.83467 | 0.17586 |  | 174.961 | -0.98719
8 | YES | 174.961 | 2.7558 | 0.0019929 | -0.98719 | -0.90201 | 0.085179 |  | 150.142 | -1.0285
8 | YES | 150.142 | 2.8096 | 0.001072 | -1.0285 | -0.96788 | 0.060573 |  | 124.952 | -1.1309
8 | YES | 124.952 | 2.9861 | 0.00079902 | -1.1309 | -1.0347 | 0.096132 |  | 99.7344 | -1.1517
8 | YES | 99.7344 | 3.1033 | 0.0014975 | -1.1517 | -1.1017 | 0.050008 |  | 75.0419 | -1.2314
8 | YES | 75.0419 | 3.2596 | 0.0038877 | -1.2314 | -1.1672 | 0.064178 |  | 40.413 | -1.2553
8 | YES | 40.413 | 3.4167 | 0.0019968 | -1.2553 | -1.2591 | -0.0038022 |  | 30.5348 | -1.2545
8 | YES | 30.5348 | 3.3345 | 0.0049397 | -1.2545 | -1.2853 | -0.030807 |  | 30.5348 | -1.4965
8 | YES | 30.5348 | 3.3345 | 0.0049397 | -1.4965 | -1.2853 | 0.21119 |  | 20.275 | -1.1716
8 | YES | 20.275 | 3.2694 | 0.0015147 | -1.1716 | -1.3125 | -0.14094 |  | 9.9533 | -1.2191
8 | YES | 9.9533 | 3.3589 | 0.0022456 | -1.2191 | -1.3399 | -0.1208 |  | 5.2407 | -1.2875
8 | YES | 5.2407 | 3.4405 | 0.0022065 | -1.2875 | -1.3524 | -0.064914 |  | 2.3193 | -1.2528
8 | YES | 2.3193 | 3.4542 | 0.0013743 | -1.2528 | -1.3602 | -0.10737 |  | 175.577 | -1.1299
9 | YES | 175.577 | 2.9001 | 0.0019947 | -1.1299 | -0.90038 | 0.22955 |  | 150.277 | -1.3167
9 | YES | 150.277 | 2.9603 | 0.0010539 | -1.3167 | -0.96752 | 0.34918 |  | 125.206 | -1.0143
9 | YES | 125.206 | 3.1137 | 0.0016188 | -1.0143 | -1.0341 | -0.019734 |  | 99.9993 | -1.1507
9 | YES | 99.9993 | 3.1503 | 0.0018896 | -1.1507 | -1.1009 | 0.049731 |  | 74.5362 | -1.5692
9 | YES | 74.5362 | 3.2088 | 0.001315 | -1.5692 | -1.1685 | 0.40071 |  | 49.9354 | -1.1387
9 | YES | 49.9354 | 3.2253 | 0.0013127 | -1.1387 | -1.2338 | -0.095141 |  | 40.2109 | -1.1436
9 | YES | 40.2109 | 2.8404 | 0.0030207 | -1.1436 | -1.2596 | -0.11602 |  | 30.0634 | -0.91897
9 | YES | 30.0634 | 2.777 | 0.0011723 | -0.91897 | -1.2865 | -0.36757 |  | 19.8283 | -0.66167
9 | YES | 19.8283 | 2.9123 | 0.0021962 | -0.66167 | -1.3137 | -0.65204 |  | 9.7016 | -1.2153
9 | YES | 9.7016 | 3.6487 | 0.0041504 | -1.2153 | -1.3406 | -0.12533 |  | 4.6399 | -1.4303
9 | YES | 4.6399 | 3.7187 | 0.0014812 | -1.4303 | -1.354 | 0.076281 |  | 1.5897 | -1.5512
9 | YES | 1.5897 | 3.7189 | 0.00065398 | -1.5512 | -1.3621 | 0.18904 |  | 310.053 | -0.83371
10 | YES | 310.053 | 2.1853 | 0.0014913 | -0.83371 | -0.5435 | 0.29021 |  | 300.001 | -0.75734
10 | YES | 300.001 | 2.1067 | 0.0036596 | -0.75734 | -0.57017 | 0.18717 |  | 249.985 | -0.86203
10 | YES | 249.985 | 2.043 | 0.00082121 | -0.86203 | -0.70291 | 0.15912 |  | 200.039 | -0.87897
10 | YES | 200.039 | 2.114 | 0.0017535 | -0.87897 | -0.83546 | 0.043517 |  | 175.027 | -0.82663
10 | YES | 175.027 | 2.1614 | 0.00099275 | -0.82663 | -0.90184 | -0.075208 |  | 150.02 | -0.84039
10 | YES | 150.02 | 2.2426 | 0.0015021 | -0.84039 | -0.9682 | -0.12781 |  | 124.967 | -0.88602
10 | YES | 124.967 | 2.468 | 0.0014259 | -0.88602 | -1.0347 | -0.14867 |  | 100.176 | -0.96526
10 | YES | 100.176 | 2.6637 | 0.0021125 | -0.96526 | -1.1005 | -0.13522 |  | 75.1188 | -1.2354
10 | YES | 75.1188 | 2.9876 | 0.0015307 | -1.2354 | -1.167 | 0.068422 |  | 50.0965 | -1.1975
10 | YES | 50.0965 | 2.8675 | 0.0016284 | -1.1975 | -1.2334 | -0.035864 |  | 40.2836 | -1.0343
10 | YES | 40.2836 | 2.8467 | 0.00094862 | -1.0343 | -1.2594 | -0.22516 |  | 30.0779 | -1.1856
10 | YES | 30.0779 | 2.8484 | 0.0018941 | -1.1856 | -1.2865 | -0.10089 |  | 20.1112 | -1.0735
10 | YES | 20.1112 | 2.9955 | 0.0017892 | -1.0735 | -1.313 | -0.23949 |  | 10.1999 | -1.6114
10 | YES | 10.1999 | 4.3846 | 0.0017672 | -1.6114 | -1.3393 | 0.27213 |  | 5.2541 | -1.9867
10 | YES | 5.2541 | 4.4924 | 0.001174 | -1.9867 | -1.3524 | 0.63426 |  | 1.6808 | -1.7292
10 | YES | 1.6808 | 4.5088 | 0.0016263 | -1.7292 | -1.3619 | 0.3673 |  | 372.67 | -0.76655
11 | YES | 372.67 | 1.8615 | 0.0013598 | -0.76655 | -0.37732 | 0.38923 |  | 349.872 | -0.71092
11 | YES | 349.872 | 1.9021 | 0.0073457 | -0.71092 | -0.43782 | 0.27309 |  | 299.899 | -0.7919
11 | YES | 299.899 | 1.9921 | 0.00062458 | -0.7919 | -0.57044 | 0.22146 |  | 249.773 | -0.78428
11 | YES | 249.773 | 2.0347 | 0.0019374 | -0.78428 | -0.70347 | 0.08081 |  | 199.882 | -0.79413
11 | YES | 199.882 | 2.1149 | 0.00071811 | -0.79413 | -0.83587 | -0.041742 |  | 174.692 | -1.0308
11 | YES | 174.692 | 2.1962 | 0.0013709 | -1.0308 | -0.90272 | 0.12805 |  | 150.058 | -0.93454
11 | YES | 150.058 | 2.2815 | 0.0030442 | -0.93454 | -0.9681 | -0.033563 |  | 124.78 | -1.1144
11 | YES | 124.78 | 2.5196 | 0.0033291 | -1.1144 | -1.0352 | 0.079176 |  | 99.9088 | -0.84642
11 | YES | 99.9088 | 2.3226 | 0.00233 | -0.84642 | -1.1012 | -0.25477 |  | 74.8 | -1.1737
11 | YES | 74.8 | 2.3213 | 0.003395 | -1.1737 | -1.1678 | 0.005836 |  | 49.9199 | -0.94079
11 | YES | 49.9199 | 2.4142 | 0.010427 | -0.94079 | -1.2339 | -0.29306 |  | 40.1629 | -1.0502
11 | YES | 40.1629 | 2.6878 | 0.026679 | -1.0502 | -1.2597 | -0.20956 |  | 19.961 | -1.3928
11 | YES | 19.961 | 3.5582 | 0.032358 | -1.3928 | -1.3134 | 0.079441 |  | 9.9586 | -1.7684
11 | YES | 9.9586 | 4.1226 | 0.073385 | -1.7684 | -1.3399 | 0.42854 |  | 5.2056 | -1.6147
11 | YES | 5.2056 | 4.3373 | 0.014591 | -1.6147 | -1.3525 | 0.26215 |  | 1.7664 | -1.6705
11 | YES | 1.7664 | 4.5755 | 0.0059289 | -1.6705 | -1.3616 | 0.30888 |  | 340.021 | -0.74529
12 | YES | 340.021 | 1.9607 | 0.0018773 | -0.74529 | -0.46397 | 0.28132 |  | 299.463 | -0.72101
12 | YES | 299.463 | 2.017 | 0.0014171 | -0.72101 | -0.5716 | 0.14941 |  | 250.243 | -0.81095
12 | YES | 250.243 | 2.0541 | 0.00038409 | -0.81095 | -0.70222 | 0.10873 |  | 199.683 | -0.77797
12 | YES | 199.683 | 2.152 | 0.0017534 | -0.77797 | -0.8364 | -0.058427 |  | 174.672 | -0.79903
12 | YES | 174.672 | 2.264 | 0.0014079 | -0.79903 | -0.90278 | -0.10375 |  | 149.647 | -0.84325
12 | YES | 149.647 | 2.3908 | 0.0019161 | -0.84325 | -0.96919 | -0.12594 |  | 124.71 | -0.82095
12 | YES | 124.71 | 2.3441 | 0.00054543 | -0.82095 | -1.0354 | -0.21442 |  | 99.9992 | -0.93299
12 | YES | 99.9992 | 2.574 | 0.00045636 | -0.93299 | -1.1009 | -0.16796 |  | 74.821 | -0.91317
12 | YES | 74.821 | 2.5228 | 0.0018288 | -0.91317 | -1.1678 | -0.25459 |  | 50.0241 | -0.78807
12 | YES | 50.0241 | 2.2669 | 0.00065492 | -0.78807 | -1.2336 | -0.44551 |  | 40.106 | -0.92683
12 | YES | 40.106 | 2.4772 | 0.0042766 | -0.92683 | -1.2599 | -0.33306 |  | 30.1896 | -1.0121
12 | YES | 30.1896 | 2.8579 | 0.0015253 | -1.0121 | -1.2862 | -0.27413 |  | 20.0473 | -1.4359
12 | YES | 20.0473 | 3.5491 | 0.0027415 | -1.4359 | -1.3131 | 0.12274 |  | 9.9047 | -1.5361
12 | YES | 9.9047 | 4.4419 | 0.0037537 | -1.5361 | -1.34 | 0.19605 |  | 5.0873 | -1.647
12 | YES | 5.0873 | 4.572 | 0.0019142 | -1.647 | -1.3528 | 0.29422 |  | 1.8965 | -1.6963
12 | YES | 1.8965 | 4.5837 | 0.00087346 | -1.6963 | -1.3613 | 0.33496 |  | 275.184 | -0.8494
13 | YES | 275.184 | 1.9426 | 0.0011141 | -0.8494 | -0.63603 | 0.21336 |  | 250.05 | -0.69889
13 | YES | 250.05 | 1.9541 | 0.001413 | -0.69889 | -0.70274 | -0.0038427 |  | 200.053 | -0.69654
13 | YES | 200.053 | 2.0225 | 0.0019793 | -0.69654 | -0.83542 | -0.13888 |  | 174.464 | -0.78212
13 | YES | 174.464 | 2.0639 | 0.0018328 | -0.78212 | -0.90333 | -0.12121 |  | 150.006 | -1.1339
13 | YES | 150.006 | 2.1951 | 0.0022581 | -1.1339 | -0.96824 | 0.16563 |  | 124.876 | -0.857
13 | YES | 124.876 | 2.346 | 0.0024427 | -0.857 | -1.0349 | -0.17793 |  | 100.109 | -0.95692
13 | YES | 100.109 | 2.4571 | 0.0027706 | -0.95692 | -1.1007 | -0.14374 |  | 74.8535 | -0.84271
13 | YES | 74.8535 | 2.4643 | 0.0021033 | -0.84271 | -1.1677 | -0.32497 |  | 50.1091 | -0.66783
13 | YES | 50.1091 | 2.5152 | 0.0024923 | -0.66783 | -1.2334 | -0.56552 |  | 40.206 | -0.92455
13 | YES | 40.206 | 2.7454 | 0.01255 | -0.92455 | -1.2596 | -0.33508 |  | 30.1437 | -1.1811
13 | YES | 30.1437 | 2.9619 | 0.010664 | -1.1811 | -1.2863 | -0.10522 |  | 20.203 | -1.5093
13 | YES | 20.203 | 3.4817 | 0.082422 | -1.5093 | -1.3127 | 0.19658 |  | 10.0564 | -1.8821
13 | YES | 10.0564 | 4.4209 | 0.022527 | -1.8821 | -1.3396 | 0.54247 |  | 5.0958 | -1.8568
13 | YES | 5.0958 | 4.5602 | 0.0061208 | -1.8568 | -1.3528 | 0.50398 |  | 1.9085 | -1.7627
13 | YES | 1.9085 | 4.5853 | 0.0018394 | -1.7627 | -1.3613 | 0.40142 |  | 316.832 | -0.7483
14 | YES | 316.832 | 1.8967 | 0.0023439 | -0.7483 | -0.52551 | 0.22279 |  | 300.016 | -0.70599
14 | YES | 300.016 | 1.917 | 0.0017344 | -0.70599 | -0.57013 | 0.13586 |  | 250.279 | -0.9191
14 | YES | 250.279 | 1.9599 | 0.0005386 | -0.9191 | -0.70213 | 0.21697 |  | 199.862 | -0.73608
14 | YES | 199.862 | 2.0679 | 0.001688 | -0.73608 | -0.83593 | -0.099847 |  | 175.252 | -0.77931
14 | YES | 175.252 | 2.2057 | 0.0012705 | -0.77931 | -0.90124 | -0.12193 |  | 150.242 | -0.80469
14 | YES | 150.242 | 2.2903 | 0.0063138 | -0.80469 | -0.96761 | -0.16292 |  | 124.737 | -0.86928
14 | YES | 124.737 | 2.4217 | 0.0057012 | -0.86928 | -1.0353 | -0.16602 |  | 100.157 | -0.84585
14 | YES | 100.157 | 2.3931 | 0.00065398 | -0.84585 | -1.1005 | -0.25468 |  | 75.1673 | -0.79336
14 | YES | 75.1673 | 2.3126 | 0.0014999 | -0.79336 | -1.1668 | -0.37349 |  | 49.8643 | -0.98908
14 | YES | 49.8643 | 2.5959 | 0.0017681 | -0.98908 | -1.234 | -0.24492 |  | 40.2218 | -1.1643
14 | YES | 40.2218 | 3.3327 | 0.005416 | -1.1643 | -1.2596 | -0.09532 |  | 29.8998 | -1.3892
14 | YES | 29.8998 | 3.8398 | 0.0024037 | -1.3892 | -1.287 | 0.10219 |  | 20.2758 | -1.5396
14 | YES | 20.2758 | 4.0304 | 0.0042531 | -1.5396 | -1.3125 | 0.22709 |  | 10.1761 | -1.3884
14 | YES | 10.1761 | 4.3705 | 0.0012247 | -1.3884 | -1.3393 | 0.049123 |  | 5.3946 | -1.5966
14 | YES | 5.3946 | 4.4354 | 0.0026729 | -1.5966 | -1.352 | 0.24457 |  | 2.1456 | -1.594
14 | YES | 2.1456 | 4.513 | 0.0017744 | -1.594 | -1.3606 | 0.23339 |  | 188.75 | -0.91527
15 | YES | 188.75 | 2.0777 | 0.00086603 | -0.91527 | -0.86542 | 0.049856 |  | 174.696 | -0.79469
15 | YES | 174.696 | 2.1083 | 0.0021441 | -0.79469 | -0.90271 | -0.10802 |  | 150.187 | -0.77049
15 | YES | 150.187 | 2.1915 | 0.0028229 | -0.77049 | -0.96776 | -0.19727 |  | 125.052 | -0.83222
15 | YES | 125.052 | 2.3058 | 0.00085135 | -0.83222 | -1.0345 | -0.20224 |  | 100.088 | -0.80521
15 | YES | 100.088 | 2.2588 | 0.0022506 | -0.80521 | -1.1007 | -0.29551 |  | 74.9715 | -0.79574
15 | YES | 74.9715 | 2.1533 | 0.0024101 | -0.79574 | -1.1674 | -0.37162 |  | 49.9686 | -0.94049
15 | YES | 49.9686 | 2.5965 | 0.00769 | -0.94049 | -1.2337 | -0.29323 |  | 39.9753 | -1.5717
15 | YES | 39.9753 | 3.2743 | 0.029715 | -1.5717 | -1.2602 | 0.31145 |  | 30.2688 | -1.4408
15 | YES | 30.2688 | 3.7542 | 0.0054148 | -1.4408 | -1.286 | 0.15479 |  | 20.1448 | -1.4453
15 | YES | 20.1448 | 3.8327 | 0.0044506 | -1.4453 | -1.3129 | 0.13242 |  | 10.0866 | -1.5984
15 | YES | 10.0866 | 4.2636 | 0.0069756 | -1.5984 | -1.3396 | 0.25884 |  | 5.1861 | -1.5879
15 | YES | 5.1861 | 4.3391 | 0.0037943 | -1.5879 | -1.3526 | 0.23532 |  | 1.9275 | -1.6261
15 | YES | 1.9275 | 4.4039 | 0.0042631 | -1.6261 | -1.3612 | 0.26488 |  | 165.084 | -0.84145
16 | YES | 165.084 | 1.8476 | 0.0011619 | -0.84145 | -0.92822 | -0.086769 |  | 150.005 | -0.72895
16 | YES | 150.005 | 1.872 | 0.0084979 | -0.72895 | -0.96824 | -0.23929 |  | 124.87 | -0.74898
16 | YES | 124.87 | 1.977 | 0.0014943 | -0.74898 | -1.0349 | -0.28596 |  | 100.004 | -0.75646
16 | YES | 100.004 | 1.9685 | 0.0079057 | -0.75646 | -1.1009 | -0.34447 |  | 74.8941 | -0.97544
16 | YES | 74.8941 | 2.0526 | 0.0021598 | -0.97544 | -1.1676 | -0.19214 |  | 50.3432 | -0.97513
16 | YES | 50.3432 | 2.4929 | 0.0081024 | -0.97513 | -1.2327 | -0.2576 |  | 40.1958 | -1.3681
16 | YES | 40.1958 | 2.9779 | 0.10133 | -1.3681 | -1.2597 | 0.10841 |  | 30.075 | -1.5512
16 | YES | 30.075 | 3.7348 | 0.066807 | -1.5512 | -1.2865 | 0.26469 |  | 20.0364 | -1.7787
16 | YES | 20.0364 | 4.3003 | 0.033634 | -1.7787 | -1.3132 | 0.46551 |  | 10.1725 | -1.722
16 | YES | 10.1725 | 4.445 | 0.0027095 | -1.722 | -1.3393 | 0.38265 |  | 4.9316 | -1.6707
16 | YES | 4.9316 | 4.4583 | 0.0050078 | -1.6707 | -1.3532 | 0.31746 |  | 1.686 | -1.6293
16 | YES | 1.686 | 4.4707 | 0.0018532 | -1.6293 | -1.3619 | 0.26745 |  | 350.547 | -0.7863
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7863 | -0.43603 | 0.35027 |  | 350.547 | -0.7813
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7813 | -0.43603 | 0.34527 |  | 300.413 | -0.61523
17 | YES | 300.413 | 1.9158 | 0.00093775 | -0.61523 | -0.56908 | 0.046151 |  | 249.829 | -0.72518
17 | YES | 249.829 | 1.9758 | 0.00099161 | -0.72518 | -0.70332 | 0.02186 |  | 200.363 | -1.0292
17 | YES | 200.363 | 2.0128 | 0.00092755 | -1.0292 | -0.8346 | 0.1946 |  | 175.126 | -0.7752
17 | YES | 175.126 | 2.1298 | 0.00074856 | -0.7752 | -0.90157 | -0.12638 |  | 149.927 | -0.96417
17 | YES | 149.927 | 2.1218 | 0.0007033 | -0.96417 | -0.96845 | -0.0042738 |  | 124.734 | -0.90002
17 | YES | 124.734 | 2.186 | 0.00061185 | -0.90002 | -1.0353 | -0.13528 |  | 99.8058 | -0.70816
17 | YES | 99.8058 | 1.9368 | 0.0018664 | -0.70816 | -1.1015 | -0.39331 |  | 74.8425 | -0.75636
17 | YES | 74.8425 | 2.2816 | 0.0015 | -0.75636 | -1.1677 | -0.41135 |  | 50.1359 | -1.7333
17 | YES | 50.1359 | 4.1537 | 0.0094609 | -1.7333 | -1.2333 | 0.50006 |  | 40.3131 | -1.6338
17 | YES | 40.3131 | 4.3472 | 0.003735 | -1.6338 | -1.2593 | 0.37449 |  | 20.1101 | -1.702
17 | NO | 30.2081 | 4.4076 | 0.0017444 | -4.4854 | -1.2862 | 3.1992 |  | 10.319 | -1.3213
17 | YES | 20.1101 | 4.397 | 0.0017841 | -1.702 | -1.313 | 0.38906 |  | 5.1317 | -1.7093
17 | YES | 10.319 | 4.6757 | 0.001782 | -1.3213 | -1.3389 | -0.017687 |  | 2.2104 | -1.762
17 | YES | 5.1317 | 4.7157 | 0.0056502 | -1.7093 | -1.3527 | 0.35663 |  | 312.193 | -0.74222
17 | YES | 2.2104 | 4.77 | 0.0019683 | -1.762 | -1.3605 | 0.40151 |  | 300.322 | -0.72805
18 | YES | 312.193 | 1.8728 | 0.0021506 | -0.74222 | -0.53782 | 0.20441 |  | 250.544 | -0.74515
18 | YES | 300.322 | 1.8729 | 0.001244 | -0.72805 | -0.56932 | 0.15873 |  | 200.207 | -0.81876
18 | YES | 250.544 | 1.9048 | 0.0019003 | -0.74515 | -0.70142 | 0.043724 |  | 175.174 | -0.82617
18 | YES | 200.207 | 1.9462 | 0.001049 | -0.81876 | -0.83501 | -0.016252 |  | 150.243 | -0.72261
18 | YES | 175.174 | 1.8698 | 0.0017449 | -0.82617 | -0.90145 | -0.075272 |  | 124.303 | -0.67167
18 | YES | 150.243 | 1.8704 | 0.0018725 | -0.72261 | -0.96761 | -0.245 |  | 99.7041 | -1.1105
18 | YES | 124.303 | 1.8573 | 0.0011059 | -0.67167 | -1.0365 | -0.36478 |  | 74.792 | -1.1766
18 | YES | 99.7041 | 1.9255 | 0.0031388 | -1.1105 | -1.1017 | 0.0087679 |  | 50.1137 | -1.4046
18 | YES | 74.792 | 2.6254 | 0.034142 | -1.1766 | -1.1678 | 0.0087548 |  | 39.9031 | -1.4981
18 | YES | 50.1137 | 3.2384 | 0.032568 | -1.4046 | -1.2333 | 0.1713 |  | 30.4567 | -1.5236
18 | YES | 39.9031 | 3.5309 | 0.012688 | -1.4981 | -1.2604 | 0.23762 |  | 19.8381 | -1.6106
18 | YES | 30.4567 | 3.6134 | 0.0034662 | -1.5236 | -1.2855 | 0.23806 |  | 9.9176 | -1.5227
18 | YES | 19.8381 | 3.9144 | 0.017966 | -1.6106 | -1.3137 | 0.29692 |  | 4.8786 | -1.528
18 | YES | 9.9176 | 4.1413 | 0.0043618 | -1.5227 | -1.34 | 0.18273 |  | 1.9734 | -1.5553
18 | YES | 4.8786 | 4.178 | 0.0015186 | -1.528 | -1.3534 | 0.17457 |  | 145.098 | -0.78
18 | YES | 1.9734 | 4.2067 | 0.0022344 | -1.5553 | -1.3611 | 0.19416 |  | 124.892 | -0.6772
19 | YES | 145.098 | 1.916 | 0 | -0.78 | -0.98126 | -0.20126 |  | 100.175 | -0.66893
19 | YES | 124.892 | 1.9098 | 0.00074856 | -0.6772 | -1.0349 | -0.35769 |  | 74.2929 | -0.66621
19 | YES | 100.175 | 1.9011 | 0.00051211 | -0.66893 | -1.1005 | -0.43155 |  | 50.1467 | -1.3011
19 | YES | 74.2929 | 1.9028 | 0.0014996 | -0.66621 | -1.1692 | -0.50296 |  | 40.2307 | -1.1415
19 | YES | 50.1467 | 2.7889 | 0.027121 | -1.3011 | -1.2333 | 0.06787 |  | 30.1963 | -1.2884
19 | YES | 40.2307 | 2.9675 | 0.0028229 | -1.1415 | -1.2596 | -0.11807 |  | 20.2351 | -1.1676
19 | YES | 30.1963 | 3.2236 | 0.011798 | -1.2884 | -1.2862 | 0.0021742 |  | 9.7635 | -1.2956
19 | YES | 20.2351 | 3.2764 | 0.0030684 | -1.1676 | -1.3126 | -0.14507 |  | 4.6938 | -1.3126
19 | YES | 9.7635 | 3.3764 | 0.013102 | -1.2956 | -1.3404 | -0.044802 |  | 1.5803 | -1.2637
19 | YES | 4.6938 | 3.5024 | 0.0049005 | -1.3126 | -1.3539 | -0.041316 |  | 231.357 | -1.1755
19 | YES | 1.5803 | 3.5423 | 0.0034886 | -1.2637 | -1.3621 | -0.098419 |  | 199.967 | -1.186
20 | YES | 231.357 | 3.2565 | 0.0019577 | -1.1755 | -0.75234 | 0.42311 |  | 174.934 | -1.092
20 | YES | 199.967 | 3.249 | 0.0012974 | -1.186 | -0.83565 | 0.35033 |  | 149.814 | -1.1885
20 | YES | 174.934 | 3.249 | 0.0012682 | -1.092 | -0.90208 | 0.18993 |  | 125.263 | -1.6057
20 | YES | 149.814 | 3.2455 | 0.001119 | -1.1885 | -0.96875 | 0.21975 |  | 99.9825 | -1.1909
20 | YES | 125.263 | 3.2443 | 0.0014919 | -1.6057 | -1.0339 | 0.57176 |  | 75.3592 | -1.1426
20 | YES | 99.9825 | 3.2461 | 0.0019845 | -1.1909 | -1.101 | 0.089867 |  | 50.0965 | -1.3199
20 | YES | 75.3592 | 3.2424 | 0.0021689 | -1.1426 | -1.1663 | -0.02372 |  | 40.1804 | -1.1456
20 | YES | 50.0965 | 3.2401 | 0.00089973 | -1.3199 | -1.2334 | 0.086496 |  | 29.974 | -1.0232
20 | YES | 40.1804 | 3.2424 | 0.001479 | -1.1456 | -1.2597 | -0.11408 |  | 20.0329 | -1.2513
20 | YES | 29.974 | 3.2437 | 0.0018542 | -1.0232 | -1.2868 | -0.26354 |  | 9.9679 | -1.1512
20 | YES | 20.0329 | 3.2427 | 0.0015187 | -1.2513 | -1.3132 | -0.061838 |  | 5.1428 | -1.173
20 | YES | 9.9679 | 3.2448 | 0.0016259 | -1.1512 | -1.3399 | -0.18868 |  | 2.0859 | -1.1843
20 | YES | 5.1428 | 3.242 | 0.0016532 | -1.173 | -1.3527 | -0.1797
20 | YES | 2.0859 | 3.2467 | 0.0015492 | -1.1843 | -1.3608 | -0.17653
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX-CTD - DOX-BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=DOX-CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.87983 |  | X | CALC Y
m= | -0.35932 | 0.0081574 |  | Std Dev of Resid |  | 0.12061 |  | 0.16698 | -0.12942
b= | -0.06942 | 0.024923 |  | Resid Sum of Squares |  | 3.8548 |  | 4.77 | -1.7834
# of Values Extracted = |  |  | 269
# of Values Used in Fit = |  |  | 267
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | DOX-CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 110.209 | 0.14646 | 0.0023558 | -0.59954 | -0.12204 | 0.4775 |  | 0.16698 | -0.053025
1 | YES | 100.191 | 0.16698 | 0.00027273 | -0.053025 | -0.12942 | -0.076393 |  | 2.6501 | -0.58691
1 | YES | 75.0704 | 2.6501 | 0.0053135 | -0.58691 | -1.0217 | -0.43475 |  | 3.6213 | -1.2407
1 | YES | 1.5835 | 3.6213 | 0.073599 | -1.2407 | -1.3706 | -0.12988 |  | 2.5645 | -0.95754
2 | YES | 231.111 | 2.5645 | 0.0012848 | -0.95754 | -0.99089 | -0.03335 |  | 2.6258 | -1.0462
2 | YES | 200.09 | 2.6258 | 0.0028774 | -1.0462 | -1.0129 | 0.033313 |  | 2.6535 | -1.1205
2 | YES | 175.687 | 2.6535 | 0.001119 | -1.1205 | -1.0229 | 0.097619 |  | 2.8042 | -1.3028
2 | YES | 150.335 | 2.8042 | 0.0040114 | -1.3028 | -1.077 | 0.22574 |  | 2.8934 | -1.1616
2 | YES | 125.097 | 2.8934 | 0.0035549 | -1.1616 | -1.1091 | 0.052545 |  | 3.0364 | -1.2356
2 | YES | 99.9256 | 3.0364 | 0.0029657 | -1.2356 | -1.1604 | 0.075203 |  | 3.3423 | -1.2787
2 | YES | 74.4687 | 3.3423 | 0.0074156 | -1.2787 | -1.2704 | 0.008332 |  | 3.4055 | -1.3215
2 | YES | 49.1603 | 3.4055 | 0.0020856 | -1.3215 | -1.2931 | 0.028341 |  | 3.4705 | -1.3115
2 | YES | 40.6798 | 3.4705 | 0.0019233 | -1.3115 | -1.3164 | -0.0048787 |  | 3.5254 | -1.2646
2 | YES | 29.7398 | 3.5254 | 0.0016215 | -1.2646 | -1.3362 | -0.071601 |  | 3.5421 | -1.3259
2 | YES | 19.6317 | 3.5421 | 0.00039364 | -1.3259 | -1.3422 | -0.01622 |  | 3.5499 | -1.2771
2 | YES | 9.6901 | 3.5499 | 0.0014419 | -1.2771 | -1.345 | -0.067918 |  | 3.5604 | -1.2555
2 | YES | 4.9828 | 3.5604 | 0.001072 | -1.2555 | -1.3488 | -0.093218 |  | 3.5762 | -1.3018
2 | YES | 1.6667 | 3.5762 | 0.0024279 | -1.3018 | -1.3544 | -0.052586 |  | 1.4684 | -0.58065
3 | YES | 244.526 | 1.4684 | 0.001215 | -0.58065 | -0.59703 | -0.016383 |  | 1.4684 | -0.81563
3 | YES | 244.611 | 1.4684 | 0.0012711 | -0.81563 | -0.59704 | 0.21859 |  | 1.4764 | -0.60664
3 | YES | 203.398 | 1.4764 | 0.0019011 | -0.60664 | -0.59991 | 0.0067368 |  | 1.6429 | -0.68709
3 | YES | 175.87 | 1.6429 | 0.0012649 | -0.68709 | -0.65975 | 0.027338 |  | 1.7275 | -1.0215
3 | YES | 149.282 | 1.7275 | 0.0018305 | -1.0215 | -0.69013 | 0.33141 |  | 2.1854 | -0.81862
3 | YES | 125.206 | 2.1854 | 0.0023953 | -0.81862 | -0.85467 | -0.036055 |  | 2.3844 | -0.98561
3 | YES | 99.6255 | 2.3844 | 0.0015021 | -0.98561 | -0.92618 | 0.059428 |  | 2.5879 | -1.0281
3 | YES | 75.1051 | 2.5879 | 0.0012538 | -1.0281 | -0.99929 | 0.028845 |  | 3.2923 | -1.2957
3 | YES | 50.4953 | 3.2923 | 0.0013106 | -1.2957 | -1.2524 | 0.04323 |  | 3.7275 | -1.5615
3 | YES | 41.1572 | 3.7275 | 0.0018261 | -1.5615 | -1.4088 | 0.15275 |  | 4.41 | -1.709
3 | YES | 29.3588 | 4.41 | 0.0036958 | -1.709 | -1.654 | 0.054984 |  | 4.5408 | -1.6802
3 | YES | 20.5425 | 4.5408 | 0.0016749 | -1.6802 | -1.701 | -0.020843 |  | 4.5747 | -1.6583
3 | YES | 10.4123 | 4.5747 | 0.0010919 | -1.6583 | -1.7132 | -0.054964 |  | 4.5799 | -1.6551
3 | YES | 5.0724 | 4.5799 | 0.0015894 | -1.6551 | -1.7151 | -0.060006 |  | 4.5846 | -1.7904
3 | YES | 1.8669 | 4.5846 | 0.0023584 | -1.7904 | -1.7168 | 0.073646 |  | 1.4957 | -0.54429
4 | YES | 221.623 | 1.4957 | 0.001313 | -0.54429 | -0.60686 | -0.062572 |  | 1.4889 | -0.55505
4 | YES | 201.791 | 1.4889 | 0.00038409 | -0.55505 | -0.60443 | -0.049382 |  | 1.5309 | -0.65709
4 | YES | 172.797 | 1.5309 | 0.0023629 | -0.65709 | -0.61951 | 0.037582 |  | 1.7571 | -0.72293
4 | YES | 150.467 | 1.7571 | 0.0038257 | -0.72293 | -0.70077 | 0.022152 |  | 1.9126 | -1.0374
4 | YES | 124.65 | 1.9126 | 0.0019911 | -1.0374 | -0.75667 | 0.28069 |  | 2.1878 | -0.90424
4 | YES | 99.9203 | 2.1878 | 0.024804 | -0.90424 | -0.85553 | 0.04871 |  | 2.68 | -1.091
4 | YES | 75.1964 | 2.68 | 0.01257 | -1.091 | -1.0324 | 0.058597 |  | 2.9895 | -1.1775
4 | YES | 50.3963 | 2.9895 | 0.005905 | -1.1775 | -1.1436 | 0.03386 |  | 2.9873 | -1.1277
4 | YES | 40.4829 | 2.9873 | 0.00256 | -1.1277 | -1.1428 | -0.015095 |  | 3.2054 | -1.1636
4 | YES | 29.9081 | 3.2054 | 0.021531 | -1.1636 | -1.2212 | -0.057618 |  | 3.6917 | -1.3863
4 | YES | 20.177 | 3.6917 | 0.016997 | -1.3863 | -1.3959 | -0.009683 |  | 4.1154 | -1.8326
4 | YES | 9.7949 | 4.1154 | 0.032814 | -1.8326 | -1.5482 | 0.28441 |  | 4.4224 | -1.6496
4 | YES | 5.2267 | 4.4224 | 0.010058 | -1.6496 | -1.6585 | -0.0088309 |  | 4.4949 | -1.3901
4 | YES | 2.1258 | 4.4949 | 0.00088504 | -1.3901 | -1.6845 | -0.29445 |  | 1.551 | -0.64399
5 | YES | 184.313 | 1.551 | 0.0014171 | -0.64399 | -0.62673 | 0.01726 |  | 1.5436 | -0.51441
5 | YES | 175.011 | 1.5436 | 0.0014294 | -0.51441 | -0.62407 | -0.10966 |  | 1.5529 | -0.50614
5 | YES | 149.623 | 1.5529 | 0.002364 | -0.50614 | -0.6274 | -0.12126 |  | 1.6841 | -0.59293
5 | YES | 124.73 | 1.6841 | 0.0020007 | -0.59293 | -0.67454 | -0.081617 |  | 1.9403 | -0.67674
5 | YES | 99.845 | 1.9403 | 0.0018642 | -0.67674 | -0.7666 | -0.089854 |  | 2.3818 | -0.91923
5 | YES | 74.951 | 2.3818 | 0.0014246 | -0.91923 | -0.92524 | -0.006011 |  | 2.845 | -0.93303
5 | YES | 50.2894 | 2.845 | 0.0020933 | -0.93303 | -1.0917 | -0.15865 |  | 3.1446 | -1.1514
5 | YES | 40.0385 | 3.1446 | 0.0018432 | -1.1514 | -1.1993 | -0.047903 |  | 3.6185 | -1.2485
5 | YES | 30.0153 | 3.6185 | 0.0021177 | -1.2485 | -1.3696 | -0.12117 |  | 4.3044 | -1.3606
5 | YES | 19.4498 | 4.3044 | 0.0035562 | -1.3606 | -1.6161 | -0.25553 |  | 4.3526 | -1.7584
5 | YES | 9.701 | 4.3526 | 0.002138 | -1.7584 | -1.6334 | 0.12497 |  | 4.3721 | -1.6149
5 | YES | 4.9995 | 4.3721 | 0.0018055 | -1.6149 | -1.6404 | -0.025539 |  | 4.3726 | -1.6004
5 | YES | 1.5612 | 4.3726 | 0.0019213 | -1.6004 | -1.6406 | -0.040191 |  | 1.6122 | -0.56283
6 | YES | 171.173 | 1.6122 | 0.001202 | -0.56283 | -0.64871 | -0.085882 |  | 1.6273 | -0.57172
6 | YES | 149.783 | 1.6273 | 0.0015823 | -0.57172 | -0.65414 | -0.082418 |  | 1.6392 | -0.59684
6 | YES | 124.912 | 1.6392 | 0.0018545 | -0.59684 | -0.65841 | -0.061574 |  | 1.8678 | -0.8592
6 | YES | 98.8446 | 1.8678 | 0.0042556 | -0.8592 | -0.74056 | 0.11864 |  | 2.0953 | -0.8727
6 | YES | 74.2203 | 2.0953 | 0.0018423 | -0.8727 | -0.82231 | 0.050395 |  | 2.4122 | -0.92179
6 | YES | 49.47 | 2.4122 | 0.0020652 | -0.92179 | -0.93618 | -0.014387 |  | 2.5562 | -1.0238
6 | YES | 39.8415 | 2.5562 | 0.0072749 | -1.0238 | -0.98793 | 0.035854 |  | 2.9242 | -1.3838
6 | YES | 32.8962 | 2.9242 | 0.0096326 | -1.3838 | -1.1202 | 0.26364 |  | 3.1593 | -1.2837
6 | YES | 22.9881 | 3.1593 | 0.0019928 | -1.2837 | -1.2046 | 0.079142 |  | 3.2449 | -1.3501
6 | YES | 13.0632 | 3.2449 | 0.0019118 | -1.3501 | -1.2354 | 0.11466 |  | 3.3821 | -1.4829
6 | YES | 4.2473 | 3.3821 | 0.0014905 | -1.4829 | -1.2847 | 0.19816 |  | 3.3723 | -1.2927
6 | YES | 1.983 | 3.3723 | 0.0024434 | -1.2927 | -1.2812 | 0.011512 |  | 2.323 | -0.862
7 | YES | 77.2534 | 2.323 | 0 | -0.862 | -0.90412 | -0.042125 |  | 2.4244 | -0.81963
7 | YES | 75.1932 | 2.4244 | 0.01458 | -0.81963 | -0.94055 | -0.12092 |  | 2.7389 | -1.0201
7 | YES | 49.5689 | 2.7389 | 0.016676 | -1.0201 | -1.0536 | -0.033439 |  | 2.9771 | -1.1619
7 | YES | 39.9511 | 2.9771 | 0.0073387 | -1.1619 | -1.1392 | 0.022716 |  | 3.2204 | -1.2746
7 | YES | 30.3125 | 3.2204 | 0.017383 | -1.2746 | -1.2266 | 0.048061 |  | 3.4166 | -1.2694
7 | YES | 20.0896 | 3.4166 | 0.0022552 | -1.2694 | -1.2971 | -0.027652 |  | 3.4661 | -1.247
7 | YES | 10.9783 | 3.4661 | 0.00075419 | -1.247 | -1.3148 | -0.067898 |  | 3.5019 | -1.3381
7 | YES | 2.8781 | 3.5019 | 0.0094609 | -1.3381 | -1.3277 | 0.010343 |  | 3.5019 | -1.2791
7 | YES | 2.8781 | 3.5019 | 0.0094609 | -1.2791 | -1.3277 | -0.048657 |  | 2.7516 | -1.1024
8 | YES | 210.266 | 2.7516 | 0.0015037 | -1.1024 | -1.0581 | 0.044284 |  | 2.7605 | -1.0105
8 | YES | 200.335 | 2.7605 | 0.0011837 | -1.0105 | -1.0613 | -0.050787 |  | 2.7558 | -0.98719
8 | YES | 174.961 | 2.7558 | 0.0019929 | -0.98719 | -1.0596 | -0.072453 |  | 2.8096 | -1.0285
8 | YES | 150.142 | 2.8096 | 0.001072 | -1.0285 | -1.079 | -0.050503 |  | 2.9861 | -1.1309
8 | YES | 124.952 | 2.9861 | 0.00079902 | -1.1309 | -1.1424 | -0.011545 |  | 3.1033 | -1.1517
8 | YES | 99.7344 | 3.1033 | 0.0014975 | -1.1517 | -1.1845 | -0.032858 |  | 3.2596 | -1.2314
8 | YES | 75.0419 | 3.2596 | 0.0038877 | -1.2314 | -1.2407 | -0.00932 |  | 3.4167 | -1.2553
8 | YES | 40.413 | 3.4167 | 0.0019968 | -1.2553 | -1.2971 | -0.041842 |  | 3.3345 | -1.2545
8 | YES | 30.5348 | 3.3345 | 0.0049397 | -1.2545 | -1.2676 | -0.013092 |  | 3.3345 | -1.4965
8 | YES | 30.5348 | 3.3345 | 0.0049397 | -1.4965 | -1.2676 | 0.22891 |  | 3.2694 | -1.1716
8 | YES | 20.275 | 3.2694 | 0.0015147 | -1.1716 | -1.2442 | -0.072601 |  | 3.3589 | -1.2191
8 | YES | 9.9533 | 3.3589 | 0.0022456 | -1.2191 | -1.2763 | -0.057219 |  | 3.4405 | -1.2875
8 | YES | 5.2407 | 3.4405 | 0.0022065 | -1.2875 | -1.3057 | -0.018153 |  | 3.4542 | -1.2528
8 | YES | 2.3193 | 3.4542 | 0.0013743 | -1.2528 | -1.3106 | -0.057776 |  | 2.9001 | -1.1299
9 | YES | 175.577 | 2.9001 | 0.0019947 | -1.1299 | -1.1115 | 0.018451 |  | 2.9603 | -1.3167
9 | YES | 150.277 | 2.9603 | 0.0010539 | -1.3167 | -1.1331 | 0.18358 |  | 3.1137 | -1.0143
9 | YES | 125.206 | 3.1137 | 0.0016188 | -1.0143 | -1.1882 | -0.17391 |  | 3.1503 | -1.1507
9 | YES | 99.9993 | 3.1503 | 0.0018896 | -1.1507 | -1.2014 | -0.050719 |  | 3.2088 | -1.5692
9 | YES | 74.5362 | 3.2088 | 0.001315 | -1.5692 | -1.2224 | 0.34683 |  | 3.2253 | -1.1387
9 | YES | 49.9354 | 3.2253 | 0.0013127 | -1.1387 | -1.2284 | -0.089682 |  | 2.8404 | -1.1436
9 | YES | 40.2109 | 2.8404 | 0.0030207 | -1.1436 | -1.09 | 0.053562 |  | 2.777 | -0.91897
9 | YES | 30.0634 | 2.777 | 0.0011723 | -0.91897 | -1.0673 | -0.14829 |  | 2.9123 | -0.66167
9 | YES | 19.8283 | 2.9123 | 0.0021962 | -0.66167 | -1.1159 | -0.45421 |  | 3.6487 | -1.2153
9 | YES | 9.7016 | 3.6487 | 0.0041504 | -1.2153 | -1.3805 | -0.16523 |  | 3.7187 | -1.4303
9 | YES | 4.6399 | 3.7187 | 0.0014812 | -1.4303 | -1.4056 | 0.02467 |  | 3.7189 | -1.5512
9 | YES | 1.5897 | 3.7189 | 0.00065398 | -1.5512 | -1.4057 | 0.14547 |  | 2.1853 | -0.83371
10 | YES | 310.053 | 2.1853 | 0.0014913 | -0.83371 | -0.85464 | -0.020931 |  | 2.1067 | -0.75734
10 | YES | 300.001 | 2.1067 | 0.0036596 | -0.75734 | -0.82639 | -0.06905 |  | 2.043 | -0.86203
10 | YES | 249.985 | 2.043 | 0.00082121 | -0.86203 | -0.80351 | 0.058518 |  | 2.114 | -0.87897
10 | YES | 200.039 | 2.114 | 0.0017535 | -0.87897 | -0.82903 | 0.049941 |  | 2.1614 | -0.82663
10 | YES | 175.027 | 2.1614 | 0.00099275 | -0.82663 | -0.84605 | -0.019419 |  | 2.2426 | -0.84039
10 | YES | 150.02 | 2.2426 | 0.0015021 | -0.84039 | -0.87524 | -0.03485 |  | 2.468 | -0.88602
10 | YES | 124.967 | 2.468 | 0.0014259 | -0.88602 | -0.95622 | -0.070202 |  | 2.6637 | -0.96526
10 | YES | 100.176 | 2.6637 | 0.0021125 | -0.96526 | -1.0266 | -0.061296 |  | 2.9876 | -1.2354
10 | YES | 75.1188 | 2.9876 | 0.0015307 | -1.2354 | -1.1429 | 0.09247 |  | 2.8675 | -1.1975
10 | YES | 50.0965 | 2.8675 | 0.0016284 | -1.1975 | -1.0998 | 0.097752 |  | 2.8467 | -1.0343
10 | YES | 40.2836 | 2.8467 | 0.00094862 | -1.0343 | -1.0923 | -0.058042 |  | 2.8484 | -1.1856
10 | YES | 30.0779 | 2.8484 | 0.0018941 | -1.1856 | -1.0929 | 0.092715 |  | 2.9955 | -1.0735
10 | YES | 20.1112 | 2.9955 | 0.0017892 | -1.0735 | -1.1458 | -0.072309 |  | 4.3846 | -1.6114
10 | YES | 10.1999 | 4.3846 | 0.0017672 | -1.6114 | -1.6449 | -0.033516 |  | 4.4924 | -1.9867
10 | YES | 5.2541 | 4.4924 | 0.001174 | -1.9867 | -1.6836 | 0.30303 |  | 4.5088 | -1.7292
10 | YES | 1.6808 | 4.5088 | 0.0016263 | -1.7292 | -1.6895 | 0.039629 |  | 1.8615 | -0.76655
11 | YES | 372.67 | 1.8615 | 0.0013598 | -0.76655 | -0.73828 | 0.028274 |  | 1.9021 | -0.71092
11 | YES | 349.872 | 1.9021 | 0.0073457 | -0.71092 | -0.75288 | -0.041962 |  | 1.9921 | -0.7919
11 | YES | 299.899 | 1.9921 | 0.00062458 | -0.7919 | -0.78522 | 0.006676 |  | 2.0347 | -0.78428
11 | YES | 249.773 | 2.0347 | 0.0019374 | -0.78428 | -0.80054 | -0.016258 |  | 2.1149 | -0.79413
11 | YES | 199.882 | 2.1149 | 0.00071811 | -0.79413 | -0.82934 | -0.035207 |  | 2.1962 | -1.0308
11 | YES | 174.692 | 2.1962 | 0.0013709 | -1.0308 | -0.85857 | 0.1722 |  | 2.2815 | -0.93454
11 | YES | 150.058 | 2.2815 | 0.0030442 | -0.93454 | -0.8892 | 0.045339 |  | 2.5196 | -1.1144
11 | YES | 124.78 | 2.5196 | 0.0033291 | -1.1144 | -0.97478 | 0.13958 |  | 2.3226 | -0.84642
11 | YES | 99.9088 | 2.3226 | 0.00233 | -0.84642 | -0.90397 | -0.057553 |  | 2.3213 | -1.1737
11 | YES | 74.8 | 2.3213 | 0.003395 | -1.1737 | -0.90353 | 0.27013 |  | 2.4142 | -0.94079
11 | YES | 49.9199 | 2.4142 | 0.010427 | -0.94079 | -0.9369 | 0.0038956 |  | 2.6878 | -1.0502
11 | YES | 40.1629 | 2.6878 | 0.026679 | -1.0502 | -1.0352 | 0.014981 |  | 3.5582 | -1.3928
11 | YES | 19.961 | 3.5582 | 0.032358 | -1.3928 | -1.348 | 0.044841 |  | 4.1226 | -1.7684
11 | YES | 9.9586 | 4.1226 | 0.073385 | -1.7684 | -1.5507 | 0.21769 |  | 4.3373 | -1.6147
11 | YES | 5.2056 | 4.3373 | 0.014591 | -1.6147 | -1.6279 | -0.013248 |  | 4.5755 | -1.6705
11 | YES | 1.7664 | 4.5755 | 0.0059289 | -1.6705 | -1.7135 | -0.04297 |  | 1.9607 | -0.74529
12 | YES | 340.021 | 1.9607 | 0.0018773 | -0.74529 | -0.77395 | -0.028657 |  | 2.017 | -0.72101
12 | YES | 299.463 | 2.017 | 0.0014171 | -0.72101 | -0.79417 | -0.07316 |  | 2.0541 | -0.81095
12 | YES | 250.243 | 2.0541 | 0.00038409 | -0.81095 | -0.80748 | 0.003466 |  | 2.152 | -0.77797
12 | YES | 199.683 | 2.152 | 0.0017534 | -0.77797 | -0.84269 | -0.064713 |  | 2.264 | -0.79903
12 | YES | 174.672 | 2.264 | 0.0014079 | -0.79903 | -0.88291 | -0.083881 |  | 2.3908 | -0.84325
12 | YES | 149.647 | 2.3908 | 0.0019161 | -0.84325 | -0.92847 | -0.085221 |  | 2.3441 | -0.82095
12 | YES | 124.71 | 2.3441 | 0.00054543 | -0.82095 | -0.91169 | -0.090738 |  | 2.574 | -0.93299
12 | YES | 99.9992 | 2.574 | 0.00045636 | -0.93299 | -0.99432 | -0.061326 |  | 2.5228 | -0.91317
12 | YES | 74.821 | 2.5228 | 0.0018288 | -0.91317 | -0.97593 | -0.062754 |  | 2.2669 | -0.78807
12 | YES | 50.0241 | 2.2669 | 0.00065492 | -0.78807 | -0.88398 | -0.095911 |  | 2.4772 | -0.92683
12 | YES | 40.106 | 2.4772 | 0.0042766 | -0.92683 | -0.95952 | -0.032686 |  | 2.8579 | -1.0121
12 | YES | 30.1896 | 2.8579 | 0.0015253 | -1.0121 | -1.0963 | -0.084253 |  | 3.5491 | -1.4359
12 | YES | 20.0473 | 3.5491 | 0.0027415 | -1.4359 | -1.3447 | 0.09117 |  | 4.4419 | -1.5361
12 | YES | 9.9047 | 4.4419 | 0.0037537 | -1.5361 | -1.6655 | -0.12939 |  | 4.572 | -1.647
12 | YES | 5.0873 | 4.572 | 0.0019142 | -1.647 | -1.7122 | -0.065172 |  | 4.5837 | -1.6963
12 | YES | 1.8965 | 4.5837 | 0.00087346 | -1.6963 | -1.7165 | -0.020198 |  | 1.9426 | -0.8494
13 | YES | 275.184 | 1.9426 | 0.0011141 | -0.8494 | -0.76744 | 0.081958 |  | 1.9541 | -0.69889
13 | YES | 250.05 | 1.9541 | 0.001413 | -0.69889 | -0.77157 | -0.072681 |  | 2.0225 | -0.69654
13 | YES | 200.053 | 2.0225 | 0.0019793 | -0.69654 | -0.79613 | -0.099597 |  | 2.0639 | -0.78212
13 | YES | 174.464 | 2.0639 | 0.0018328 | -0.78212 | -0.81102 | -0.028893 |  | 2.1951 | -1.1339
13 | YES | 150.006 | 2.1951 | 0.0022581 | -1.1339 | -0.85818 | 0.27569 |  | 2.346 | -0.857
13 | YES | 124.876 | 2.346 | 0.0024427 | -0.857 | -0.91239 | -0.055389 |  | 2.4571 | -0.95692
13 | YES | 100.109 | 2.4571 | 0.0027706 | -0.95692 | -0.9523 | 0.0046155 |  | 2.4643 | -0.84271
13 | YES | 74.8535 | 2.4643 | 0.0021033 | -0.84271 | -0.95489 | -0.11218 |  | 2.5152 | -0.66783
13 | YES | 50.1091 | 2.5152 | 0.0024923 | -0.66783 | -0.97318 | -0.30535 |  | 2.7454 | -0.92455
13 | YES | 40.206 | 2.7454 | 0.01255 | -0.92455 | -1.0559 | -0.13137 |  | 2.9619 | -1.1811
13 | YES | 30.1437 | 2.9619 | 0.010664 | -1.1811 | -1.1337 | 0.047432 |  | 3.4817 | -1.5093
13 | YES | 20.203 | 3.4817 | 0.082422 | -1.5093 | -1.3205 | 0.18883 |  | 4.4209 | -1.8821
13 | YES | 10.0564 | 4.4209 | 0.022527 | -1.8821 | -1.6579 | 0.22417 |  | 4.5602 | -1.8568
13 | YES | 5.0958 | 4.5602 | 0.0061208 | -1.8568 | -1.708 | 0.14879 |  | 4.5853 | -1.7627
13 | YES | 1.9085 | 4.5853 | 0.0018394 | -1.7627 | -1.717 | 0.045668 |  | 1.8967 | -0.7483
14 | YES | 316.832 | 1.8967 | 0.0023439 | -0.7483 | -0.75095 | -0.0026476 |  | 1.917 | -0.70599
14 | YES | 300.016 | 1.917 | 0.0017344 | -0.70599 | -0.75824 | -0.052251 |  | 1.9599 | -0.9191
14 | YES | 250.279 | 1.9599 | 0.0005386 | -0.9191 | -0.77365 | 0.14544 |  | 2.0679 | -0.73608
14 | YES | 199.862 | 2.0679 | 0.001688 | -0.73608 | -0.81247 | -0.076389 |  | 2.2057 | -0.77931
14 | YES | 175.252 | 2.2057 | 0.0012705 | -0.77931 | -0.86197 | -0.082667 |  | 2.2903 | -0.80469
14 | YES | 150.242 | 2.2903 | 0.0063138 | -0.80469 | -0.89238 | -0.087684 |  | 2.4217 | -0.86928
14 | YES | 124.737 | 2.4217 | 0.0057012 | -0.86928 | -0.9396 | -0.070316 |  | 2.3931 | -0.84585
14 | YES | 100.157 | 2.3931 | 0.00065398 | -0.84585 | -0.92933 | -0.08348 |  | 2.3126 | -0.79336
14 | YES | 75.1673 | 2.3126 | 0.0014999 | -0.79336 | -0.9004 | -0.10704 |  | 2.5959 | -0.98908
14 | YES | 49.8643 | 2.5959 | 0.0017681 | -0.98908 | -1.0022 | -0.013108 |  | 3.3327 | -1.1643
14 | YES | 40.2218 | 3.3327 | 0.005416 | -1.1643 | -1.2669 | -0.10267 |  | 3.8398 | -1.3892
14 | YES | 29.8998 | 3.8398 | 0.0024037 | -1.3892 | -1.4492 | -0.059985 |  | 4.0304 | -1.5396
14 | YES | 20.2758 | 4.0304 | 0.0042531 | -1.5396 | -1.5176 | 0.021983 |  | 4.3705 | -1.3884
14 | YES | 10.1761 | 4.3705 | 0.0012247 | -1.3884 | -1.6399 | -0.2514 |  | 4.4354 | -1.5966
14 | YES | 5.3946 | 4.4354 | 0.0026729 | -1.5966 | -1.6632 | -0.06657 |  | 4.513 | -1.594
14 | YES | 2.1456 | 4.513 | 0.0017744 | -1.594 | -1.691 | -0.096999 |  | 2.0777 | -0.91527
15 | YES | 188.75 | 2.0777 | 0.00086603 | -0.91527 | -0.81599 | 0.099279 |  | 2.1083 | -0.79469
15 | YES | 174.696 | 2.1083 | 0.0021441 | -0.79469 | -0.82698 | -0.032288 |  | 2.1915 | -0.77049
15 | YES | 150.187 | 2.1915 | 0.0028229 | -0.77049 | -0.85688 | -0.086389 |  | 2.3058 | -0.83222
15 | YES | 125.052 | 2.3058 | 0.00085135 | -0.83222 | -0.89794 | -0.065714 |  | 2.2588 | -0.80521
15 | YES | 100.088 | 2.2588 | 0.0022506 | -0.80521 | -0.88105 | -0.075846 |  | 2.1533 | -0.79574
15 | YES | 74.9715 | 2.1533 | 0.0024101 | -0.79574 | -0.84313 | -0.047389 |  | 2.5965 | -0.94049
15 | YES | 49.9686 | 2.5965 | 0.00769 | -0.94049 | -1.0024 | -0.061915 |  | 3.2743 | -1.5717
15 | YES | 39.9753 | 3.2743 | 0.029715 | -1.5717 | -1.246 | 0.32574 |  | 3.7542 | -1.4408
15 | YES | 30.2688 | 3.7542 | 0.0054148 | -1.4408 | -1.4184 | 0.022397 |  | 3.8327 | -1.4453
15 | YES | 20.1448 | 3.8327 | 0.0044506 | -1.4453 | -1.4466 | -0.0013067 |  | 4.2636 | -1.5984
15 | YES | 10.0866 | 4.2636 | 0.0069756 | -1.5984 | -1.6014 | -0.0030249 |  | 4.3391 | -1.5879
15 | YES | 5.1861 | 4.3391 | 0.0037943 | -1.5879 | -1.6286 | -0.040667 |  | 4.4039 | -1.6261
15 | YES | 1.9275 | 4.4039 | 0.0042631 | -1.6261 | -1.6518 | -0.025738 |  | 1.8476 | -0.84145
16 | YES | 165.084 | 1.8476 | 0.0011619 | -0.84145 | -0.73328 | 0.10817 |  | 1.872 | -0.72895
16 | YES | 150.005 | 1.872 | 0.0084979 | -0.72895 | -0.74209 | -0.013137 |  | 1.977 | -0.74898
16 | YES | 124.87 | 1.977 | 0.0014943 | -0.74898 | -0.77981 | -0.030822 |  | 1.9685 | -0.75646
16 | YES | 100.004 | 1.9685 | 0.0079057 | -0.75646 | -0.77676 | -0.020296 |  | 2.0526 | -0.97544
16 | YES | 74.8941 | 2.0526 | 0.0021598 | -0.97544 | -0.80695 | 0.16849 |  | 2.4929 | -0.97513
16 | YES | 50.3432 | 2.4929 | 0.0081024 | -0.97513 | -0.96516 | 0.0099693 |  | 2.9779 | -1.3681
16 | YES | 40.1958 | 2.9779 | 0.10133 | -1.3681 | -1.1395 | 0.22861 |  | 3.7348 | -1.5512
16 | YES | 30.075 | 3.7348 | 0.066807 | -1.5512 | -1.4114 | 0.1398 |  | 4.3003 | -1.7787
16 | YES | 20.0364 | 4.3003 | 0.033634 | -1.7787 | -1.6146 | 0.16405 |  | 4.445 | -1.722
16 | YES | 10.1725 | 4.445 | 0.0027095 | -1.722 | -1.6666 | 0.05538 |  | 4.4583 | -1.6707
16 | YES | 4.9316 | 4.4583 | 0.0050078 | -1.6707 | -1.6714 | -0.00068491 |  | 4.4707 | -1.6293
16 | YES | 1.686 | 4.4707 | 0.0018532 | -1.6293 | -1.6758 | -0.046527 |  | 1.8207 | -0.7863
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7863 | -0.72364 | 0.062662 |  | 1.8207 | -0.7813
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7813 | -0.72364 | 0.057662 |  | 1.9158 | -0.61523
17 | YES | 300.413 | 1.9158 | 0.00093775 | -0.61523 | -0.7578 | -0.14257 |  | 1.9758 | -0.72518
17 | YES | 249.829 | 1.9758 | 0.00099161 | -0.72518 | -0.77938 | -0.054193 |  | 2.0128 | -1.0292
17 | YES | 200.363 | 2.0128 | 0.00092755 | -1.0292 | -0.79266 | 0.23654 |  | 2.1298 | -0.7752
17 | YES | 175.126 | 2.1298 | 0.00074856 | -0.7752 | -0.8347 | -0.059506 |  | 2.1218 | -0.96417
17 | YES | 149.927 | 2.1218 | 0.0007033 | -0.96417 | -0.83184 | 0.13233 |  | 2.186 | -0.90002
17 | YES | 124.734 | 2.186 | 0.00061185 | -0.90002 | -0.85489 | 0.045135 |  | 1.9368 | -0.70816
17 | YES | 99.8058 | 1.9368 | 0.0018664 | -0.70816 | -0.76537 | -0.057212 |  | 2.2816 | -0.75636
17 | YES | 74.8425 | 2.2816 | 0.0015 | -0.75636 | -0.88926 | -0.1329 |  | 4.1537 | -1.7333
17 | YES | 50.1359 | 4.1537 | 0.0094609 | -1.7333 | -1.5619 | 0.17142 |  | 4.3472 | -1.6338
17 | YES | 40.3131 | 4.3472 | 0.003735 | -1.6338 | -1.6314 | 0.0023901 |  | 4.397 | -1.702
17 | NO | 30.2081 | 4.4076 | 0.0017444 | -4.4854 | -1.6532 | 2.8322 |  | 4.6757 | -1.3213
17 | YES | 20.1101 | 4.397 | 0.0017841 | -1.702 | -1.6494 | 0.052669 |  | 4.7157 | -1.7093
17 | YES | 10.319 | 4.6757 | 0.001782 | -1.3213 | -1.7495 | -0.42826 |  | 4.77 | -1.762
17 | YES | 5.1317 | 4.7157 | 0.0056502 | -1.7093 | -1.7639 | -0.05452 |  | 1.8728 | -0.74222
17 | YES | 2.2104 | 4.77 | 0.0019683 | -1.762 | -1.7834 | -0.021413 |  | 1.8729 | -0.72805
18 | YES | 312.193 | 1.8728 | 0.0021506 | -0.74222 | -0.74235 | -0.00012766 |  | 1.9048 | -0.74515
18 | YES | 300.322 | 1.8729 | 0.001244 | -0.72805 | -0.74241 | -0.014362 |  | 1.9462 | -0.81876
18 | YES | 250.544 | 1.9048 | 0.0019003 | -0.74515 | -0.75387 | -0.0087251 |  | 1.8698 | -0.82617
18 | YES | 200.207 | 1.9462 | 0.001049 | -0.81876 | -0.76875 | 0.050014 |  | 1.8704 | -0.72261
18 | YES | 175.174 | 1.8698 | 0.0017449 | -0.82617 | -0.74129 | 0.084883 |  | 1.8573 | -0.67167
18 | YES | 150.243 | 1.8704 | 0.0018725 | -0.72261 | -0.74149 | -0.01888 |  | 1.9255 | -1.1105
18 | YES | 124.303 | 1.8573 | 0.0011059 | -0.67167 | -0.7368 | -0.06513 |  | 2.6254 | -1.1766
18 | YES | 99.7041 | 1.9255 | 0.0031388 | -1.1105 | -0.76129 | 0.34921 |  | 3.2384 | -1.4046
18 | YES | 74.792 | 2.6254 | 0.034142 | -1.1766 | -1.0128 | 0.16382 |  | 3.5309 | -1.4981
18 | YES | 50.1137 | 3.2384 | 0.032568 | -1.4046 | -1.233 | 0.17161 |  | 3.6134 | -1.5236
18 | YES | 39.9031 | 3.5309 | 0.012688 | -1.4981 | -1.3382 | 0.1599 |  | 3.9144 | -1.6106
18 | YES | 30.4567 | 3.6134 | 0.0034662 | -1.5236 | -1.3678 | 0.15575 |  | 4.1413 | -1.5227
18 | YES | 19.8381 | 3.9144 | 0.017966 | -1.6106 | -1.4759 | 0.13465 |  | 4.178 | -1.528
18 | YES | 9.9176 | 4.1413 | 0.0043618 | -1.5227 | -1.5575 | -0.034725 |  | 4.2067 | -1.5553
18 | YES | 4.8786 | 4.178 | 0.0015186 | -1.528 | -1.5707 | -0.042721 |  | 1.916 | -0.78
18 | YES | 1.9734 | 4.2067 | 0.0022344 | -1.5553 | -1.581 | -0.025734 |  | 1.9098 | -0.6772
19 | YES | 145.098 | 1.916 | 0 | -0.78 | -0.75788 | 0.022119 |  | 1.9011 | -0.66893
19 | YES | 124.892 | 1.9098 | 0.00074856 | -0.6772 | -0.75565 | -0.078455 |  | 1.9028 | -0.66621
19 | YES | 100.175 | 1.9011 | 0.00051211 | -0.66893 | -0.75252 | -0.083582 |  | 2.7889 | -1.3011
19 | YES | 74.2929 | 1.9028 | 0.0014996 | -0.66621 | -0.75313 | -0.086927 |  | 2.9675 | -1.1415
19 | YES | 50.1467 | 2.7889 | 0.027121 | -1.3011 | -1.0715 | 0.22959 |  | 3.2236 | -1.2884
19 | YES | 40.2307 | 2.9675 | 0.0028229 | -1.1415 | -1.1357 | 0.0057923 |  | 3.2764 | -1.1676
19 | YES | 30.1963 | 3.2236 | 0.011798 | -1.2884 | -1.2277 | 0.060629 |  | 3.3764 | -1.2956
19 | YES | 20.2351 | 3.2764 | 0.0030684 | -1.1676 | -1.2467 | -0.079157 |  | 3.5024 | -1.3126
19 | YES | 9.7635 | 3.3764 | 0.013102 | -1.2956 | -1.2826 | 0.012993 |  | 3.5423 | -1.2637
19 | YES | 4.6938 | 3.5024 | 0.0049005 | -1.3126 | -1.3279 | -0.015363 |  | 3.2565 | -1.1755
19 | YES | 1.5803 | 3.5423 | 0.0034886 | -1.2637 | -1.3422 | -0.078519 |  | 3.249 | -1.186
20 | YES | 231.357 | 3.2565 | 0.0019577 | -1.1755 | -1.2396 | -0.06412 |  | 3.249 | -1.092
20 | YES | 199.967 | 3.249 | 0.0012974 | -1.186 | -1.2369 | -0.050884 |  | 3.2455 | -1.1885
20 | YES | 174.934 | 3.249 | 0.0012682 | -1.092 | -1.2369 | -0.14484 |  | 3.2443 | -1.6057
20 | YES | 149.814 | 3.2455 | 0.001119 | -1.1885 | -1.2356 | -0.047099 |  | 3.2461 | -1.1909
20 | YES | 125.263 | 3.2443 | 0.0014919 | -1.6057 | -1.2352 | 0.37048 |  | 3.2424 | -1.1426
20 | YES | 99.9825 | 3.2461 | 0.0019845 | -1.1909 | -1.2358 | -0.044969 |  | 3.2401 | -1.3199
20 | YES | 75.3592 | 3.2424 | 0.0021689 | -1.1426 | -1.2345 | -0.091858 |  | 3.2424 | -1.1456
20 | YES | 50.0965 | 3.2401 | 0.00089973 | -1.3199 | -1.2337 | 0.086214 |  | 3.2437 | -1.0232
20 | YES | 40.1804 | 3.2424 | 0.001479 | -1.1456 | -1.2345 | -0.088858 |  | 3.2427 | -1.2513
20 | YES | 29.974 | 3.2437 | 0.0018542 | -1.0232 | -1.235 | -0.21172 |  | 3.2448 | -1.1512
20 | YES | 20.0329 | 3.2427 | 0.0015187 | -1.2513 | -1.2346 | 0.016748 |  | 3.242 | -1.173
20 | YES | 9.9679 | 3.2448 | 0.0016259 | -1.1512 | -1.2353 | -0.084148 |  | 3.2467 | -1.1843
20 | YES | 5.1428 | 3.242 | 0.0016532 | -1.173 | -1.2343 | -0.061369
20 | YES | 2.0859 | 3.2467 | 0.0015492 | -1.1843 | -1.236 | -0.051771
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX-CTD - DOX-BOT) (3)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.00064128 |  | X | CALC Y
m= | -0.0015929 | 0.0038627 |  | Std Dev of Resid |  | 0.34782 |  | 1 | -1.1021
b= | -1.1005 | 0.047498 |  | Resid Sum of Squares |  | 32.0586 |  | 20 | -1.1323
# of Values Extracted = |  |  | 269
# of Values Used in Fit = |  |  | 267
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 110.209 | 0.14646 | 0.0023558 | -0.59954 | -1.1021 | -0.50252 |  | 1 | -0.053025
1 | YES | 100.191 | 0.16698 | 0.00027273 | -0.053025 | -1.1021 | -1.049 |  | 1 | -0.58691
1 | YES | 75.0704 | 2.6501 | 0.0053135 | -0.58691 | -1.1021 | -0.51515 |  | 1 | -1.2407
1 | YES | 1.5835 | 3.6213 | 0.073599 | -1.2407 | -1.1021 | 0.13868 |  | 2 | -0.95754
2 | YES | 231.111 | 2.5645 | 0.0012848 | -0.95754 | -1.1037 | -0.14612 |  | 2 | -1.0462
2 | YES | 200.09 | 2.6258 | 0.0028774 | -1.0462 | -1.1037 | -0.057427 |  | 2 | -1.1205
2 | YES | 175.687 | 2.6535 | 0.001119 | -1.1205 | -1.1037 | 0.016843 |  | 2 | -1.3028
2 | YES | 150.335 | 2.8042 | 0.0040114 | -1.3028 | -1.1037 | 0.19912 |  | 2 | -1.1616
2 | YES | 125.097 | 2.8934 | 0.0035549 | -1.1616 | -1.1037 | 0.057963 |  | 2 | -1.2356
2 | YES | 99.9256 | 3.0364 | 0.0029657 | -1.2356 | -1.1037 | 0.13199 |  | 2 | -1.2787
2 | YES | 74.4687 | 3.3423 | 0.0074156 | -1.2787 | -1.1037 | 0.17505 |  | 2 | -1.3215
2 | YES | 49.1603 | 3.4055 | 0.0020856 | -1.3215 | -1.1037 | 0.21779 |  | 2 | -1.3115
2 | YES | 40.6798 | 3.4705 | 0.0019233 | -1.3115 | -1.1037 | 0.20789 |  | 2 | -1.2646
2 | YES | 29.7398 | 3.5254 | 0.0016215 | -1.2646 | -1.1037 | 0.16092 |  | 2 | -1.3259
2 | YES | 19.6317 | 3.5421 | 0.00039364 | -1.3259 | -1.1037 | 0.22228 |  | 2 | -1.2771
2 | YES | 9.6901 | 3.5499 | 0.0014419 | -1.2771 | -1.1037 | 0.17341 |  | 2 | -1.2555
2 | YES | 4.9828 | 3.5604 | 0.001072 | -1.2555 | -1.1037 | 0.15189 |  | 2 | -1.3018
2 | YES | 1.6667 | 3.5762 | 0.0024279 | -1.3018 | -1.1037 | 0.19817 |  | 3 | -0.58065
3 | YES | 244.526 | 1.4684 | 0.001215 | -0.58065 | -1.1052 | -0.5246 |  | 3 | -0.81563
3 | YES | 244.611 | 1.4684 | 0.0012711 | -0.81563 | -1.1052 | -0.28962 |  | 3 | -0.60664
3 | YES | 203.398 | 1.4764 | 0.0019011 | -0.60664 | -1.1052 | -0.4986 |  | 3 | -0.68709
3 | YES | 175.87 | 1.6429 | 0.0012649 | -0.68709 | -1.1052 | -0.41816 |  | 3 | -1.0215
3 | YES | 149.282 | 1.7275 | 0.0018305 | -1.0215 | -1.1052 | -0.08371 |  | 3 | -0.81862
3 | YES | 125.206 | 2.1854 | 0.0023953 | -0.81862 | -1.1052 | -0.28663 |  | 3 | -0.98561
3 | YES | 99.6255 | 2.3844 | 0.0015021 | -0.98561 | -1.1052 | -0.11964 |  | 3 | -1.0281
3 | YES | 75.1051 | 2.5879 | 0.0012538 | -1.0281 | -1.1052 | -0.07711 |  | 3 | -1.2957
3 | YES | 50.4953 | 3.2923 | 0.0013106 | -1.2957 | -1.1052 | 0.19041 |  | 3 | -1.5615
3 | YES | 41.1572 | 3.7275 | 0.0018261 | -1.5615 | -1.1052 | 0.45628 |  | 3 | -1.709
3 | YES | 29.3588 | 4.41 | 0.0036958 | -1.709 | -1.1052 | 0.60376 |  | 3 | -1.6802
3 | YES | 20.5425 | 4.5408 | 0.0016749 | -1.6802 | -1.1052 | 0.57494 |  | 3 | -1.6583
3 | YES | 10.4123 | 4.5747 | 0.0010919 | -1.6583 | -1.1052 | 0.55301 |  | 3 | -1.6551
3 | YES | 5.0724 | 4.5799 | 0.0015894 | -1.6551 | -1.1052 | 0.54983 |  | 3 | -1.7904
3 | YES | 1.8669 | 4.5846 | 0.0023584 | -1.7904 | -1.1052 | 0.68516 |  | 4 | -0.54429
4 | YES | 221.623 | 1.4957 | 0.001313 | -0.54429 | -1.1068 | -0.56255 |  | 4 | -0.55505
4 | YES | 201.791 | 1.4889 | 0.00038409 | -0.55505 | -1.1068 | -0.55179 |  | 4 | -0.65709
4 | YES | 172.797 | 1.5309 | 0.0023629 | -0.65709 | -1.1068 | -0.44975 |  | 4 | -0.72293
4 | YES | 150.467 | 1.7571 | 0.0038257 | -0.72293 | -1.1068 | -0.38392 |  | 4 | -1.0374
4 | YES | 124.65 | 1.9126 | 0.0019911 | -1.0374 | -1.1068 | -0.069483 |  | 4 | -0.90424
4 | YES | 99.9203 | 2.1878 | 0.024804 | -0.90424 | -1.1068 | -0.2026 |  | 4 | -1.091
4 | YES | 75.1964 | 2.68 | 0.01257 | -1.091 | -1.1068 | -0.015843 |  | 4 | -1.1775
4 | YES | 50.3963 | 2.9895 | 0.005905 | -1.1775 | -1.1068 | 0.070637 |  | 4 | -1.1277
4 | YES | 40.4829 | 2.9873 | 0.00256 | -1.1277 | -1.1068 | 0.020877 |  | 4 | -1.1636
4 | YES | 29.9081 | 3.2054 | 0.021531 | -1.1636 | -1.1068 | 0.056737 |  | 4 | -1.3863
4 | YES | 20.177 | 3.6917 | 0.016997 | -1.3863 | -1.1068 | 0.27942 |  | 4 | -1.8326
4 | YES | 9.7949 | 4.1154 | 0.032814 | -1.8326 | -1.1068 | 0.72575 |  | 4 | -1.6496
4 | YES | 5.2267 | 4.4224 | 0.010058 | -1.6496 | -1.1068 | 0.5428 |  | 4 | -1.3901
4 | YES | 2.1258 | 4.4949 | 0.00088504 | -1.3901 | -1.1068 | 0.28325 |  | 5 | -0.64399
5 | YES | 184.313 | 1.551 | 0.0014171 | -0.64399 | -1.1084 | -0.46444 |  | 5 | -0.51441
5 | YES | 175.011 | 1.5436 | 0.0014294 | -0.51441 | -1.1084 | -0.59403 |  | 5 | -0.50614
5 | YES | 149.623 | 1.5529 | 0.002364 | -0.50614 | -1.1084 | -0.6023 |  | 5 | -0.59293
5 | YES | 124.73 | 1.6841 | 0.0020007 | -0.59293 | -1.1084 | -0.51551 |  | 5 | -0.67674
5 | YES | 99.845 | 1.9403 | 0.0018642 | -0.67674 | -1.1084 | -0.43169 |  | 5 | -0.91923
5 | YES | 74.951 | 2.3818 | 0.0014246 | -0.91923 | -1.1084 | -0.1892 |  | 5 | -0.93303
5 | YES | 50.2894 | 2.845 | 0.0020933 | -0.93303 | -1.1084 | -0.1754 |  | 5 | -1.1514
5 | YES | 40.0385 | 3.1446 | 0.0018432 | -1.1514 | -1.1084 | 0.042995 |  | 5 | -1.2485
5 | YES | 30.0153 | 3.6185 | 0.0021177 | -1.2485 | -1.1084 | 0.14003 |  | 5 | -1.3606
5 | YES | 19.4498 | 4.3044 | 0.0035562 | -1.3606 | -1.1084 | 0.25213 |  | 5 | -1.7584
5 | YES | 9.701 | 4.3526 | 0.002138 | -1.7584 | -1.1084 | 0.64994 |  | 5 | -1.6149
5 | YES | 4.9995 | 4.3721 | 0.0018055 | -1.6149 | -1.1084 | 0.50644 |  | 5 | -1.6004
5 | YES | 1.5612 | 4.3726 | 0.0019213 | -1.6004 | -1.1084 | 0.49196 |  | 6 | -0.56283
6 | YES | 171.173 | 1.6122 | 0.001202 | -0.56283 | -1.11 | -0.5472 |  | 6 | -0.57172
6 | YES | 149.783 | 1.6273 | 0.0015823 | -0.57172 | -1.11 | -0.53831 |  | 6 | -0.59684
6 | YES | 124.912 | 1.6392 | 0.0018545 | -0.59684 | -1.11 | -0.51319 |  | 6 | -0.8592
6 | YES | 98.8446 | 1.8678 | 0.0042556 | -0.8592 | -1.11 | -0.25083 |  | 6 | -0.8727
6 | YES | 74.2203 | 2.0953 | 0.0018423 | -0.8727 | -1.11 | -0.23733 |  | 6 | -0.92179
6 | YES | 49.47 | 2.4122 | 0.0020652 | -0.92179 | -1.11 | -0.18824 |  | 6 | -1.0238
6 | YES | 39.8415 | 2.5562 | 0.0072749 | -1.0238 | -1.11 | -0.086248 |  | 6 | -1.3838
6 | YES | 32.8962 | 2.9242 | 0.0096326 | -1.3838 | -1.11 | 0.27376 |  | 6 | -1.2837
6 | YES | 22.9881 | 3.1593 | 0.0019928 | -1.2837 | -1.11 | 0.17372 |  | 6 | -1.3501
6 | YES | 13.0632 | 3.2449 | 0.0019118 | -1.3501 | -1.11 | 0.24003 |  | 6 | -1.4829
6 | YES | 4.2473 | 3.3821 | 0.0014905 | -1.4829 | -1.11 | 0.37283 |  | 6 | -1.2927
6 | YES | 1.983 | 3.3723 | 0.0024434 | -1.2927 | -1.11 | 0.18265 |  | 7 | -0.862
7 | YES | 77.2534 | 2.323 | 0 | -0.862 | -1.1116 | -0.24962 |  | 7 | -0.81963
7 | YES | 75.1932 | 2.4244 | 0.01458 | -0.81963 | -1.1116 | -0.29199 |  | 7 | -1.0201
7 | YES | 49.5689 | 2.7389 | 0.016676 | -1.0201 | -1.1116 | -0.091501 |  | 7 | -1.1619
7 | YES | 39.9511 | 2.9771 | 0.0073387 | -1.1619 | -1.1116 | 0.050259 |  | 7 | -1.2746
7 | YES | 30.3125 | 3.2204 | 0.017383 | -1.2746 | -1.1116 | 0.16301 |  | 7 | -1.2694
7 | YES | 20.0896 | 3.4166 | 0.0022552 | -1.2694 | -1.1116 | 0.1578 |  | 7 | -1.247
7 | YES | 10.9783 | 3.4661 | 0.00075419 | -1.247 | -1.1116 | 0.13533 |  | 7 | -1.3381
7 | YES | 2.8781 | 3.5019 | 0.0094609 | -1.3381 | -1.1116 | 0.22646 |  | 7 | -1.2791
7 | YES | 2.8781 | 3.5019 | 0.0094609 | -1.2791 | -1.1116 | 0.16746 |  | 8 | -1.1024
8 | YES | 210.266 | 2.7516 | 0.0015037 | -1.1024 | -1.1132 | -0.010804 |  | 8 | -1.0105
8 | YES | 200.335 | 2.7605 | 0.0011837 | -1.0105 | -1.1132 | -0.10268 |  | 8 | -0.98719
8 | YES | 174.961 | 2.7558 | 0.0019929 | -0.98719 | -1.1132 | -0.12602 |  | 8 | -1.0285
8 | YES | 150.142 | 2.8096 | 0.001072 | -1.0285 | -1.1132 | -0.084764 |  | 8 | -1.1309
8 | YES | 124.952 | 2.9861 | 0.00079902 | -1.1309 | -1.1132 | 0.017646 |  | 8 | -1.1517
8 | YES | 99.7344 | 3.1033 | 0.0014975 | -1.1517 | -1.1132 | 0.038446 |  | 8 | -1.2314
8 | YES | 75.0419 | 3.2596 | 0.0038877 | -1.2314 | -1.1132 | 0.11815 |  | 8 | -1.2553
8 | YES | 40.413 | 3.4167 | 0.0019968 | -1.2553 | -1.1132 | 0.14207 |  | 8 | -1.2545
8 | YES | 30.5348 | 3.3345 | 0.0049397 | -1.2545 | -1.1132 | 0.14128 |  | 8 | -1.4965
8 | YES | 30.5348 | 3.3345 | 0.0049397 | -1.4965 | -1.1132 | 0.38328 |  | 8 | -1.1716
8 | YES | 20.275 | 3.2694 | 0.0015147 | -1.1716 | -1.1132 | 0.058376 |  | 8 | -1.2191
8 | YES | 9.9533 | 3.3589 | 0.0022456 | -1.2191 | -1.1132 | 0.10591 |  | 8 | -1.2875
8 | YES | 5.2407 | 3.4405 | 0.0022065 | -1.2875 | -1.1132 | 0.1743 |  | 8 | -1.2528
8 | YES | 2.3193 | 3.4542 | 0.0013743 | -1.2528 | -1.1132 | 0.1396 |  | 9 | -1.1299
9 | YES | 175.577 | 2.9001 | 0.0019947 | -1.1299 | -1.1148 | 0.015123 |  | 9 | -1.3167
9 | YES | 150.277 | 2.9603 | 0.0010539 | -1.3167 | -1.1148 | 0.20189 |  | 9 | -1.0143
9 | YES | 125.206 | 3.1137 | 0.0016188 | -1.0143 | -1.1148 | -0.10049 |  | 9 | -1.1507
9 | YES | 99.9993 | 3.1503 | 0.0018896 | -1.1507 | -1.1148 | 0.035873 |  | 9 | -1.5692
9 | YES | 74.5362 | 3.2088 | 0.001315 | -1.5692 | -1.1148 | 0.45442 |  | 9 | -1.1387
9 | YES | 49.9354 | 3.2253 | 0.0013127 | -1.1387 | -1.1148 | 0.023863 |  | 9 | -1.1436
9 | YES | 40.2109 | 2.8404 | 0.0030207 | -1.1436 | -1.1148 | 0.028793 |  | 9 | -0.91897
9 | YES | 30.0634 | 2.777 | 0.0011723 | -0.91897 | -1.1148 | -0.19583 |  | 9 | -0.66167
9 | YES | 19.8283 | 2.9123 | 0.0021962 | -0.66167 | -1.1148 | -0.45314 |  | 9 | -1.2153
9 | YES | 9.7016 | 3.6487 | 0.0041504 | -1.2153 | -1.1148 | 0.10045 |  | 9 | -1.4303
9 | YES | 4.6399 | 3.7187 | 0.0014812 | -1.4303 | -1.1148 | 0.31549 |  | 9 | -1.5512
9 | YES | 1.5897 | 3.7189 | 0.00065398 | -1.5512 | -1.1148 | 0.43634 |  | 10 | -0.83371
10 | YES | 310.053 | 2.1853 | 0.0014913 | -0.83371 | -1.1164 | -0.28269 |  | 10 | -0.75734
10 | YES | 300.001 | 2.1067 | 0.0036596 | -0.75734 | -1.1164 | -0.35906 |  | 10 | -0.86203
10 | YES | 249.985 | 2.043 | 0.00082121 | -0.86203 | -1.1164 | -0.25437 |  | 10 | -0.87897
10 | YES | 200.039 | 2.114 | 0.0017535 | -0.87897 | -1.1164 | -0.23742 |  | 10 | -0.82663
10 | YES | 175.027 | 2.1614 | 0.00099275 | -0.82663 | -1.1164 | -0.28977 |  | 10 | -0.84039
10 | YES | 150.02 | 2.2426 | 0.0015021 | -0.84039 | -1.1164 | -0.27601 |  | 10 | -0.88602
10 | YES | 124.967 | 2.468 | 0.0014259 | -0.88602 | -1.1164 | -0.23038 |  | 10 | -0.96526
10 | YES | 100.176 | 2.6637 | 0.0021125 | -0.96526 | -1.1164 | -0.15114 |  | 10 | -1.2354
10 | YES | 75.1188 | 2.9876 | 0.0015307 | -1.2354 | -1.1164 | 0.119 |  | 10 | -1.1975
10 | YES | 50.0965 | 2.8675 | 0.0016284 | -1.1975 | -1.1164 | 0.08112 |  | 10 | -1.0343
10 | YES | 40.2836 | 2.8467 | 0.00094862 | -1.0343 | -1.1164 | -0.08213 |  | 10 | -1.1856
10 | YES | 30.0779 | 2.8484 | 0.0018941 | -1.1856 | -1.1164 | 0.06922 |  | 10 | -1.0735
10 | YES | 20.1112 | 2.9955 | 0.0017892 | -1.0735 | -1.1164 | -0.04293 |  | 10 | -1.6114
10 | YES | 10.1999 | 4.3846 | 0.0017672 | -1.6114 | -1.1164 | 0.49499 |  | 10 | -1.9867
10 | YES | 5.2541 | 4.4924 | 0.001174 | -1.9867 | -1.1164 | 0.87025 |  | 10 | -1.7292
10 | YES | 1.6808 | 4.5088 | 0.0016263 | -1.7292 | -1.1164 | 0.61277 |  | 11 | -0.76655
11 | YES | 372.67 | 1.8615 | 0.0013598 | -0.76655 | -1.118 | -0.35144 |  | 11 | -0.71092
11 | YES | 349.872 | 1.9021 | 0.0073457 | -0.71092 | -1.118 | -0.40708 |  | 11 | -0.7919
11 | YES | 299.899 | 1.9921 | 0.00062458 | -0.7919 | -1.118 | -0.32609 |  | 11 | -0.78428
11 | YES | 249.773 | 2.0347 | 0.0019374 | -0.78428 | -1.118 | -0.33371 |  | 11 | -0.79413
11 | YES | 199.882 | 2.1149 | 0.00071811 | -0.79413 | -1.118 | -0.32386 |  | 11 | -1.0308
11 | YES | 174.692 | 2.1962 | 0.0013709 | -1.0308 | -1.118 | -0.087223 |  | 11 | -0.93454
11 | YES | 150.058 | 2.2815 | 0.0030442 | -0.93454 | -1.118 | -0.18346 |  | 11 | -1.1144
11 | YES | 124.78 | 2.5196 | 0.0033291 | -1.1144 | -1.118 | -0.0036326 |  | 11 | -0.84642
11 | YES | 99.9088 | 2.3226 | 0.00233 | -0.84642 | -1.118 | -0.27157 |  | 11 | -1.1737
11 | YES | 74.8 | 2.3213 | 0.003395 | -1.1737 | -1.118 | 0.055667 |  | 11 | -0.94079
11 | YES | 49.9199 | 2.4142 | 0.010427 | -0.94079 | -1.118 | -0.1772 |  | 11 | -1.0502
11 | YES | 40.1629 | 2.6878 | 0.026679 | -1.0502 | -1.118 | -0.067803 |  | 11 | -1.3928
11 | YES | 19.961 | 3.5582 | 0.032358 | -1.3928 | -1.118 | 0.27481 |  | 11 | -1.7684
11 | YES | 9.9586 | 4.1226 | 0.073385 | -1.7684 | -1.118 | 0.65045 |  | 11 | -1.6147
11 | YES | 5.2056 | 4.3373 | 0.014591 | -1.6147 | -1.118 | 0.49668 |  | 11 | -1.6705
11 | YES | 1.7664 | 4.5755 | 0.0059289 | -1.6705 | -1.118 | 0.55253 |  | 12 | -0.74529
12 | YES | 340.021 | 1.9607 | 0.0018773 | -0.74529 | -1.1196 | -0.3743 |  | 12 | -0.72101
12 | YES | 299.463 | 2.017 | 0.0014171 | -0.72101 | -1.1196 | -0.39858 |  | 12 | -0.81095
12 | YES | 250.243 | 2.0541 | 0.00038409 | -0.81095 | -1.1196 | -0.30863 |  | 12 | -0.77797
12 | YES | 199.683 | 2.152 | 0.0017534 | -0.77797 | -1.1196 | -0.34161 |  | 12 | -0.79903
12 | YES | 174.672 | 2.264 | 0.0014079 | -0.79903 | -1.1196 | -0.32055 |  | 12 | -0.84325
12 | YES | 149.647 | 2.3908 | 0.0019161 | -0.84325 | -1.1196 | -0.27634 |  | 12 | -0.82095
12 | YES | 124.71 | 2.3441 | 0.00054543 | -0.82095 | -1.1196 | -0.29864 |  | 12 | -0.93299
12 | YES | 99.9992 | 2.574 | 0.00045636 | -0.93299 | -1.1196 | -0.18659 |  | 12 | -0.91317
12 | YES | 74.821 | 2.5228 | 0.0018288 | -0.91317 | -1.1196 | -0.20641 |  | 12 | -0.78807
12 | YES | 50.0241 | 2.2669 | 0.00065492 | -0.78807 | -1.1196 | -0.33152 |  | 12 | -0.92683
12 | YES | 40.106 | 2.4772 | 0.0042766 | -0.92683 | -1.1196 | -0.19275 |  | 12 | -1.0121
12 | YES | 30.1896 | 2.8579 | 0.0015253 | -1.0121 | -1.1196 | -0.10751 |  | 12 | -1.4359
12 | YES | 20.0473 | 3.5491 | 0.0027415 | -1.4359 | -1.1196 | 0.31628 |  | 12 | -1.5361
12 | YES | 9.9047 | 4.4419 | 0.0037537 | -1.5361 | -1.1196 | 0.41651 |  | 12 | -1.647
12 | YES | 5.0873 | 4.572 | 0.0019142 | -1.647 | -1.1196 | 0.52746 |  | 12 | -1.6963
12 | YES | 1.8965 | 4.5837 | 0.00087346 | -1.6963 | -1.1196 | 0.57667 |  | 13 | -0.8494
13 | YES | 275.184 | 1.9426 | 0.0011141 | -0.8494 | -1.1212 | -0.27178 |  | 13 | -0.69889
13 | YES | 250.05 | 1.9541 | 0.001413 | -0.69889 | -1.1212 | -0.42229 |  | 13 | -0.69654
13 | YES | 200.053 | 2.0225 | 0.0019793 | -0.69654 | -1.1212 | -0.42464 |  | 13 | -0.78212
13 | YES | 174.464 | 2.0639 | 0.0018328 | -0.78212 | -1.1212 | -0.33905 |  | 13 | -1.1339
13 | YES | 150.006 | 2.1951 | 0.0022581 | -1.1339 | -1.1212 | 0.012692 |  | 13 | -0.857
13 | YES | 124.876 | 2.346 | 0.0024427 | -0.857 | -1.1212 | -0.26418 |  | 13 | -0.95692
13 | YES | 100.109 | 2.4571 | 0.0027706 | -0.95692 | -1.1212 | -0.16426 |  | 13 | -0.84271
13 | YES | 74.8535 | 2.4643 | 0.0021033 | -0.84271 | -1.1212 | -0.27847 |  | 13 | -0.66783
13 | YES | 50.1091 | 2.5152 | 0.0024923 | -0.66783 | -1.1212 | -0.45335 |  | 13 | -0.92455
13 | YES | 40.206 | 2.7454 | 0.01255 | -0.92455 | -1.1212 | -0.19662 |  | 13 | -1.1811
13 | YES | 30.1437 | 2.9619 | 0.010664 | -1.1811 | -1.1212 | 0.059942 |  | 13 | -1.5093
13 | YES | 20.203 | 3.4817 | 0.082422 | -1.5093 | -1.1212 | 0.38812 |  | 13 | -1.8821
13 | YES | 10.0564 | 4.4209 | 0.022527 | -1.8821 | -1.1212 | 0.76093 |  | 13 | -1.8568
13 | YES | 5.0958 | 4.5602 | 0.0061208 | -1.8568 | -1.1212 | 0.73561 |  | 13 | -1.7627
13 | YES | 1.9085 | 4.5853 | 0.0018394 | -1.7627 | -1.1212 | 0.64151 |  | 14 | -0.7483
14 | YES | 316.832 | 1.8967 | 0.0023439 | -0.7483 | -1.1228 | -0.37447 |  | 14 | -0.70599
14 | YES | 300.016 | 1.917 | 0.0017344 | -0.70599 | -1.1228 | -0.41678 |  | 14 | -0.9191
14 | YES | 250.279 | 1.9599 | 0.0005386 | -0.9191 | -1.1228 | -0.20367 |  | 14 | -0.73608
14 | YES | 199.862 | 2.0679 | 0.001688 | -0.73608 | -1.1228 | -0.38669 |  | 14 | -0.77931
14 | YES | 175.252 | 2.2057 | 0.0012705 | -0.77931 | -1.1228 | -0.34347 |  | 14 | -0.80469
14 | YES | 150.242 | 2.2903 | 0.0063138 | -0.80469 | -1.1228 | -0.31808 |  | 14 | -0.86928
14 | YES | 124.737 | 2.4217 | 0.0057012 | -0.86928 | -1.1228 | -0.25349 |  | 14 | -0.84585
14 | YES | 100.157 | 2.3931 | 0.00065398 | -0.84585 | -1.1228 | -0.27692 |  | 14 | -0.79336
14 | YES | 75.1673 | 2.3126 | 0.0014999 | -0.79336 | -1.1228 | -0.32941 |  | 14 | -0.98908
14 | YES | 49.8643 | 2.5959 | 0.0017681 | -0.98908 | -1.1228 | -0.13369 |  | 14 | -1.1643
14 | YES | 40.2218 | 3.3327 | 0.005416 | -1.1643 | -1.1228 | 0.041499 |  | 14 | -1.3892
14 | YES | 29.8998 | 3.8398 | 0.0024037 | -1.3892 | -1.1228 | 0.2664 |  | 14 | -1.5396
14 | YES | 20.2758 | 4.0304 | 0.0042531 | -1.5396 | -1.1228 | 0.41684 |  | 14 | -1.3884
14 | YES | 10.1761 | 4.3705 | 0.0012247 | -1.3884 | -1.1228 | 0.26568 |  | 14 | -1.5966
14 | YES | 5.3946 | 4.4354 | 0.0026729 | -1.5966 | -1.1228 | 0.47382 |  | 14 | -1.594
14 | YES | 2.1456 | 4.513 | 0.0017744 | -1.594 | -1.1228 | 0.47126 |  | 15 | -0.91527
15 | YES | 188.75 | 2.0777 | 0.00086603 | -0.91527 | -1.1244 | -0.20909 |  | 15 | -0.79469
15 | YES | 174.696 | 2.1083 | 0.0021441 | -0.79469 | -1.1244 | -0.32967 |  | 15 | -0.77049
15 | YES | 150.187 | 2.1915 | 0.0028229 | -0.77049 | -1.1244 | -0.35388 |  | 15 | -0.83222
15 | YES | 125.052 | 2.3058 | 0.00085135 | -0.83222 | -1.1244 | -0.29214 |  | 15 | -0.80521
15 | YES | 100.088 | 2.2588 | 0.0022506 | -0.80521 | -1.1244 | -0.31916 |  | 15 | -0.79574
15 | YES | 74.9715 | 2.1533 | 0.0024101 | -0.79574 | -1.1244 | -0.32862 |  | 15 | -0.94049
15 | YES | 49.9686 | 2.5965 | 0.00769 | -0.94049 | -1.1244 | -0.18388 |  | 15 | -1.5717
15 | YES | 39.9753 | 3.2743 | 0.029715 | -1.5717 | -1.1244 | 0.44733 |  | 15 | -1.4408
15 | YES | 30.2688 | 3.7542 | 0.0054148 | -1.4408 | -1.1244 | 0.31643 |  | 15 | -1.4453
15 | YES | 20.1448 | 3.8327 | 0.0044506 | -1.4453 | -1.1244 | 0.32093 |  | 15 | -1.5984
15 | YES | 10.0866 | 4.2636 | 0.0069756 | -1.5984 | -1.1244 | 0.47404 |  | 15 | -1.5879
15 | YES | 5.1861 | 4.3391 | 0.0037943 | -1.5879 | -1.1244 | 0.46353 |  | 15 | -1.6261
15 | YES | 1.9275 | 4.4039 | 0.0042631 | -1.6261 | -1.1244 | 0.50174 |  | 16 | -0.84145
16 | YES | 165.084 | 1.8476 | 0.0011619 | -0.84145 | -1.126 | -0.2845 |  | 16 | -0.72895
16 | YES | 150.005 | 1.872 | 0.0084979 | -0.72895 | -1.126 | -0.39701 |  | 16 | -0.74898
16 | YES | 124.87 | 1.977 | 0.0014943 | -0.74898 | -1.126 | -0.37697 |  | 16 | -0.75646
16 | YES | 100.004 | 1.9685 | 0.0079057 | -0.75646 | -1.126 | -0.36949 |  | 16 | -0.97544
16 | YES | 74.8941 | 2.0526 | 0.0021598 | -0.97544 | -1.126 | -0.15052 |  | 16 | -0.97513
16 | YES | 50.3432 | 2.4929 | 0.0081024 | -0.97513 | -1.126 | -0.15082 |  | 16 | -1.3681
16 | YES | 40.1958 | 2.9779 | 0.10133 | -1.3681 | -1.126 | 0.24211 |  | 16 | -1.5512
16 | YES | 30.075 | 3.7348 | 0.066807 | -1.5512 | -1.126 | 0.42525 |  | 16 | -1.7787
16 | YES | 20.0364 | 4.3003 | 0.033634 | -1.7787 | -1.126 | 0.65271 |  | 16 | -1.722
16 | YES | 10.1725 | 4.445 | 0.0027095 | -1.722 | -1.126 | 0.59603 |  | 16 | -1.6707
16 | YES | 4.9316 | 4.4583 | 0.0050078 | -1.6707 | -1.126 | 0.54474 |  | 16 | -1.6293
16 | YES | 1.686 | 4.4707 | 0.0018532 | -1.6293 | -1.126 | 0.50335 |  | 17 | -0.7863
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7863 | -1.1275 | -0.34125 |  | 17 | -0.7813
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7813 | -1.1275 | -0.34625 |  | 17 | -0.61523
17 | YES | 300.413 | 1.9158 | 0.00093775 | -0.61523 | -1.1275 | -0.51232 |  | 17 | -0.72518
17 | YES | 249.829 | 1.9758 | 0.00099161 | -0.72518 | -1.1275 | -0.40237 |  | 17 | -1.0292
17 | YES | 200.363 | 2.0128 | 0.00092755 | -1.0292 | -1.1275 | -0.09835 |  | 17 | -0.7752
17 | YES | 175.126 | 2.1298 | 0.00074856 | -0.7752 | -1.1275 | -0.35235 |  | 17 | -0.96417
17 | YES | 149.927 | 2.1218 | 0.0007033 | -0.96417 | -1.1275 | -0.16338 |  | 17 | -0.90002
17 | YES | 124.734 | 2.186 | 0.00061185 | -0.90002 | -1.1275 | -0.22752 |  | 17 | -0.70816
17 | YES | 99.8058 | 1.9368 | 0.0018664 | -0.70816 | -1.1275 | -0.41939 |  | 17 | -0.75636
17 | YES | 74.8425 | 2.2816 | 0.0015 | -0.75636 | -1.1275 | -0.37119 |  | 17 | -1.7333
17 | YES | 50.1359 | 4.1537 | 0.0094609 | -1.7333 | -1.1275 | 0.60579 |  | 17 | -1.6338
17 | YES | 40.3131 | 4.3472 | 0.003735 | -1.6338 | -1.1275 | 0.50629 |  | 17 | -1.702
17 | NO | 30.2081 | 4.4076 | 0.0017444 | -4.4854 | -1.1275 | 3.3579 |  | 17 | -1.3213
17 | YES | 20.1101 | 4.397 | 0.0017841 | -1.702 | -1.1275 | 0.57447 |  | 17 | -1.7093
17 | YES | 10.319 | 4.6757 | 0.001782 | -1.3213 | -1.1275 | 0.19371 |  | 17 | -1.762
17 | YES | 5.1317 | 4.7157 | 0.0056502 | -1.7093 | -1.1275 | 0.58179 |  | 18 | -0.74222
17 | YES | 2.2104 | 4.77 | 0.0019683 | -1.762 | -1.1275 | 0.63443 |  | 18 | -0.72805
18 | YES | 312.193 | 1.8728 | 0.0021506 | -0.74222 | -1.1291 | -0.38692 |  | 18 | -0.74515
18 | YES | 300.322 | 1.8729 | 0.001244 | -0.72805 | -1.1291 | -0.40109 |  | 18 | -0.81876
18 | YES | 250.544 | 1.9048 | 0.0019003 | -0.74515 | -1.1291 | -0.38399 |  | 18 | -0.82617
18 | YES | 200.207 | 1.9462 | 0.001049 | -0.81876 | -1.1291 | -0.31038 |  | 18 | -0.72261
18 | YES | 175.174 | 1.8698 | 0.0017449 | -0.82617 | -1.1291 | -0.30297 |  | 18 | -0.67167
18 | YES | 150.243 | 1.8704 | 0.0018725 | -0.72261 | -1.1291 | -0.40653 |  | 18 | -1.1105
18 | YES | 124.303 | 1.8573 | 0.0011059 | -0.67167 | -1.1291 | -0.45747 |  | 18 | -1.1766
18 | YES | 99.7041 | 1.9255 | 0.0031388 | -1.1105 | -1.1291 | -0.018643 |  | 18 | -1.4046
18 | YES | 74.792 | 2.6254 | 0.034142 | -1.1766 | -1.1291 | 0.047457 |  | 18 | -1.4981
18 | YES | 50.1137 | 3.2384 | 0.032568 | -1.4046 | -1.1291 | 0.2755 |  | 18 | -1.5236
18 | YES | 39.9031 | 3.5309 | 0.012688 | -1.4981 | -1.1291 | 0.36892 |  | 18 | -1.6106
18 | YES | 30.4567 | 3.6134 | 0.0034662 | -1.5236 | -1.1291 | 0.39442 |  | 18 | -1.5227
18 | YES | 19.8381 | 3.9144 | 0.017966 | -1.6106 | -1.1291 | 0.48146 |  | 18 | -1.528
18 | YES | 9.9176 | 4.1413 | 0.0043618 | -1.5227 | -1.1291 | 0.3936 |  | 18 | -1.5553
18 | YES | 4.8786 | 4.178 | 0.0015186 | -1.528 | -1.1291 | 0.39882 |  | 19 | -0.78
18 | YES | 1.9734 | 4.2067 | 0.0022344 | -1.5553 | -1.1291 | 0.42612 |  | 19 | -0.6772
19 | YES | 145.098 | 1.916 | 0 | -0.78 | -1.1307 | -0.35074 |  | 19 | -0.66893
19 | YES | 124.892 | 1.9098 | 0.00074856 | -0.6772 | -1.1307 | -0.45354 |  | 19 | -0.66621
19 | YES | 100.175 | 1.9011 | 0.00051211 | -0.66893 | -1.1307 | -0.4618 |  | 19 | -1.3011
19 | YES | 74.2929 | 1.9028 | 0.0014996 | -0.66621 | -1.1307 | -0.46453 |  | 19 | -1.1415
19 | YES | 50.1467 | 2.7889 | 0.027121 | -1.3011 | -1.1307 | 0.17038 |  | 19 | -1.2884
19 | YES | 40.2307 | 2.9675 | 0.0028229 | -1.1415 | -1.1307 | 0.010764 |  | 19 | -1.1676
19 | YES | 30.1963 | 3.2236 | 0.011798 | -1.2884 | -1.1307 | 0.15763 |  | 19 | -1.2956
19 | YES | 20.2351 | 3.2764 | 0.0030684 | -1.1676 | -1.1307 | 0.036824 |  | 19 | -1.3126
19 | YES | 9.7635 | 3.3764 | 0.013102 | -1.2956 | -1.1307 | 0.16488 |  | 19 | -1.2637
19 | YES | 4.6938 | 3.5024 | 0.0049005 | -1.3126 | -1.1307 | 0.18182 |  | 20 | -1.1755
19 | YES | 1.5803 | 3.5423 | 0.0034886 | -1.2637 | -1.1307 | 0.13298 |  | 20 | -1.186
20 | YES | 231.357 | 3.2565 | 0.0019577 | -1.1755 | -1.1323 | 0.043122 |  | 20 | -1.092
20 | YES | 199.967 | 3.249 | 0.0012974 | -1.186 | -1.1323 | 0.053652 |  | 20 | -1.1885
20 | YES | 174.934 | 3.249 | 0.0012682 | -1.092 | -1.1323 | -0.040318 |  | 20 | -1.6057
20 | YES | 149.814 | 3.2455 | 0.001119 | -1.1885 | -1.1323 | 0.056172 |  | 20 | -1.1909
20 | YES | 125.263 | 3.2443 | 0.0014919 | -1.6057 | -1.1323 | 0.47333 |  | 20 | -1.1426
20 | YES | 99.9825 | 3.2461 | 0.0019845 | -1.1909 | -1.1323 | 0.058532 |  | 20 | -1.3199
20 | YES | 75.3592 | 3.2424 | 0.0021689 | -1.1426 | -1.1323 | 0.010292 |  | 20 | -1.1456
20 | YES | 50.0965 | 3.2401 | 0.00089973 | -1.3199 | -1.1323 | 0.18755 |  | 20 | -1.0232
20 | YES | 40.1804 | 3.2424 | 0.001479 | -1.1456 | -1.1323 | 0.013292 |  | 20 | -1.2513
20 | YES | 29.974 | 3.2437 | 0.0018542 | -1.0232 | -1.1323 | -0.10908 |  | 20 | -1.1512
20 | YES | 20.0329 | 3.2427 | 0.0015187 | -1.2513 | -1.1323 | 0.119 |  | 20 | -1.173
20 | YES | 9.9679 | 3.2448 | 0.0016259 | -1.1512 | -1.1323 | 0.018872 |  | 20 | -1.1843
20 | YES | 5.1428 | 3.242 | 0.0016532 | -1.173 | -1.1323 | 0.040652
20 | YES | 2.0859 | 3.2467 | 0.0015492 | -1.1843 | -1.1323 | 0.051942
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX-CTD - DOX-BOT) (4)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=DOX-CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.88791 |  | X | CALC Y
m= | -0.33014 | 0.03135 |  | Std Dev of Resid |  | 0.15283 |  | 1.8207 | -0.74484
b= | -0.14375 | 0.10017 |  | Resid Sum of Squares |  | 0.327 |  | 4.77 | -1.7185
# of Values Extracted = |  |  | 17
# of Values Used in Fit = |  |  | 16
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | DOX-CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7863 | -0.74484 | 0.041457 |  | 1.8207 | -0.7863
17 | YES | 350.547 | 1.8207 | 0.0012724 | -0.7813 | -0.74484 | 0.036457 |  | 1.8207 | -0.7813
17 | YES | 300.413 | 1.9158 | 0.00093775 | -0.61523 | -0.77623 | -0.161 |  | 1.9158 | -0.61523
17 | YES | 249.829 | 1.9758 | 0.00099161 | -0.72518 | -0.79605 | -0.070872 |  | 1.9758 | -0.72518
17 | YES | 200.363 | 2.0128 | 0.00092755 | -1.0292 | -0.80826 | 0.22094 |  | 2.0128 | -1.0292
17 | YES | 175.126 | 2.1298 | 0.00074856 | -0.7752 | -0.84689 | -0.071691 |  | 2.1298 | -0.7752
17 | YES | 149.927 | 2.1218 | 0.0007033 | -0.96417 | -0.84426 | 0.11992 |  | 2.1218 | -0.96417
17 | YES | 124.734 | 2.186 | 0.00061185 | -0.90002 | -0.86544 | 0.034589 |  | 2.186 | -0.90002
17 | YES | 99.8058 | 1.9368 | 0.0018664 | -0.70816 | -0.78318 | -0.075028 |  | 1.9368 | -0.70816
17 | YES | 74.8425 | 2.2816 | 0.0015 | -0.75636 | -0.89702 | -0.14065 |  | 2.2816 | -0.75636
17 | YES | 50.1359 | 4.1537 | 0.0094609 | -1.7333 | -1.515 | 0.21829 |  | 4.1537 | -1.7333
17 | YES | 40.3131 | 4.3472 | 0.003735 | -1.6338 | -1.5789 | 0.054908 |  | 4.3472 | -1.6338
17 | NO | 30.2081 | 4.4076 | 0.0017444 | -4.4854 | -1.5989 | 2.8865 |  | 4.397 | -1.702
17 | YES | 20.1101 | 4.397 | 0.0017841 | -1.702 | -1.5954 | 0.10664 |  | 4.6757 | -1.3213
17 | YES | 10.319 | 4.6757 | 0.001782 | -1.3213 | -1.6874 | -0.36615 |  | 4.7157 | -1.7093
17 | YES | 5.1317 | 4.7157 | 0.0056502 | -1.7093 | -1.7006 | 0.0087512 |  | 4.77 | -1.762
17 | YES | 2.2104 | 4.77 | 0.0019683 | -1.762 | -1.7185 | 0.043445
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX-BOT | DOX-CTD |  | DOX-CTD - DOX-BOT
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX-BOT | DOX-CTD | DOX-CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2004-37-0001 | 2004-37-0001 | 121 | 110.209 | 110.209 | 0.033192 | -0.00047302 | 0.746 | 0.14646 | 0.0023558 | -0.59954
2004-37-0001 | 2004-37-0001 | 121 | 100.191 | 100.191 | 0.031799 | 0.0003891 | 0.22 | 0.16698 | 0.00027273 | -0.053025
2004-37-0001 | 2004-37-0001 | 121 | 75.07 | 75.0704 | 0.032888 | 0.00037384 | 3.237 | 2.6501 | 0.0053135 | -0.58691
2004-37-0001 | 2004-37-0001 | 121 | 1.584 | 1.5835 | 0.041537 | -0.00046277 | 4.862 | 3.6213 | 0.073599 | -1.2407
2004-37-0002 | 2004-37-0002 | 121 | 231.111 | 231.111 | 0.091313 | 9.1553e-05 | 3.522 | 2.5645 | 0.0012848 | -0.95754
2004-37-0002 | 2004-37-0002 | 121 | 200.09 | 200.09 | 0.11028 | -0.00022888 | 3.672 | 2.6258 | 0.0028774 | -1.0462
2004-37-0002 | 2004-37-0002 | 121 | 175.687 | 175.687 | 0.14308 | -0.00028992 | 3.774 | 2.6535 | 0.001119 | -1.1205
2004-37-0002 | 2004-37-0002 | 121 | 150.335 | 150.335 | 0.046535 | 0.00041199 | 4.107 | 2.8042 | 0.0040114 | -1.3028
2004-37-0002 | 2004-37-0002 | 121 | 125.097 | 125.097 | 0.077909 | 0.00020599 | 4.055 | 2.8934 | 0.0035549 | -1.1616
2004-37-0002 | 2004-37-0002 | 121 | 99.926 | 99.9256 | 0.12079 | -0.0003891 | 4.272 | 3.0364 | 0.0029657 | -1.2356
2004-37-0002 | 2004-37-0002 | 121 | 74.469 | 74.4687 | 0.10592 | -0.00028992 | 4.621 | 3.3423 | 0.0074156 | -1.2787
2004-37-0002 | 2004-37-0002 | 121 | 49.16 | 49.1603 | 0.14873 | 0.00027466 | 4.727 | 3.4055 | 0.0020856 | -1.3215
2004-37-0002 | 2004-37-0002 | 121 | 40.68 | 40.6798 | 0.15796 | -0.00022507 | 4.782 | 3.4705 | 0.0019233 | -1.3115
2004-37-0002 | 2004-37-0002 | 121 | 29.74 | 29.7398 | 0.15619 | -0.00024796 | 4.79 | 3.5254 | 0.0016215 | -1.2646
2004-37-0002 | 2004-37-0002 | 121 | 19.632 | 19.6317 | 0.081225 | -0.00032997 | 4.868 | 3.5421 | 0.00039364 | -1.3259
2004-37-0002 | 2004-37-0002 | 121 | 9.69 | 9.6901 | 0.050976 | 0.00012398 | 4.827 | 3.5499 | 0.0014419 | -1.2771
2004-37-0002 | 2004-37-0002 | 121 | 4.983 | 4.9828 | 0.091058 | -0.00018168 | 4.816 | 3.5604 | 0.001072 | -1.2555
2004-37-0002 | 2004-37-0002 | 121 | 1.667 | 1.6667 | 0.09663 | -0.00027275 | 4.878 | 3.5762 | 0.0024279 | -1.3018
2004-37-0006 | 2004-37-0006 | 232 | 244.277 | 244.526 | 0.41922 | 0.24901 | 2.049 | 1.4684 | 0.001215 | -0.58065
2004-37-0006 | 2004-37-0006 | 227 | 244.841 | 244.611 | 0.40547 | -0.23027 | 2.284 | 1.4684 | 0.0012711 | -0.81563
2004-37-0006 | 2004-37-0006 | 121 | 203.398 | 203.398 | 0.16284 | 3.0518e-05 | 2.083 | 1.4764 | 0.0019011 | -0.60664
2004-37-0006 | 2004-37-0006 | 121 | 175.87 | 175.87 | 0.52865 | -0.00010681 | 2.33 | 1.6429 | 0.0012649 | -0.68709
2004-37-0006 | 2004-37-0006 | 121 | 149.282 | 149.282 | 0.084323 | -0.00012207 | 2.749 | 1.7275 | 0.0018305 | -1.0215
2004-37-0006 | 2004-37-0006 | 121 | 125.206 | 125.206 | 0.54219 | 0.00045013 | 3.004 | 2.1854 | 0.0023953 | -0.81862
2004-37-0006 | 2004-37-0006 | 121 | 99.625 | 99.6255 | 0.25296 | 0.00047302 | 3.37 | 2.3844 | 0.0015021 | -0.98561
2004-37-0006 | 2004-37-0006 | 121 | 75.105 | 75.1051 | 0.62073 | 0.00010681 | 3.616 | 2.5879 | 0.0012538 | -1.0281
2004-37-0006 | 2004-37-0006 | 99 | 50.236 | 50.4953 | 0.65779 | 0.25926 | 4.588 | 3.2923 | 0.0013106 | -1.2957
2004-37-0006 | 2004-37-0006 | 121 | 41.157 | 41.1572 | 0.22189 | 0.00021362 | 5.289 | 3.7275 | 0.0018261 | -1.5615
2004-37-0006 | 2004-37-0006 | 121 | 29.359 | 29.3588 | 0.22153 | -0.00020599 | 6.119 | 4.41 | 0.0036958 | -1.709
2004-37-0006 | 2004-37-0006 | 121 | 20.542 | 20.5425 | 0.33018 | 0.00047874 | 6.221 | 4.5408 | 0.0016749 | -1.6802
2004-37-0006 | 2004-37-0006 | 121 | 10.412 | 10.4123 | 0.16118 | 0.00025654 | 6.233 | 4.5747 | 0.0010919 | -1.6583
2004-37-0006 | 2004-37-0006 | 121 | 5.072 | 5.0724 | 0.39748 | 0.00037193 | 6.235 | 4.5799 | 0.0015894 | -1.6551
2004-37-0006 | 2004-37-0006 | 121 | 1.867 | 1.8669 | 0.25435 | -7.4387e-05 | 6.375 | 4.5846 | 0.0023584 | -1.7904
2004-37-0009 | 2004-37-0009 | 121 | 221.623 | 221.623 | 0.16604 | -0.00038147 | 2.04 | 1.4957 | 0.001313 | -0.54429
2004-37-0009 | 2004-37-0009 | 121 | 201.791 | 201.791 | 0.19714 | 0.00048828 | 2.044 | 1.4889 | 0.00038409 | -0.55505
2004-37-0009 | 2004-37-0009 | 121 | 172.797 | 172.797 | 0.36851 | 6.1035e-05 | 2.188 | 1.5309 | 0.0023629 | -0.65709
2004-37-0009 | 2004-37-0009 | 121 | 150.467 | 150.467 | 0.28835 | 0.00030518 | 2.48 | 1.7571 | 0.0038257 | -0.72293
2004-37-0009 | 2004-37-0009 | 121 | 124.65 | 124.65 | 0.21453 | -0.00024414 | 2.95 | 1.9126 | 0.0019911 | -1.0374
2004-37-0009 | 2004-37-0009 | 121 | 99.92 | 99.9203 | 0.45028 | 0.00033569 | 3.092 | 2.1878 | 0.024804 | -0.90424
2004-37-0009 | 2004-37-0009 | 121 | 75.196 | 75.1964 | 0.090183 | 0.00036621 | 3.771 | 2.68 | 0.01257 | -1.091
2004-37-0009 | 2004-37-0009 | 121 | 50.396 | 50.3963 | 0.13584 | 0.00030518 | 4.167 | 2.9895 | 0.005905 | -1.1775
2004-37-0009 | 2004-37-0009 | 121 | 40.483 | 40.4829 | 0.32895 | -0.00014114 | 4.115 | 2.9873 | 0.00256 | -1.1277
2004-37-0009 | 2004-37-0009 | 121 | 29.908 | 29.9081 | 0.22786 | 0.00012398 | 4.369 | 3.2054 | 0.021531 | -1.1636
2004-37-0009 | 2004-37-0009 | 121 | 20.177 | 20.177 | 0.13059 | -3.2425e-05 | 5.078 | 3.6917 | 0.016997 | -1.3863
2004-37-0009 | 2004-37-0009 | 121 | 9.795 | 9.7949 | 0.22183 | -0.00010776 | 5.948 | 4.1154 | 0.032814 | -1.8326
2004-37-0009 | 2004-37-0009 | 121 | 5.227 | 5.2267 | 0.20322 | -0.00028944 | 6.072 | 4.4224 | 0.010058 | -1.6496
2004-37-0009 | 2004-37-0009 | 121 | 2.126 | 2.1258 | 0.053365 | -0.00016522 | 5.885 | 4.4949 | 0.00088504 | -1.3901
2004-37-0012 | 2004-37-0012 | 121 | 184.313 | 184.313 | 0.15212 | -0.00038147 | 2.195 | 1.551 | 0.0014171 | -0.64399
2004-37-0012 | 2004-37-0012 | 121 | 175.011 | 175.011 | 0.15659 | -0.00018311 | 2.058 | 1.5436 | 0.0014294 | -0.51441
2004-37-0012 | 2004-37-0012 | 121 | 149.623 | 149.623 | 0.11428 | -1.5259e-05 | 2.059 | 1.5529 | 0.002364 | -0.50614
2004-37-0012 | 2004-37-0012 | 121 | 124.73 | 124.73 | 0.2462 | 9.9182e-05 | 2.277 | 1.6841 | 0.0020007 | -0.59293
2004-37-0012 | 2004-37-0012 | 121 | 99.845 | 99.845 | 0.073991 | 1.5259e-05 | 2.617 | 1.9403 | 0.0018642 | -0.67674
2004-37-0012 | 2004-37-0012 | 121 | 74.951 | 74.951 | 0.1695 | 4.5776e-05 | 3.301 | 2.3818 | 0.0014246 | -0.91923
2004-37-0012 | 2004-37-0012 | 121 | 50.289 | 50.2894 | 0.11557 | 0.00039673 | 3.778 | 2.845 | 0.0020933 | -0.93303
2004-37-0012 | 2004-37-0012 | 121 | 40.038 | 40.0385 | 0.19671 | 0.00048065 | 4.296 | 3.1446 | 0.0018432 | -1.1514
2004-37-0012 | 2004-37-0012 | 121 | 30.015 | 30.0153 | 0.15407 | 0.00032234 | 4.867 | 3.6185 | 0.0021177 | -1.2485
2004-37-0012 | 2004-37-0012 | 121 | 19.45 | 19.4498 | 0.067043 | -0.0002327 | 5.665 | 4.3044 | 0.0035562 | -1.3606
2004-37-0012 | 2004-37-0012 | 121 | 9.701 | 9.701 | 0.041786 | -1.7166e-05 | 6.111 | 4.3526 | 0.002138 | -1.7584
2004-37-0012 | 2004-37-0012 | 121 | 5 | 4.9995 | 0.042434 | -0.00049591 | 5.987 | 4.3721 | 0.0018055 | -1.6149
2004-37-0012 | 2004-37-0012 | 121 | 1.561 | 1.5612 | 0.059847 | 0.00023138 | 5.973 | 4.3726 | 0.0019213 | -1.6004
2004-37-0015 | 2004-37-0015 | 121 | 171.173 | 171.173 | 0.2669 | -0.00027466 | 2.175 | 1.6122 | 0.001202 | -0.56283
2004-37-0015 | 2004-37-0015 | 121 | 149.783 | 149.783 | 0.052145 | 0.00039673 | 2.199 | 1.6273 | 0.0015823 | -0.57172
2004-37-0015 | 2004-37-0015 | 121 | 124.912 | 124.912 | 0.11476 | -0.0001297 | 2.236 | 1.6392 | 0.0018545 | -0.59684
2004-37-0015 | 2004-37-0015 | 121 | 98.845 | 98.8446 | 0.18849 | -0.00036621 | 2.727 | 1.8678 | 0.0042556 | -0.8592
2004-37-0015 | 2004-37-0015 | 121 | 74.22 | 74.2203 | 0.13609 | 0.00030518 | 2.968 | 2.0953 | 0.0018423 | -0.8727
2004-37-0015 | 2004-37-0015 | 121 | 49.47 | 49.47 | 0.20035 | -2.6703e-05 | 3.334 | 2.4122 | 0.0020652 | -0.92179
2004-37-0015 | 2004-37-0015 | 121 | 39.841 | 39.8415 | 0.11944 | 0.00048065 | 3.58 | 2.5562 | 0.0072749 | -1.0238
2004-37-0015 | 2004-37-0015 | 121 | 32.896 | 32.8962 | 0.23794 | 0.0002327 | 4.308 | 2.9242 | 0.0096326 | -1.3838
2004-37-0015 | 2004-37-0015 | 121 | 22.988 | 22.9881 | 0.13646 | 5.722e-05 | 4.443 | 3.1593 | 0.0019928 | -1.2837
2004-37-0015 | 2004-37-0015 | 121 | 13.063 | 13.0632 | 0.088728 | 0.00016594 | 4.595 | 3.2449 | 0.0019118 | -1.3501
2004-37-0015 | 2004-37-0015 | 121 | 4.247 | 4.2473 | 0.071249 | 0.00030565 | 4.865 | 3.3821 | 0.0014905 | -1.4829
2004-37-0015 | 2004-37-0015 | 121 | 1.983 | 1.983 | 0.10832 | -4.1366e-05 | 4.665 | 3.3723 | 0.0024434 | -1.2927
2004-37-0018 | 2004-37-0018 | 121 | 77.253 | 77.2534 | 0.21499 | 0.00043487 | 3.185 | 2.323 | 0 | -0.862
2004-37-0018 | 2004-37-0018 | 121 | 75.193 | 75.1932 | 0.32109 | 0.00016785 | 3.244 | 2.4244 | 0.01458 | -0.81963
2004-37-0018 | 2004-37-0018 | 121 | 49.569 | 49.5689 | 0.15183 | -0.00014877 | 3.759 | 2.7389 | 0.016676 | -1.0201
2004-37-0018 | 2004-37-0018 | 56 | 39.921 | 39.9511 | 0.60842 | 0.030086 | 4.139 | 2.9771 | 0.0073387 | -1.1619
2004-37-0018 | 2004-37-0018 | 121 | 30.313 | 30.3125 | 0.15851 | -0.00047874 | 4.495 | 3.2204 | 0.017383 | -1.2746
2004-37-0018 | 2004-37-0018 | 113 | 20.015 | 20.0896 | 0.43066 | 0.074558 | 4.686 | 3.4166 | 0.0022552 | -1.2694
2004-37-0018 | 2004-37-0018 | 78 | 10.82 | 10.9783 | 0.5708 | 0.1583 | 4.713 | 3.4661 | 0.00075419 | -1.247
2004-37-0018 | 2004-37-0018 | 242 | 2.963 | 2.8781 | 0.26105 | -0.084872 | 4.84 | 3.5019 | 0.0094609 | -1.3381
2004-37-0018 | 2004-37-0018 | 242 | 2.793 | 2.8781 | 0.26105 | 0.085128 | 4.781 | 3.5019 | 0.0094609 | -1.2791
2004-37-0023 | 2004-37-0023 | 121 | 210.266 | 210.266 | 0.12467 | 0.00027466 | 3.854 | 2.7516 | 0.0015037 | -1.1024
2004-37-0023 | 2004-37-0023 | 121 | 200.335 | 200.335 | 0.22634 | -0.00022888 | 3.771 | 2.7605 | 0.0011837 | -1.0105
2004-37-0023 | 2004-37-0023 | 121 | 174.961 | 174.961 | 0.11813 | -0.00030518 | 3.743 | 2.7558 | 0.0019929 | -0.98719
2004-37-0023 | 2004-37-0023 | 121 | 150.142 | 150.142 | 0.13349 | 0.00016785 | 3.838 | 2.8096 | 0.001072 | -1.0285
2004-37-0023 | 2004-37-0023 | 121 | 124.952 | 124.952 | 0.16788 | 0.00019836 | 4.117 | 2.9861 | 0.00079902 | -1.1309
2004-37-0023 | 2004-37-0023 | 121 | 99.734 | 99.7344 | 0.19218 | 0.00041962 | 4.255 | 3.1033 | 0.0014975 | -1.1517
2004-37-0023 | 2004-37-0023 | 121 | 75.042 | 75.0419 | 0.085199 | -8.3923e-05 | 4.491 | 3.2596 | 0.0038877 | -1.2314
2004-37-0023 | 2004-37-0023 | 121 | 40.413 | 40.413 | 0.077654 | -3.0518e-05 | 4.672 | 3.4167 | 0.0019968 | -1.2553
2004-37-0023 | 2004-37-0023 | 242 | 30.545 | 30.5348 | 0.076088 | -0.01022 | 4.589 | 3.3345 | 0.0049397 | -1.2545
2004-37-0023 | 2004-37-0023 | 242 | 30.525 | 30.5348 | 0.076088 | 0.0097809 | 4.831 | 3.3345 | 0.0049397 | -1.4965
2004-37-0023 | 2004-37-0023 | 121 | 20.275 | 20.275 | 0.047285 | -7.6294e-06 | 4.441 | 3.2694 | 0.0015147 | -1.1716
2004-37-0023 | 2004-37-0023 | 121 | 9.953 | 9.9533 | 0.1388 | 0.00033855 | 4.578 | 3.3589 | 0.0022456 | -1.2191
2004-37-0023 | 2004-37-0023 | 121 | 5.241 | 5.2407 | 0.10654 | -0.00028133 | 4.728 | 3.4405 | 0.0022065 | -1.2875
2004-37-0023 | 2004-37-0023 | 121 | 2.319 | 2.3193 | 0.11462 | 0.00030589 | 4.707 | 3.4542 | 0.0013743 | -1.2528
2004-37-0032 | 2004-37-0032 | 121 | 175.577 | 175.577 | 0.037293 | -0.00027466 | 4.03 | 2.9001 | 0.0019947 | -1.1299
2004-37-0032 | 2004-37-0032 | 121 | 150.277 | 150.277 | 0.031574 | 0.00027466 | 4.277 | 2.9603 | 0.0010539 | -1.3167
2004-37-0032 | 2004-37-0032 | 121 | 125.206 | 125.206 | 0.04638 | -0.0001297 | 4.128 | 3.1137 | 0.0016188 | -1.0143
2004-37-0032 | 2004-37-0032 | 121 | 99.999 | 99.9993 | 0.035977 | 0.0003128 | 4.301 | 3.1503 | 0.0018896 | -1.1507
2004-37-0032 | 2004-37-0032 | 121 | 74.536 | 74.5362 | 0.036442 | 0.00024414 | 4.778 | 3.2088 | 0.001315 | -1.5692
2004-37-0032 | 2004-37-0032 | 121 | 49.935 | 49.9354 | 0.045554 | 0.00043869 | 4.364 | 3.2253 | 0.0013127 | -1.1387
2004-37-0032 | 2004-37-0032 | 121 | 40.211 | 40.2109 | 0.035173 | -6.485e-05 | 3.984 | 2.8404 | 0.0030207 | -1.1436
2004-37-0032 | 2004-37-0032 | 121 | 30.063 | 30.0634 | 0.037399 | 0.00044632 | 3.696 | 2.777 | 0.0011723 | -0.91897
2004-37-0032 | 2004-37-0032 | 121 | 19.828 | 19.8283 | 0.044305 | 0.00032234 | 3.574 | 2.9123 | 0.0021962 | -0.66167
2004-37-0032 | 2004-37-0032 | 121 | 9.702 | 9.7016 | 0.041738 | -0.00043774 | 4.864 | 3.6487 | 0.0041504 | -1.2153
2004-37-0032 | 2004-37-0032 | 121 | 4.64 | 4.6399 | 0.072597 | -0.00013208 | 5.149 | 3.7187 | 0.0014812 | -1.4303
2004-37-0032 | 2004-37-0032 | 121 | 1.59 | 1.5897 | 0.06171 | -0.00029755 | 5.27 | 3.7189 | 0.00065398 | -1.5512
2004-37-0034 | 2004-37-0034 | 121 | 310.053 | 310.053 | 0.057574 | 0.00030518 | 3.019 | 2.1853 | 0.0014913 | -0.83371
2004-37-0034 | 2004-37-0034 | 121 | 300.001 | 300.001 | 0.035623 | -3.0518e-05 | 2.864 | 2.1067 | 0.0036596 | -0.75734
2004-37-0034 | 2004-37-0034 | 121 | 249.985 | 249.985 | 0.077451 | -0.00038147 | 2.905 | 2.043 | 0.00082121 | -0.86203
2004-37-0034 | 2004-37-0034 | 121 | 200.039 | 200.039 | 0.079243 | 0.00032043 | 2.993 | 2.114 | 0.0017535 | -0.87897
2004-37-0034 | 2004-37-0034 | 121 | 175.027 | 175.027 | 0.059999 | -0.00039673 | 2.988 | 2.1614 | 0.00099275 | -0.82663
2004-37-0034 | 2004-37-0034 | 121 | 150.02 | 150.02 | 0.054285 | -0.00016785 | 3.083 | 2.2426 | 0.0015021 | -0.84039
2004-37-0034 | 2004-37-0034 | 121 | 124.967 | 124.967 | 0.045786 | -0.00041962 | 3.354 | 2.468 | 0.0014259 | -0.88602
2004-37-0034 | 2004-37-0034 | 121 | 100.176 | 100.176 | 0.041966 | 0.00022125 | 3.629 | 2.6637 | 0.0021125 | -0.96526
2004-37-0034 | 2004-37-0034 | 121 | 75.119 | 75.1188 | 0.048883 | -0.00024414 | 4.223 | 2.9876 | 0.0015307 | -1.2354
2004-37-0034 | 2004-37-0034 | 121 | 50.096 | 50.0965 | 0.057914 | 0.00045395 | 4.065 | 2.8675 | 0.0016284 | -1.1975
2004-37-0034 | 2004-37-0034 | 121 | 40.284 | 40.2836 | 0.074006 | -0.00041199 | 3.881 | 2.8467 | 0.00094862 | -1.0343
2004-37-0034 | 2004-37-0034 | 121 | 30.078 | 30.0779 | 0.043801 | -7.4387e-05 | 4.034 | 2.8484 | 0.0018941 | -1.1856
2004-37-0034 | 2004-37-0034 | 121 | 20.111 | 20.1112 | 0.038931 | 0.00016594 | 4.069 | 2.9955 | 0.0017892 | -1.0735
2004-37-0034 | 2004-37-0034 | 121 | 10.2 | 10.1999 | 0.037342 | -0.00013161 | 5.996 | 4.3846 | 0.0017672 | -1.6114
2004-37-0034 | 2004-37-0034 | 121 | 5.254 | 5.2541 | 0.050509 | 0.00014019 | 6.479 | 4.4924 | 0.001174 | -1.9867
2004-37-0034 | 2004-37-0034 | 121 | 1.681 | 1.6808 | 0.040036 | -0.00023138 | 6.238 | 4.5088 | 0.0016263 | -1.7292
2004-37-0039 | 2004-37-0039 | 121 | 372.67 | 372.67 | 0.04212 | 0.00036621 | 2.628 | 1.8615 | 0.0013598 | -0.76655
2004-37-0039 | 2004-37-0039 | 121 | 349.872 | 349.872 | 0.057227 | 0.00045776 | 2.613 | 1.9021 | 0.0073457 | -0.71092
2004-37-0039 | 2004-37-0039 | 121 | 299.899 | 299.899 | 0.054311 | -0.00015259 | 2.784 | 1.9921 | 0.00062458 | -0.7919
2004-37-0039 | 2004-37-0039 | 121 | 249.773 | 249.773 | 0.051537 | 0.00038147 | 2.819 | 2.0347 | 0.0019374 | -0.78428
2004-37-0039 | 2004-37-0039 | 121 | 199.882 | 199.882 | 0.039359 | -0.00039673 | 2.909 | 2.1149 | 0.00071811 | -0.79413
2004-37-0039 | 2004-37-0039 | 121 | 174.692 | 174.692 | 0.051473 | -1.5259e-05 | 3.227 | 2.1962 | 0.0013709 | -1.0308
2004-37-0039 | 2004-37-0039 | 121 | 150.058 | 150.058 | 0.045399 | 0.00048828 | 3.216 | 2.2815 | 0.0030442 | -0.93454
2004-37-0039 | 2004-37-0039 | 121 | 124.78 | 124.78 | 0.045925 | -0.00026703 | 3.634 | 2.5196 | 0.0033291 | -1.1144
2004-37-0039 | 2004-37-0039 | 121 | 99.909 | 99.9088 | 0.044429 | -0.00020599 | 3.169 | 2.3226 | 0.00233 | -0.84642
2004-37-0039 | 2004-37-0039 | 121 | 74.8 | 74.8 | 0.04206 | 1.5259e-05 | 3.495 | 2.3213 | 0.003395 | -1.1737
2004-37-0039 | 2004-37-0039 | 121 | 49.92 | 49.9199 | 0.055611 | -8.0109e-05 | 3.355 | 2.4142 | 0.010427 | -0.94079
2004-37-0039 | 2004-37-0039 | 121 | 40.163 | 40.1629 | 0.047174 | -0.00011444 | 3.738 | 2.6878 | 0.026679 | -1.0502
2004-37-0039 | 2004-37-0039 | 121 | 19.961 | 19.961 | 0.046367 | 7.6294e-06 | 4.951 | 3.5582 | 0.032358 | -1.3928
2004-37-0039 | 2004-37-0039 | 121 | 9.959 | 9.9586 | 0.044259 | -0.00042057 | 5.891 | 4.1226 | 0.073385 | -1.7684
2004-37-0039 | 2004-37-0039 | 121 | 5.206 | 5.2056 | 0.050991 | -0.00041294 | 5.952 | 4.3373 | 0.014591 | -1.6147
2004-37-0039 | 2004-37-0039 | 121 | 1.766 | 1.7664 | 0.075822 | 0.00035536 | 6.246 | 4.5755 | 0.0059289 | -1.6705
2004-37-0042 | 2004-37-0042 | 121 | 340.021 | 340.021 | 0.10793 | -0.00021362 | 2.706 | 1.9607 | 0.0018773 | -0.74529
2004-37-0042 | 2004-37-0042 | 121 | 299.463 | 299.463 | 0.093835 | 0.00030518 | 2.738 | 2.017 | 0.0014171 | -0.72101
2004-37-0042 | 2004-37-0042 | 121 | 250.243 | 250.243 | 0.051154 | 0.00012207 | 2.865 | 2.0541 | 0.00038409 | -0.81095
2004-37-0042 | 2004-37-0042 | 121 | 199.683 | 199.683 | 0.064742 | 0.00047302 | 2.93 | 2.152 | 0.0017534 | -0.77797
2004-37-0042 | 2004-37-0042 | 121 | 174.672 | 174.672 | 0.098589 | -0.00032043 | 3.063 | 2.264 | 0.0014079 | -0.79903
2004-37-0042 | 2004-37-0042 | 121 | 149.647 | 149.647 | 0.065016 | -0.00015259 | 3.234 | 2.3908 | 0.0019161 | -0.84325
2004-37-0042 | 2004-37-0042 | 121 | 124.71 | 124.71 | 0.092927 | 9.9182e-05 | 3.165 | 2.3441 | 0.00054543 | -0.82095
2004-37-0042 | 2004-37-0042 | 121 | 99.999 | 99.9992 | 0.064475 | 0.00022125 | 3.507 | 2.574 | 0.00045636 | -0.93299
2004-37-0042 | 2004-37-0042 | 121 | 74.821 | 74.821 | 0.046128 | -4.5776e-05 | 3.436 | 2.5228 | 0.0018288 | -0.91317
2004-37-0042 | 2004-37-0042 | 121 | 50.024 | 50.0241 | 0.048165 | 0.00013351 | 3.055 | 2.2669 | 0.00065492 | -0.78807
2004-37-0042 | 2004-37-0042 | 121 | 40.106 | 40.106 | 0.03861 | 4.9591e-05 | 3.404 | 2.4772 | 0.0042766 | -0.92683
2004-37-0042 | 2004-37-0042 | 121 | 30.19 | 30.1896 | 0.047896 | -0.00039673 | 3.87 | 2.8579 | 0.0015253 | -1.0121
2004-37-0042 | 2004-37-0042 | 121 | 20.047 | 20.0473 | 0.045031 | 0.0003376 | 4.985 | 3.5491 | 0.0027415 | -1.4359
2004-37-0042 | 2004-37-0042 | 121 | 9.905 | 9.9047 | 0.05464 | -0.00026417 | 5.978 | 4.4419 | 0.0037537 | -1.5361
2004-37-0042 | 2004-37-0042 | 121 | 5.087 | 5.0873 | 0.033979 | 0.00027275 | 6.219 | 4.572 | 0.0019142 | -1.647
2004-37-0042 | 2004-37-0042 | 121 | 1.896 | 1.8965 | 0.053797 | 0.00048757 | 6.28 | 4.5837 | 0.00087346 | -1.6963
2004-37-0044 | 2004-37-0044 | 121 | 275.184 | 275.184 | 0.036581 | -0.00033569 | 2.792 | 1.9426 | 0.0011141 | -0.8494
2004-37-0044 | 2004-37-0044 | 121 | 250.05 | 250.05 | 0.074986 | 6.1035e-05 | 2.653 | 1.9541 | 0.001413 | -0.69889
2004-37-0044 | 2004-37-0044 | 121 | 200.053 | 200.053 | 0.046686 | -0.00019836 | 2.719 | 2.0225 | 0.0019793 | -0.69654
2004-37-0044 | 2004-37-0044 | 121 | 174.464 | 174.464 | 0.037939 | 0.00045776 | 2.846 | 2.0639 | 0.0018328 | -0.78212
2004-37-0044 | 2004-37-0044 | 121 | 150.006 | 150.006 | 0.039567 | -0.0004425 | 3.329 | 2.1951 | 0.0022581 | -1.1339
2004-37-0044 | 2004-37-0044 | 121 | 124.876 | 124.876 | 0.043795 | 0.00039673 | 3.203 | 2.346 | 0.0024427 | -0.857
2004-37-0044 | 2004-37-0044 | 121 | 100.109 | 100.109 | 0.038927 | 0.00028229 | 3.414 | 2.4571 | 0.0027706 | -0.95692
2004-37-0044 | 2004-37-0044 | 121 | 74.853 | 74.8535 | 0.043937 | 0.00049591 | 3.307 | 2.4643 | 0.0021033 | -0.84271
2004-37-0044 | 2004-37-0044 | 121 | 50.109 | 50.1091 | 0.034978 | 6.485e-05 | 3.183 | 2.5152 | 0.0024923 | -0.66783
2004-37-0044 | 2004-37-0044 | 121 | 40.206 | 40.206 | 0.059858 | 7.6294e-06 | 3.67 | 2.7454 | 0.01255 | -0.92455
2004-37-0044 | 2004-37-0044 | 121 | 30.144 | 30.1437 | 0.036378 | -0.0003376 | 4.143 | 2.9619 | 0.010664 | -1.1811
2004-37-0044 | 2004-37-0044 | 121 | 20.203 | 20.203 | 0.050338 | 4.1962e-05 | 4.991 | 3.4817 | 0.082422 | -1.5093
2004-37-0044 | 2004-37-0044 | 121 | 10.056 | 10.0564 | 0.031726 | 0.00042152 | 6.303 | 4.4209 | 0.022527 | -1.8821
2004-37-0044 | 2004-37-0044 | 121 | 5.096 | 5.0958 | 0.047284 | -0.00018215 | 6.417 | 4.5602 | 0.0061208 | -1.8568
2004-37-0044 | 2004-37-0044 | 121 | 1.909 | 1.9085 | 0.081969 | -0.00047934 | 6.348 | 4.5853 | 0.0018394 | -1.7627
2004-37-0047 | 2004-37-0047 | 121 | 316.832 | 316.832 | 0.032909 | 0.00036621 | 2.645 | 1.8967 | 0.0023439 | -0.7483
2004-37-0047 | 2004-37-0047 | 121 | 300.016 | 300.016 | 0.031463 | -0.00030518 | 2.623 | 1.917 | 0.0017344 | -0.70599
2004-37-0047 | 2004-37-0047 | 121 | 250.279 | 250.279 | 0.048894 | 0.00042725 | 2.879 | 1.9599 | 0.0005386 | -0.9191
2004-37-0047 | 2004-37-0047 | 174 | 199.825 | 199.862 | 0.060269 | 0.037384 | 2.804 | 2.0679 | 0.001688 | -0.73608
2004-37-0047 | 2004-37-0047 | 121 | 175.252 | 175.252 | 0.046691 | 0.00019836 | 2.985 | 2.2057 | 0.0012705 | -0.77931
2004-37-0047 | 2004-37-0047 | 121 | 150.242 | 150.242 | 0.035113 | -0.00018311 | 3.095 | 2.2903 | 0.0063138 | -0.80469
2004-37-0047 | 2004-37-0047 | 121 | 124.737 | 124.737 | 0.030997 | 0.00022125 | 3.291 | 2.4217 | 0.0057012 | -0.86928
2004-37-0047 | 2004-37-0047 | 121 | 100.157 | 100.157 | 0.03864 | -0.00038147 | 3.239 | 2.3931 | 0.00065398 | -0.84585
2004-37-0047 | 2004-37-0047 | 121 | 75.167 | 75.1673 | 0.033499 | 0.00027466 | 3.106 | 2.3126 | 0.0014999 | -0.79336
2004-37-0047 | 2004-37-0047 | 121 | 49.864 | 49.8643 | 0.060866 | 0.00032425 | 3.585 | 2.5959 | 0.0017681 | -0.98908
2004-37-0047 | 2004-37-0047 | 121 | 40.222 | 40.2218 | 0.042267 | -0.00022507 | 4.497 | 3.3327 | 0.005416 | -1.1643
2004-37-0047 | 2004-37-0047 | 121 | 29.9 | 29.8998 | 0.034029 | -0.00023842 | 5.229 | 3.8398 | 0.0024037 | -1.3892
2004-37-0047 | 2004-37-0047 | 121 | 20.276 | 20.2758 | 0.033622 | -0.00024605 | 5.57 | 4.0304 | 0.0042531 | -1.5396
2004-37-0047 | 2004-37-0047 | 121 | 10.176 | 10.1761 | 0.041091 | 9.1553e-05 | 5.759 | 4.3705 | 0.0012247 | -1.3884
2004-37-0047 | 2004-37-0047 | 121 | 5.395 | 5.3946 | 0.040043 | -0.00040483 | 6.032 | 4.4354 | 0.0026729 | -1.5966
2004-37-0047 | 2004-37-0047 | 121 | 2.146 | 2.1456 | 0.085288 | -0.00043798 | 6.107 | 4.513 | 0.0017744 | -1.594
2004-37-0049 | 2004-37-0049 | 121 | 188.75 | 188.75 | 0.031107 | -0.00021362 | 2.993 | 2.0777 | 0.00086603 | -0.91527
2004-37-0049 | 2004-37-0049 | 121 | 174.696 | 174.696 | 0.036057 | -0.00013733 | 2.903 | 2.1083 | 0.0021441 | -0.79469
2004-37-0049 | 2004-37-0049 | 121 | 150.187 | 150.187 | 0.036757 | -0.00035095 | 2.962 | 2.1915 | 0.0028229 | -0.77049
2004-37-0049 | 2004-37-0049 | 121 | 125.052 | 125.052 | 0.031017 | 0.00048828 | 3.138 | 2.3058 | 0.00085135 | -0.83222
2004-37-0049 | 2004-37-0049 | 121 | 100.088 | 100.088 | 0.034134 | -0.00023651 | 3.064 | 2.2588 | 0.0022506 | -0.80521
2004-37-0049 | 2004-37-0049 | 121 | 74.972 | 74.9715 | 0.036497 | -0.00045776 | 2.949 | 2.1533 | 0.0024101 | -0.79574
2004-37-0049 | 2004-37-0049 | 121 | 49.969 | 49.9686 | 0.036958 | -0.00035477 | 3.537 | 2.5965 | 0.00769 | -0.94049
2004-37-0049 | 2004-37-0049 | 121 | 39.975 | 39.9753 | 0.034124 | 0.00030518 | 4.846 | 3.2743 | 0.029715 | -1.5717
2004-37-0049 | 2004-37-0049 | 121 | 30.269 | 30.2688 | 0.039407 | -0.00023842 | 5.195 | 3.7542 | 0.0054148 | -1.4408
2004-37-0049 | 2004-37-0049 | 121 | 20.145 | 20.1448 | 0.031031 | -0.00022316 | 5.278 | 3.8327 | 0.0044506 | -1.4453
2004-37-0049 | 2004-37-0049 | 121 | 10.087 | 10.0866 | 0.031693 | -0.00036335 | 5.862 | 4.2636 | 0.0069756 | -1.5984
2004-37-0049 | 2004-37-0049 | 121 | 5.186 | 5.1861 | 0.031583 | 0.00010776 | 5.927 | 4.3391 | 0.0037943 | -1.5879
2004-37-0049 | 2004-37-0049 | 121 | 1.927 | 1.9275 | 0.03842 | 0.00044632 | 6.03 | 4.4039 | 0.0042631 | -1.6261
2004-37-0052 | 2004-37-0052 | 121 | 165.084 | 165.084 | 0.051394 | 0.00018311 | 2.689 | 1.8476 | 0.0011619 | -0.84145
2004-37-0052 | 2004-37-0052 | 121 | 150.005 | 150.005 | 0.068708 | 0.00013733 | 2.601 | 1.872 | 0.0084979 | -0.72895
2004-37-0052 | 2004-37-0052 | 121 | 124.87 | 124.87 | 0.04958 | 0.00015259 | 2.726 | 1.977 | 0.0014943 | -0.74898
2004-37-0052 | 2004-37-0052 | 121 | 100.004 | 100.004 | 0.042386 | 0.00035095 | 2.725 | 1.9685 | 0.0079057 | -0.75646
2004-37-0052 | 2004-37-0052 | 121 | 74.894 | 74.8941 | 0.070257 | 0.00015259 | 3.028 | 2.0526 | 0.0021598 | -0.97544
2004-37-0052 | 2004-37-0052 | 121 | 50.343 | 50.3432 | 0.050383 | 0.00017548 | 3.468 | 2.4929 | 0.0081024 | -0.97513
2004-37-0052 | 2004-37-0052 | 121 | 40.196 | 40.1958 | 0.089476 | -0.0001564 | 4.346 | 2.9779 | 0.10133 | -1.3681
2004-37-0052 | 2004-37-0052 | 121 | 30.075 | 30.075 | 0.070103 | -4.1962e-05 | 5.286 | 3.7348 | 0.066807 | -1.5512
2004-37-0052 | 2004-37-0052 | 121 | 20.036 | 20.0364 | 0.059456 | 0.00037193 | 6.079 | 4.3003 | 0.033634 | -1.7787
2004-37-0052 | 2004-37-0052 | 121 | 10.172 | 10.1725 | 0.057234 | 0.00046253 | 6.167 | 4.445 | 0.0027095 | -1.722
2004-37-0052 | 2004-37-0052 | 121 | 4.932 | 4.9316 | 0.047378 | -0.00037193 | 6.129 | 4.4583 | 0.0050078 | -1.6707
2004-37-0052 | 2004-37-0052 | 121 | 1.686 | 1.686 | 0.087807 | 2.4796e-05 | 6.1 | 4.4707 | 0.0018532 | -1.6293
2004-37-0055 | 2004-37-0055 | 214 | 350.535 | 350.547 | 0.044601 | 0.011597 | 2.607 | 1.8207 | 0.0012724 | -0.7863
2004-37-0055 | 2004-37-0055 | 214 | 350.555 | 350.547 | 0.044601 | -0.0083923 | 2.602 | 1.8207 | 0.0012724 | -0.7813
2004-37-0055 | 2004-37-0055 | 121 | 300.413 | 300.413 | 0.092752 | -0.00027466 | 2.531 | 1.9158 | 0.00093775 | -0.61523
2004-37-0055 | 2004-37-0055 | 121 | 249.829 | 249.829 | 0.036081 | -0.00019836 | 2.701 | 1.9758 | 0.00099161 | -0.72518
2004-37-0055 | 2004-37-0055 | 121 | 200.363 | 200.363 | 0.069437 | 0.00033569 | 3.042 | 2.0128 | 0.00092755 | -1.0292
2004-37-0055 | 2004-37-0055 | 121 | 175.126 | 175.126 | 0.079502 | 0.00021362 | 2.905 | 2.1298 | 0.00074856 | -0.7752
2004-37-0055 | 2004-37-0055 | 121 | 149.927 | 149.927 | 0.1269 | 0.00047302 | 3.086 | 2.1218 | 0.0007033 | -0.96417
2004-37-0055 | 2004-37-0055 | 121 | 124.734 | 124.734 | 0.11505 | 7.6294e-05 | 3.086 | 2.186 | 0.00061185 | -0.90002
2004-37-0055 | 2004-37-0055 | 121 | 99.806 | 99.8058 | 0.11811 | -0.00020599 | 2.645 | 1.9368 | 0.0018664 | -0.70816
2004-37-0055 | 2004-37-0055 | 121 | 74.842 | 74.8425 | 0.043147 | 0.00045013 | 3.038 | 2.2816 | 0.0015 | -0.75636
2004-37-0055 | 2004-37-0055 | 121 | 50.136 | 50.1359 | 0.076562 | -9.9182e-05 | 5.887 | 4.1537 | 0.0094609 | -1.7333
2004-37-0055 | 2004-37-0055 | 121 | 40.313 | 40.3131 | 0.054058 | 8.3923e-05 | 5.981 | 4.3472 | 0.003735 | -1.6338
2004-37-0055 | 2004-37-0055 | 121 | 30.208 | 30.2081 | 0.064954 | 7.4387e-05 | 8.893 | 4.4076 | 0.0017444 | -4.4854
2004-37-0055 | 2004-37-0055 | 121 | 20.11 | 20.1101 | 0.06344 | 8.2016e-05 | 6.099 | 4.397 | 0.0017841 | -1.702
2004-37-0055 | 2004-37-0055 | 121 | 10.319 | 10.319 | 0.074176 | 3.2425e-05 | 5.997 | 4.6757 | 0.001782 | -1.3213
2004-37-0055 | 2004-37-0055 | 121 | 5.132 | 5.1317 | 0.070647 | -0.00029755 | 6.425 | 4.7157 | 0.0056502 | -1.7093
2004-37-0055 | 2004-37-0055 | 121 | 2.21 | 2.2104 | 0.053352 | 0.00041318 | 6.532 | 4.77 | 0.0019683 | -1.762
2004-37-0057 | 2004-37-0057 | 121 | 312.193 | 312.193 | 0.29332 | 9.1553e-05 | 2.615 | 1.8728 | 0.0021506 | -0.74222
2004-37-0057 | 2004-37-0057 | 121 | 300.322 | 300.322 | 0.37226 | 0.00030518 | 2.601 | 1.8729 | 0.001244 | -0.72805
2004-37-0057 | 2004-37-0057 | 121 | 250.544 | 250.544 | 0.51816 | -7.6294e-05 | 2.65 | 1.9048 | 0.0019003 | -0.74515
2004-37-0057 | 2004-37-0057 | 121 | 200.207 | 200.207 | 0.51014 | 0.00030518 | 2.765 | 1.9462 | 0.001049 | -0.81876
2004-37-0057 | 2004-37-0057 | 121 | 175.174 | 175.174 | 0.10306 | -6.1035e-05 | 2.696 | 1.8698 | 0.0017449 | -0.82617
2004-37-0057 | 2004-37-0057 | 121 | 150.243 | 150.243 | 0.15259 | 0.00047302 | 2.593 | 1.8704 | 0.0018725 | -0.72261
2004-37-0057 | 2004-37-0057 | 121 | 124.303 | 124.303 | 0.25841 | 0.00033569 | 2.529 | 1.8573 | 0.0011059 | -0.67167
2004-37-0057 | 2004-37-0057 | 121 | 99.704 | 99.7041 | 0.22924 | 0.00012207 | 3.036 | 1.9255 | 0.0031388 | -1.1105
2004-37-0057 | 2004-37-0057 | 121 | 74.792 | 74.792 | 0.090358 | 3.0518e-05 | 3.802 | 2.6254 | 0.034142 | -1.1766
2004-37-0057 | 2004-37-0057 | 121 | 50.114 | 50.1137 | 0.38928 | -0.00030518 | 4.643 | 3.2384 | 0.032568 | -1.4046
2004-37-0057 | 2004-37-0057 | 121 | 39.903 | 39.9031 | 0.21328 | 6.485e-05 | 5.029 | 3.5309 | 0.012688 | -1.4981
2004-37-0057 | 2004-37-0057 | 121 | 30.457 | 30.4567 | 0.29277 | -0.00032234 | 5.137 | 3.6134 | 0.0034662 | -1.5236
2004-37-0057 | 2004-37-0057 | 121 | 19.838 | 19.8381 | 0.53024 | 9.1553e-05 | 5.525 | 3.9144 | 0.017966 | -1.6106
2004-37-0057 | 2004-37-0057 | 121 | 9.918 | 9.9176 | 0.29984 | -0.0003891 | 5.664 | 4.1413 | 0.0043618 | -1.5227
2004-37-0057 | 2004-37-0057 | 121 | 4.879 | 4.8786 | 0.50794 | -0.00039673 | 5.706 | 4.178 | 0.0015186 | -1.528
2004-37-0057 | 2004-37-0057 | 121 | 1.973 | 1.9734 | 0.074921 | 0.00037181 | 5.762 | 4.2067 | 0.0022344 | -1.5553
2004-37-0059 | 2004-37-0059 | 121 | 145.098 | 145.098 | 0.34851 | -7.6294e-05 | 2.696 | 1.916 | 0 | -0.78
2004-37-0059 | 2004-37-0059 | 121 | 124.892 | 124.892 | 0.12766 | -0.00030518 | 2.587 | 1.9098 | 0.00074856 | -0.6772
2004-37-0059 | 2004-37-0059 | 121 | 100.175 | 100.175 | 0.11453 | 0.00047302 | 2.57 | 1.9011 | 0.00051211 | -0.66893
2004-37-0059 | 2004-37-0059 | 121 | 74.293 | 74.2929 | 0.477 | -0.00011444 | 2.569 | 1.9028 | 0.0014996 | -0.66621
2004-37-0059 | 2004-37-0059 | 121 | 50.147 | 50.1467 | 0.24787 | -0.00030518 | 4.09 | 2.7889 | 0.027121 | -1.3011
2004-37-0059 | 2004-37-0059 | 121 | 40.231 | 40.2307 | 0.21499 | -0.00028992 | 4.109 | 2.9675 | 0.0028229 | -1.1415
2004-37-0059 | 2004-37-0059 | 121 | 30.196 | 30.1963 | 0.23677 | 0.00030708 | 4.512 | 3.2236 | 0.011798 | -1.2884
2004-37-0059 | 2004-37-0059 | 121 | 20.235 | 20.2351 | 0.16689 | 0.00012398 | 4.444 | 3.2764 | 0.0030684 | -1.1676
2004-37-0059 | 2004-37-0059 | 121 | 9.764 | 9.7635 | 0.2505 | -0.00049591 | 4.672 | 3.3764 | 0.013102 | -1.2956
2004-37-0059 | 2004-37-0059 | 121 | 4.694 | 4.6938 | 0.1665 | -0.00019789 | 4.815 | 3.5024 | 0.0049005 | -1.3126
2004-37-0059 | 2004-37-0059 | 121 | 1.58 | 1.5803 | 0.095622 | 0.00032222 | 4.806 | 3.5423 | 0.0034886 | -1.2637
2004-37-0061 | 2004-37-0061 | 121 | 231.357 | 231.357 | 0.031418 | -0.00016785 | 4.432 | 3.2565 | 0.0019577 | -1.1755
2004-37-0061 | 2004-37-0061 | 121 | 199.967 | 199.967 | 0.038252 | -0.00038147 | 4.435 | 3.249 | 0.0012974 | -1.186
2004-37-0061 | 2004-37-0061 | 121 | 174.934 | 174.934 | 0.036586 | 0.00035095 | 4.341 | 3.249 | 0.0012682 | -1.092
2004-37-0061 | 2004-37-0061 | 121 | 149.814 | 149.814 | 0.037858 | 0.00047302 | 4.434 | 3.2455 | 0.001119 | -1.1885
2004-37-0061 | 2004-37-0061 | 121 | 125.263 | 125.263 | 0.032155 | -0.00039673 | 4.85 | 3.2443 | 0.0014919 | -1.6057
2004-37-0061 | 2004-37-0061 | 121 | 99.982 | 99.9825 | 0.029542 | 0.00045776 | 4.437 | 3.2461 | 0.0019845 | -1.1909
2004-37-0061 | 2004-37-0061 | 121 | 75.359 | 75.3592 | 0.034804 | 0.00016022 | 4.385 | 3.2424 | 0.0021689 | -1.1426
2004-37-0061 | 2004-37-0061 | 121 | 50.096 | 50.0965 | 0.030348 | 0.00048065 | 4.56 | 3.2401 | 0.00089973 | -1.3199
2004-37-0061 | 2004-37-0061 | 121 | 40.18 | 40.1804 | 0.028261 | 0.00043869 | 4.388 | 3.2424 | 0.001479 | -1.1456
2004-37-0061 | 2004-37-0061 | 121 | 29.974 | 29.974 | 0.03775 | -4.9591e-05 | 4.267 | 3.2437 | 0.0018542 | -1.0232
2004-37-0061 | 2004-37-0061 | 121 | 20.033 | 20.0329 | 0.038224 | -0.00015068 | 4.494 | 3.2427 | 0.0015187 | -1.2513
2004-37-0061 | 2004-37-0061 | 121 | 9.968 | 9.9679 | 0.035681 | -0.00013256 | 4.396 | 3.2448 | 0.0016259 | -1.1512
2004-37-0061 | 2004-37-0061 | 121 | 5.143 | 5.1428 | 0.034867 | -0.00017357 | 4.415 | 3.242 | 0.0016532 | -1.173
2004-37-0061 | 2004-37-0061 | 121 | 2.086 | 2.0859 | 0.027521 | -0.00011563 | 4.431 | 3.2467 | 0.0015492 | -1.1843
Create Fit Sheet