## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 2-May-2005 | Added nutrient data to the rosette files. J.L. |

## PROCESSING NOTES

Cruise: 2004-37

Agency: OSAP

Location: Strait of Georgia / Juan de Fuca Strait

Project: JdeF / SoG

Party Chief: Masson D.

Platform: VECTOR

Date: 6 December 2004 – 11 December 2004

Processed by: Germaine Gatien

Date of Processing: 26 January 2005 –1 February 2005

Number of original CTD casts: 67

Number of casts processed: 67

Number of rosette casts: 21

Number of rosette casts processed: 20 (There was no sampling from one cast.)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Wetlabs CSTAR transmissometer (#498DR), a Seabird Dissolved Oxygen Sensor (#0766) on the secondary pump, Altimeter (#1024), a PAR sensor (#4656) and a Seapoint Fluorometer (#2229) on the primary pump with a 10X cable. The deck unit was a model 911+ (#0508). The salinometer was a Portasal model 8410 (#58879).

# SUMMARY OF QUALITY AND CONCERNS

The log book and rosette logs were in good order. All sensors performed well except for the secondary sensors used for cast #1.

The wrong PAR sensor is entered in the configuration files. It was entered as #4565 but should be #4656, as stated in the log book. The PAR data was recalibrated at the end of the processing.

The average descent rate was fairly low and the weather was often stormy, so that a lot of data was lost from 16 casts due to shed wake corruption. For most other casts the descent rate was quite steady and little data was lost.

The transmissivity, dissolved oxygen and fluorescence are unedited, except where records were removed in the editing of temperature and salinity.

This is the first data set to be processed using Dissolved Oxygen Sensor #0766. The dissolved oxygen data was aligned to correct for transit time and recalibrated based on bottle samples. A further correction for time-response errors was not done because there was a lot of scatter in the analysis, but the values appear to be high by an average of about 0.09ml/l. When more experience has been gained with this instrument it may be decided to return to this data set and apply a further correction.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 150m

- ±0.2ml/l from 150m to 400m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

The bottle chlorophyll, salinity and titrated dissolved oxygen data were obtained with flags and comments added by the analysts.

The nutrients were not ready at the time of CTD processing.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked.

The PAR sensor serial number recorded in the log book does not agree with the configuration files. It was assumed that the configuration file is correct, but checks should be made later to ensure this is correct. (Note: it was later found that the configuration file was wrong and a recalibration was done.)

A small error was found in one of the conductivity co-efficients for cast #1, this was corrected and the file saved as 2004-CTD1.con.

A copy of the con file used for all other casts was saved as 2004-37-CTD2.con.

The history of the conductivity, dissolved oxygen and pressure sensors was found. The pressure sensor had recently been checked at the factory and an offset of 3.04db is now included in the calibration. This fits what had been noted over the past 2 years.

##### Conversion of Raw Data

The raw data were converted using configuration file 2004-37-ctd1.con for cast #1 and 2004-37-ctd2.con for all other casts. A few casts were examined and all expected channels are present. The up and downcasts are reasonably similar and the pairs of sensors are close during downcasts; upcasts show larger differences.

The altimetry was extremely noisy but there is a clear signal near the bottom. Transmissivity looks reasonable while fluorescence seems a bit noisy, but has the right shape. In the log it was noted that the difference in salinity seemed large during cast #1. The sensors were then changed.

Rosette files were converted using a start time of -2s and duration of 5s.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.

##### ALIGNCTD

ALIGNCTD was not run since the deck unit had advanced both conductivity channels.

##### CELLTM

Tests were run on a few casts with (alpha, 1/beta) set to (0.1, 7), (0.01, 9), (0.02, 9), (0.03, 9), (0.02, 7) and (0.03, 7) and (0.0245,0.5). The best results varied from feature to feature and cast to cast and the differences were slight. CELLTM was run using (0.0245, 9.5) on both conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 100 | +0.0025 | -0.0004 | -0.006 | Steady, moderate |
| 38 | 270 | -0.0003 | -0.0002 | -0.0015 | “ |
| 46 | 270 | -0.0003 | -0.0001 | -0.001 | “ |
| 67 | 270 | -0.0003 | -0.0002 | -0.0025 | Noisy, moderate |
| 67 | 380 | -0.0002 very noisy | -0.0003 | -0.0025 | “ |

The difference in temperature was very large for cast #1. The salinity difference was noted at sea and the secondary T and C sensors were replaced. It looks as if the secondary temperature sensor accounted for the problem. For the later casts all differences are reasonably small. It is possible that the secondary conductivity is drifting but there is no hint of pressure-dependence; however, there are no deep casts and there is a lot of noise, so the evidence is weak.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers to the headers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files. CLEAN was run to add event numbers and the output files were named BOT.

All BOT files were plotted and no significant outliers were found. Cast #61 looks odd, well-mixed to the bottom, but Diane Masson confirms this is as expected from that site due to tidal mixing.

##### Checking Headers

The cruise track was plotted and no problems noted.

A header summary, header check and cross-reference listing were produced. There were errors in 2 station names. These were corrected in the headers.

The average surface pressure is 1.98db which is reasonable for the Vector.

The altimetry was exported to a spreadsheet. Only cast #58 lacks an altimeter header value because it was not sampled near the bottom. A few casts were examined and all values are reasonable.

##### 11. SHIFT

Conductivity

Tests were run on several casts with various shifts of conductivity and the best results overall were with an advancement of -0.8 records to the primary and -0.5 records for the secondary.

All casts were put through SHIFT using those settings.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between the two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this cruise the results varied from +1s to +2s. Overall a shift of 1s (+24 records) looked best. This is the value that has been used for most cruises in the past and was applied to this data.

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. Tests were done shifting the DO by +80 to +160 records and the best results were found using +120 records. In other recent cruises using this instrument, either +120 or +140 records has been used. A shift of +120 records was applied to all casts.

##### 12. BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. The addsamp file was edited to add sample numbers taken from the rosette sheets.

The salinity analysis spreadsheet was saved in CSV format. This was converted to SAL files.
One comment was transferred from the spreadsheet to the header of the SAL files for cast #57.

There were replicate samples for 3 bottles of cast #12. The analyst had put a 9 in front of the sample numbers in order to distinguish the two. The ADD file was edited by entering the average value for the bottle along with flag “f”. The extra line was then removed for each pair. The averages and differences for the replicates were 2.195/0.148, 3.301/0.001, 4.898/0.063ml/l.

File 2004-37chlarc.xls was obtained from the analyst. Non-standard channel names were corrected and a few errors in sample numbers were checked with the analyst and corrected. The corrected file was saved as 2004-37chlarc.csv. This was converted to CHL files. The only comments applied to all casts so Headedit was used to add a general comment to the headers and the output was called CHL1.

The sample numbers were added to the BOT files to create SAM files, which were averaged to create SAMAVG files. One line was removed from the SAMAVG file for cast #47, because there were no samples taken from that bottle. They were then merged with the SAL, ADD and CHL1 files in three steps. (Output: MRG1, MRG2 and MRG)

##### 13. COMPARE

Salinity

COMPARE was run. When a few outliers are excluded the primary salinity is higher than the bottles by about 0.0035psu and the secondary is high by 0.0019psu. The primary differences are reasonably flat with time and the secondary is even flatter. The first cast was examined separately because different secondary sensors were used. (See 2004-37-sal-comp1.xls.)

There were 9 outliers excluded from the comparison. These were investigated to see if flags should be assigned to the bottle values. Four were in areas of large salinity gradients, so minor flushing problems can explain the differences. Noisy CTD data accounted for another outlier. Four were studied further:

- For cast #1 the sample at 110db differs from both channels of the CTD by about 0.11psu. The nearest water of that salinity is about 25db above that. The bottle at 100db has salinity higher than the bottle at 110db by about 0.085psu which is very unlikely. If the two deepest samples were reversed the differences at 110db improves but they are worse for the 100db sample. The value at 110db will be flagged “c”.

- For cast #1 the sample at 100db differs from the two CTD channels by about 0.022psu and 0.028psu. There is water within a few metres that would have that difference. No flag is needed.

- For cast #32 at 176db the bottle differs from both CTD channels by 0.075psu. There is no evidence of significant shed wakes. The local gradients are high and water of the right salinity is found about 10db above, so the difference might be due to poor flushing of the bottle; the descent rate was unusually quiet in this region which might affect the flushing. No flag was assigned.

- The sample at 150db for cast #52 differs from both CTD channels by 0.006psu. In this case there is water of the right salinity within 5db and there was a very short wait before firing; it is likely that the bottle did not flush well. No flag was assigned.

Dissolved Oxygen

COMPARE was run and plots were prepared of differences between CTD and titrated values versus pressure, DO and cast number. There was only one severe outlier in COMPARE and that was a value that had been flagged “d” by the analyst. There was one other point that was a minor outlier. It is from deep in Saanich Inlet where oxygen levels were very low. It is from a bottle that was also an outlier for salinity, raising the possibility that the bottle misfired or was poorly flushed. The values look like they are from about 15db above the bottom. Leaving those 2 points out of the comparison produced a reasonably tight fit with

DOX (Titrated) = 1.3593* DOX (CTD) + 0.0694

We have no previous experience with this sensor. The differences are very flat with time. The slope is bigger than I have seen from the other sensors. A “c” flag was assigned to the sample from cast #1 at 110db. (See 2004-37-dox-comp1.xls)

Fluorescence versus titrated Chlorophyll

COMPARE was run and the CSV file used to plot chlorophyll versus fluorescence. The data fell into 2 separate groups. For FL<0.7ug/l the fluorescence was about 20% higher than the chlorophyll. For higher fluorescence the values were lower than the chlorophyll. There were only 6 points with the higher FL values and all of them are from near the mouth of Juan de Fuca Strait.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

All the DEL files were copied to *.EDT.

##### DETAILED EDITING

The primary sensors will be chosen for cast #1 since there is good evidence that the secondary sensors did not perform well.

The secondary sensors will be selected for all but cast #1. The secondary salinity was a little closer to the bottles, the differences were flatter versus time and pressure and the data is a little less noisy.

Page plots were produced. These were used to guide the editing.

The descent rate for some casts was extremely noisy, and given an average descent rate that was usually about 0.7m/s, there was a lot of corruption by shed wakes. Some casts had a very steady descent and needed little or no editing. There was only minor spiking in the salinity channel.

The following casts required heavy editing: 2, 3, 5-10, 17, 18, 56-59, 63 and 67

The following casts required moderate editing: 2, 11, 12, 21, 62

The following casts required no editing: 14, 19

All other casts required light editing.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensors were recalibrated shortly before this cruise, so there is no experience from other cruises. The secondary sensors used for the first cast were also used for 2004-29 when there were some problems that might have been due to drift with time or pressure; the salinity was found to be low by about 0.003psu. The secondary sensors used for all other casts were also recalibrated recently and there is no useful history for them.

Historic ranges – The temperatures were above the historic range maxima for casts in the northern part of the Strait of Georgia from 140db downwards (and even shallower at some stations.) There were a few spots where salinity was also slightly high. These observations are not considered to be indicative of instrumental problems.

Repeat Casts – There were no repeat casts.

##### Initial Recalibration

File 2004-37-recal1.ccf was prepared to recalibrate both salinity channels and the dissolved oxygen using the results of section 12. (Corrected primary salinity = Sal0 – 0.0035; Corrected secondary salinity = Sal1 – 0.0019; Corrected DOX = 1.3593* DOX (CTD) + 0.0694). This was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for SAL and DOX and the corrections were found to be made correctly. (See 2004-37-sal-comp2.xls and 2004-37-dox-comp2.xls.)

### Final Dissolved Oxygen Comparison

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but in the past there have sometimes been significant errors due to response-time problems. To check for such errors the downcast CTD data (after SHIFT and CALIBRATE) was compared with the bottles from the upcast. A set of downcast files were prepared by bin-averaging (0.25db bins) the recalibrated downcast files and thinning the data to the depths generally used for bottles. COMPARE was run comparing those files with the bottle data in the MRG files. The fit of differences versus pressure was found excluding severe outliers. The CTD is higher than the bottles by an average of +0.088ml/l. But there is a lot of noise and the offset is not as constant as had been noticed in the past. This is a new instrument and the waters sampled here are unusually varied. A pressure-dependent recalibration may be necessary, but this is a poor cruise on which to base such a recalibration. When there is more experience with this sensor it might be worth revisiting the calibration of this data. (See 2004-37-dox-comp3.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD1.

The SAMCOR1 files were put through CLEAN, REMOVE, HEADEDIT and saved as BOF1.

CALIBRATE was run on BOF1 and FCTD1 files to correct the PAR setting. (See section 22 for details.) The output files, FCTD and BOF, were saved to a CD-ROM for Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the full COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

(Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared.

Separate profile plots were prepared with Transmissivity, DOX, PAR and Fluorescence versus pressure.

##### FINAL CTD files steps (REMOVE, HEADEDIT and CALIBRATE)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag channels were removed from cast #1.

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag channels were removed from all casts except for cast #1.

HEADER EDIT was used to fix formats and units and to add the following comment using file 2004-37hdr.txt:

Transmissivity: The data are unedited except where records were

removed in editing temperature and salinity.

Fluorescence:URU:Seapoint - The data are unedited except as above.

The dissolved oxygen data are unedited except as above.

The dissolved oxygen data in the CTD files should be considered:

•	±0.5ml/l in the top 150m

•	±0.2ml/l from 150m to 400m

The Standards Check routine was run and HEADER EDIT adjusted and rerun until no further problems were found. These files were named CTD1.

Late in the processing it was discovered that the wrong calibrations had been used for the PAR sensor. A comparison was done between cast #31 converted with the two different PAR settings. The fit of the two is linear except very near the surface where the original file had values that overflowed. None of those values appear in the final files. File 2004-37-recal2.ccf was prepared to correct this by multiplying PAR by 0.2465277778. It was applied to the CTD1 files with output CTD.

A cross-reference listing was produced.

The sensor history was updated.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird.

SORT was used to put the data in order of increasing pressure.

REMOVE was used to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag for cast #1.

REMOVE was used to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag for casts #2-68.

HEADER EDIT was run to fix channel names and formats and to add a standard comment including an explanation of the quality flags. The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The output files were named CHE1.

CALIBRATE was run on the CHE1 files using file 2004-37-recal2.ccf to correct the PAR channel. The output was named CHE. (For details see section 22.)

Particulars

There was a note in the log referring to problems with dissolved oxygen titration. Tests were done using a different system to ensure data was ok.

1. Secondary Salinity difference was about 0.006 so sensors were changed after this cast.

6. Bottle #6 half closed at the top.

31. 3 bottles fired by mistake – no sampling.

40. May have touched bottom during drift at depth.

47. Bottles 4 and 5 fired at 200db. No sample from bottle 5.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-37

Dates: Start: December 06, 2004 End: December 11, 2004

Location: JdeF/SoG Vessel: Vector

Party Chief: Masson D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443               Cruise ID#:		2004-37

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 18/11/04 | Factory |  | Factory |
| Conductivity | 2399 | 19/11/04 | “ |  |  |
| Secondary Temp. | 2038 | 19/08/04 | “ |  |  |
| Secondary Cond. | 2173 | 19/08/04 | “ |  |  |
| Secondary Temp. | 2106 | 18/11/04 | “ |  |  |
| Secondary Cond. | 1763 | 19/11/04 | “ |  |  |
| Transmissometer | 498DR | 05/08/04 | IOS |  |  |
| Dissolved Oxygen | 0766 | 16/11/04 | Factory |  |  |
| Fluorometer | 2229 |  |  |  |  |
| PAR | 4656 | 11/02/03 |  |  |  |
| Altimeter | 1024 |  |  |  |  |
| Pressure Sensor | 63507 | 25/10/04 | Factory |  |  |