## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2004-42

Agency: OSAP

Location: Strait of Georgia

Project: Beamish

Chief Scientist: Beamish R.

Platform: W. E. RICKER

Date: June 25, 2004 – July 18, 2004

Processed by: Germaine Gatien

Date of Processing: 7 June 2005 –10 June 2005

Number of original CTD casts: 83

Number of casts processed: 83

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit is believed to have been a SeaBird model 11 (S/N 0471.)

# SUMMARY OF QUALITY AND CONCERNS

There was no Log Book available.

There was no salinity calibration sampling. Recalibration was based on the drift estimate from the post-cruise calibration report.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

There was no log book, nor any records about this cruise.

There was no bottle data available.

The configuration files were obtained. The sensors used had not been recalibrated since April 2002. It is known that there was a lot of drift during that time, since there were post-cruise calibrations in November 2004 for the T and C sensors. So the post-cruise calibrations will be used for conversion. (2004-42-CTD.con contains the post-cruise calibrations.) Later a correction can be made to allow for the drift over the 4 months between use and post-cruise calibration.

The sensor history was found.

##### Conversion of Raw Data

The raw data were converted using configuration file 2004-42-CTD.con.

All expected channels were present.

The upcast and downcast data are reasonably similar, and the two channels agree well on the downcast, but not so well on the upcasts.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2	Pass 2    Std Dev = 5     Points per block = 50

##### ALIGNCTD

ALIGNCTD was not run even though the deck unit was one of the older model since recent cruises suggest that even the secondary conductivity is being advanced. If this is not so, SHIFT can be used later to do this correction.

##### CELLTM

Tests were run on three casts for each leg running CELLTM with choices of (0.01,7), (0.01,9), (0.02,7), (0.02,9), (0.03,7),(0.03,7) and (0.0245,9.5) for (alpha, 1/beta). The casts were at least 300db deep and had a quiet descent rate.

For 2004-32 the best choice for the primary sensors was (0.02, 7) and for the secondary, (0.0245, 9.5), though several other choices were very similar. For this cruise (0.02, 7) looked best for the primary, though (0.01, 7) and (0.01, 9) were very similar. No tests were done on the secondary conductivity since the secondary salinity looks quite noisy; further, this sensor had not been selected for archiving from any other cruise in 2004 despite many deployments.

CELLTM was run using (0.02, 7) for the primary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Three casts that sampled to at least 300m were plotted to check for agreement between the pairs of T and C sensors. The differences were extremely noisy and varied from cast to cast and from one part of the cast to another; these casts are too shallow for us to expect good results from this type of study, but are still noteworthy for the high variability. For example, the temperature differences for cast #56 are about   -0.0004 at 275m and -0.0007 around 320m with no change in descent rate to account for it; this is probably related to the temperature gradient, suggesting that one sensor is resolving gradients better than the other. The conductivity variations are associated with noisy secondary conductivity. There have been problems in the past with this equipment that were believed to be pump related, but they usually involved poor data near the surface. There is no sign of that problem with this data.

The descent rate was generally steady and high.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers based on the last 4 numbers of the files name and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header check was run and indicated that 2 casts had scrambled SeaBird time/position headers. These were corrected, reconverted and put through CLEAN. No further errors were found and the speeds between casts look ok.

A header summary was produced.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 2.0db which is lower than found on High Seas Salmon cruises, but there is no indication of a Niskin bottle being attached so the routine on this cruise was probably quite different. There has been no problem in the past with drift in the pressure calibration for this sensor. One cast (#64) has a minimum pressure of-20db. There is clearly something wrong with the pressure for that cast since the temperature and conductivity values indicate the CTD was in seawater. Excluding that cast the minimum pressure was 0.2db for cast #80; the conductivity at that pressure shows the CTD was very close to the surface.

##### SHIFT

Fluorescence

A quick check of a few casts suggests that a shift of +24 records is appropriate; this has been used for most cruises in the past, including other uses of this particular equipment in 2004.

SHIFT was run using a setting of +24 records for all casts.

Conductivity

For all 2004 cruises using this equipment a SHIFT setting of -1.2 records for the primary conductivity was found effective in minimizing salinity spiking without oversmoothing. The secondary salinity has not been archived in recent cruises and is not expected to be archived for this cruise, so tests will not be done for secondary conductivity. Tests were run on casts #4 and 56 using advancements of from -1 to -1.4 records for the primary conductivity. The results were examined in T-S space with the best results those that minimize unstable spiking in salinity without oversmoothing. The setting of -1.2 looked best overall.

All data were put through SHIFT using -1.2 for the primary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00	Minimum Salinity: 5**

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 decibars to 10 decibars less than the maximum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: There were no warnings.

**While editing it was noticed that too much data had been removed near the surface when there was sampling in very low salinity surface waters, so DELETE was rerun on casts #9 -12 using Minimum Salinity: 0. The other casts had higher salinity at the surface so were not rerun.

##### DETAILED EDITING

The primary sensors were selected for editing.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Salinity generally needed light editing to remove spikes due to electrical noise and small “overshoots” in large T gradient areas due to imperfect alignment of T and C.

The descent rate was generally kept high and very steady so there was little corruption by shed wakes.

The following casts required no editing: 14, 17, 23, 39.

All other casts required only light editing. Cast #64 was edited a second time to remove the first 440 records because the pressure was bad as noted in section 10. The other data looks ok.

Note was made in the headers about any editing done.

The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that do not require editing.

##### Recalibration

Based on the drift estimate during the post-cruise calibration, the primary salinity is probably high by about 0.002, so file 2004-42-rcal.ccf was prepared to subtract 0.002psu from the salinity. (output: COR)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the COR files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

All casts were plotted on-screen to check that transmissivity and fluorescence were ok. The only problem noted was in cast #9 during which most transmissivity values were 0. The transmissivity channel will be removed from that cast only.

##### Other comparisons

Previous experience with these sensors – This equipment has been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter.

Historic ranges – While most of the data fell within the climatology, there were some excursions, mostly with high temperature and low salinity. One of the most striking was for cast #83, there was a June cruise by Diane Masson cruise during which cast #14 in the same area showed similar results. The excursions are larger in the July data, but of the same character. Most of the excursions for other casts are found close to shore where the local climatology may not be representative. When casts well away from shore were examined the only excursions were in the northern section, where near-surface temperatures which were a little high. Similar results were seen frequently in 2004, and in particular in the Queen Charlotte Strait area in September. None of these excursions are believed to be evidence of problems with the sensors.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S and Transmissivity profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts. The transmissivity channel was removed from cast #9 only. A second run of REMOVE was needed because there two Descent_Rate channels had been created accidentally.

HEADER EDIT was used to fix formats and header entries and to add the following comments:

The data was processed without access to any log records, so

times and positions are unconfirmed.

The fluorescence and transmissivity data are nominal and unedited

except that some records were removed in editing T and S.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were resolved.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

9. Transmissivity values mostly 0 even at 100db.

64. Problem in pressure at surface.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-42

Dates: Start: June 25, 2004 End: July 18, 2004

Location:

Vessel: W.E. Ricker

Party Chief:

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2004-42

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory | 18nov04 | Factory |
| Conductivity | 2424 | 16/04/02 | “ | 19nov04 | “ |
| Secondary Temp. | 2374 | 20/06/02 | “ | 18nov04 | “ |
| Secondary Cond. | 2399 | 16/04/02 | “ | 19nov04 | “ |
| Transmissometer | 197 | 16/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |