## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 23-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. |
| 28-Oct-2013 | Merged DMS data to bottle casts from Mike Arychuk’s spreadsheet file located in the \DOC directory. For information on methods see file "Method post 1999 for Dimethylsulfide Analysis.doc” in directory \\OSD_Data_Archive\Cruise_Data\Documents\Analysis Reference Papers\. |
| 6-Jun-2013 | Added Iron profile files with cast numbers 8xxx from Keith Johnson’s spreadsheet file which can be found in the cruise .DOC directory. |
| 11-Apr-2010 | Added Lisa Miller’s Dissolved Inorganic Carbon and Alkalinity data to the rosette files. J.L. |
| 19-Jan-2006 | Added loop data to the archive. |
| 15 November 2005 | TSG Fluorescence recalibrated using manufacturers 2001 calibration |

## PROCESSING NOTES

Cruise: 2005-01

Agency: OSAP

Location: Gulf of Alaska

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: February 11, 2005 – February 26, 2005

Processed by: Germaine Gatien

Date of Processing: March 20, 2005 – 6 June 2005

Number of original CTD casts: 62 (Of these, one was split into downcast and upcast.)

Number of CTD casts processed: 61

Number of rosette casts: 26 full rosette casts and 35 with 5m Niskin

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with Chelsea/Seatech transmissometer (#333DR), Altimeter OA-916D (#1024), SeaBird DO sensor #766 and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0424) and the logging computer was PAC02392. The salinometer was a Portasal model 8410 (#58879). A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a fluorometer.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was generally in good order. One sample number was assigned twice. The pressure sensor number recorded in the front of the log book matches neither the CTD# nor the configuration file, so is presumed to be wrong.

Cast #10 was interrupted after the first bottle was fired. Event #11 is the continuation of that cast. The bottle files were joined and named as event #10.

Cast #29 was missing the top 80db of data, so it was joined to the top 80db of cast #34 at the same site.

The salinity bottle analysis was interrupted when problems were found with the Portasal. After repairs the results were very different so all bottles analyzed in March were flagged “d” and excluded from the calibration comparisons.

There remain some uncertainties in the bottle salinity analysis. There were 23 bottles fired at 2000db during one cast; of those 13 fall within 0.0005psu of one value, but there is a large scatter in the other 10. There is no evidence that the problem is with the Niskin bottles or collection techniques. It seems most likely that the Portasal is at fault. Although a flat comparison was achieved between the secondary salinity and the bottles analyzed in April and May, there was a large scatter in the results, another sign that the Portasal is not working well. Salinity should be considered ±0.002psu due to the scatter.

The Thermosalinograph salinity is considered ±0.01psu.

The comparison of the dissolved oxygen data with the bottle samples was quite different during casts 1, 3 and 6 than later in the cruise. The deep waters of cast #1 were anoxic, which may have affected the sensor. Separate recalibration schemes were used for the early casts.

The dissolved oxygen data in the CTD files should be considered

- ±0.6ml/l from 0 - 200m

- ±0.4ml/l from 200 – 400m

- ±0.1ml/l from 400 - 1300m

- data was removed from 1300m down since it is considered unreliable there

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Together with the notes from Marie Robert this gave an excellent summary of problems encountered during the cruise.

The pressure sensor number recorded in the front of the log book is not what is in the configuration file, nor is it the sensor expected for CTD #0443. The configuration file is presumed to be correct.

The bottle salinity was not all available at the beginning of the processing job.

Dissolved oxygen data were obtained; flags and comments had been added.

The cruise summary sheet was completed.

The histories of the conductivity and pressure sensors were obtained.

The calibration constants were checked for all instruments and all were correct.

There were two configurations varying only in whether the Altimeter or pH sensor was mounted in the voltage 6 spot. New configuration files 2005-01-ctd1.con  and 2005-01-ctd2.con were prepared with an altimeter in the first and pH sensor in the second.

##### Conversion of Raw Data

All data were converted using file 2005-01-ctd1.con for all casts except those with a pH sensor.

Casts 10, 11, 23, 24, 44, 52, 56 and 60 were converted using file 2005-01-ctd1.con.

A few casts were examined and all expected channels are present.

Rosette files were converted using a start time of -2s and duration of 5s. The rotated time/position headers of cast #66 were fixed. The rosette files were then converted to IOS SHELL files. All files were put through CLEAN to add event numbers.

All BOT files were plotted. The only problem noted was noise in the secondary salinity at 600db. CTDEDIT was used to clean that data and the output was copied to BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### ALIGNCTD

The deck unit was one of the older ones but in recent use the secondary conductivity channels seem to have been advanced. ALIGNCTD was not run. Any alignment problems will be addressed using SHIFT when the files have been converted to IOS SHELL format.

##### CELLTM

Tests were run on 3 casts with (alpha, 1/beta) set to (0.01,9),(0.02,7), (0.03,7), (0.02,9), (0.0245,9.5) and (0.03,9).

Both (0.01, 9) and (0.02, 9) improved the primary data with variations at different depths. The overall best choice was with (0.01, 9) especially in regions of high temperature gradient. There was a lot of noise in the data and most casts had many stops for bottles, so comparing up and downcast is difficult.

The setting of (0.02,9) improved the secondary conductivity best.

CELLTM was run using (0.01, 9) for the primary and (0.02, 9) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 12 | 2000 | ~0 | -0.00007 | +0.0008 | High, very noisy |
| 35 | 2000 | -0.00025 | -0.00048 | -0.005 | High, noisy & spiky |
| 66 | 200 | -0.0005 | -0.0008 | -0.0085 | High, steady |
| 66 | 2000 | -0.00025 | -0.00073 | -0.0087 | High, steady |
| 66 | 4760 | +0.00025 | -0.00088 | -0.011 | High, steady |
| 97 | 2000 | -0.00025 | -0.00092 | -0.0115 | High, steady |
| 97 | 2500 | -0.0002 | -0.00095 | -0.115 | High, steady |

All the differences were very noisy.  There is obvious pressure dependence in the conductivity and salinity differences. The temperature differences look pressure dependent for cast #66. Two other very deep casts (#39 and 75) were checked to see if the pressure dependence was similar there and it was. It is unusual to see pressure dependence in the temperature differences. Similar observations were made during 2004-37, when the primary sensors were used throughout. The secondary sensors were different in the beginning but were switched part-way through to the ones used during this cruise. There were no deep casts during 2004-37, particularly in the late part of the cruise. So it may be that the primary sensors were the ones at fault, or that the problem is with the secondary plumbing or pump.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and pointed to a few errors in the headers. The time/position entries were scrambled for cast #66. This was fixed in the SeaBird headers and the file was reconverted to IOS Headers and put through CLEAN again. A station name was added to cast #27. The header check was rerun and no further errors were found.

The header summary was run and an error found in the geographic area entry in the header for cast #66; this was corrected.

The cruise track was plotted and no problems found.

The average surface pressure is 3.3db, which is a little higher than usual, but there is a wide range of values from 0.5 to 8.5db. The pressure sensor had been calibrated in October 2004 and an offset is included in the configuration file based on that calibration.

A few casts were examined on screen. The upcasts are much noisier than the downcasts with larger differences between sensors. For cast #2 the primary salinity was bad around 100db on the upcast.

A spreadsheet was prepared with altimeter readings from the headers of the CLN files. A few of these were checked against the original files to ensure the algorithm to calculate the header entry from the altimetry worked well. The data is very noisy so it should be considered a rough estimate, but the algorithm did provide satisfactory results..

The transmissivity and fluorometer traces look reasonable.

No other problems were noted.

##### BOTTLE FILE PREPARATION

The bottle files from casts #10 and 11 were joined and named 2005-01-0010.bot.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. The bottles for cast #11 were removed since they have been added to cast #10. The 5 deep bottles were removed from cast #27 since the sample numbers were non-standard and no sampling was done that will be analyzed at IOS.

There was an error in assigning sample numbers to cast #67 and #68. Sample # 340 was repeated. A new number was created for the single sample from cast #67 – it was renamed as sample #9340. There is no rosette sheet for this sample, so it is unknown what sampling was done; probably it was a surface bottle.

Sample numbers were then added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

There were concerns about the salinometer, so not all samples were analyzed at first. Those that were available were delivered by spreadsheet file 2005-01 sals.xls. The spreadsheet was simplified by dropping a few columns and was saved as 2005-01-sal-apr.29.csv so that it would be obvious it is incomplete if and when more data becomes available. Where duplicates exist, the average was entered in the data column, but the separate values entered as a comment. The flag was set to “f” for those samples. The records were ordered on sample number.

In May another set of salinity samples were analyzed. These were merged with the earlier samples and again ordered on sample number. The file was named 2005-01 sal-non-loop-all.csv.

Sample # 340 was used twice. For cast #67 that sample was renamed as 9340, and for cast #68 it was left as #340. (The analyst used 340 for cast #67 and 7340 for #68 so this was changed in the spreadsheet.)

The 2005-01-sal-non-loop-all.csv file was converted to individual SAL files.

Dissolved oxygen files (*.add) created by the analyst had a flag channel. There were a few duplicate values that were entered with sample numbers with an extra 9 in them. These were removed from the ADD files and the information entered into the header as a comment. There was one repeated sample number, but one of the entries has a 0 value for DO. The 0 entry was removed. In several files the flag was entered one space too far to the right. This was fixed.

The nutrient spreadsheet was simplified and saved as 2005-01-nuts.csv. That file was then converted to NUTS files. The values for sample #72 were removed because the samples were believed to be mislabelled. There are some comments for values that were not flagged; they were entered by the analyst, Wendy Richardson, for the information of users, but she does not consider the observations an indication of problems in the sampling.

Final extracted chlorophyll data were not available at the time of processing.

The SAL, ADD and NUTS files were merged with SAMAVG in three steps. (Output: MRG1, MRG2, MRG)

DO was plotted against CTD Salinity; the only outlier came from a sample that was already flagged “d”.

##### 12. COMPARE

Salinity

The bottle salinity samples were analyzed in 3 batches. There was a problem with the Portasal noted after samples were analyzed in March; the analyst found the instrument to be stable during that analysis. Following repairs, a second batch of samples were analyzed in April.  When COMPARE was run using the first 2 batches of data, the results were quite different. The March Portasal analysis was on data from casts 1 to 14, plus the first 8 bottles of cast 15, plus one bottle from cast 17. All of the March data looks notably different from that analyzed in April. Even sample #79 stands out and it is in the middle of some samples analyzed in late April. So this appeared to be a Portasal issue, rather than the drifting of sensors. The analyst stopped the work in April because of concerns over the scatter in results from duplicates and repeat bottles at one depth. The rest of the bottles were analyzed in May after further repairs.

COMPARE was rerun with all three batches. Once again, the March data stands out as different. For example, the secondary salinity appears to be high by about 0.01psu in March and low by about 0.004psu in April and May. (See 2005-01-sal-comp1.xls.)

The data from April and May are close. When the March data and a few outliers are excluded, the trend line for the fit of secondary salinity is remarkably flat with time and pressure. However, there is a lot of scatter in the comparison. The salinity is low by an average of 0.0044psu. The primary salinity shows significant time dependence, but is fairly flat with pressure; again there is a lot of scatter in the results. The primary salinity is high by an average of ~0.0023psu, but early in the cruise it is low by ~0.002psu and high by ~0.008psu at the end.

There were 8 duplicate salinity samples taken. The average difference between duplicates was 0.0004psu but this includes a lot of variability. The average of the absolute value of differences is 0.0018psu which is fairly large. If two outliers are excluded that is reduces to 0.0007 and if the March values and one other outlier are excluded the average is 0.0004psu. Those 4 were all run on one day after repairs to the salinometer.

It has sometimes been observed that the CTD behaves a little differently in motion than while stopped. So a comparison of the CTD channels during bottle stops was made to ensure this was not complicating the issue. The differences are very similar to those noted in the downcast “in-motion” CTD data, so this is not a factor. (See section 8 for details.)

COMPARE was run again using the bottle number as the reference channel to determine if there were a few bad bottles causing the large scatter in differences. There were 23 bottles fired during P18 at 2000db. This discussion will look only at the secondary salinity, but the results were similar for the primary. There are 13 cases in which the differences (secondary salinity – bottle salinity) fall within 0.0005psu of -0.0041psu. For another 7 cases the differences fall within 0.0012psu of -0.0065psu. The other 3 bottles have larger differences. The only pattern is a slight tendency to larger differences towards the end.

For P18 the most extreme outlier was from bottle #20 and bottle #23 was a slight outlier. When the data from all casts was analyzed together there is variability from one bottle to another, but given the scatter, nothing stands out as terribly significant. Perhaps bottle #25 has a slight problem, but the bottles that stood out during the P18 firings look fine in the overall comparison. There are outliers for most bottles. There is a lot of variability in the results. The CTD standard deviations suggest no significant noise in any of the CTD data to account for this. The problems must be with either the Portasal or sample collection techniques, and there is no evidence to suggest the latter had changed significantly.  (See 2005-01-sal-niskin-comp.xls)

All the salinity samples that were analyzed in March will be flagged “d”.  No other flags were assigned.

In conclusion, the secondary salinity is the best choice for the archive since it is flat with time and pressure. It is probably low by about 0.0044psu.

Dissolved Oxygen

COMPARE was run using sample number as the reference variable. The best fit was with the differences between bottles and CTD versus CTD DOX. First, data from below 1300db were excluded from the fit and then outliers were excluded as indicated by the fit residuals; the fit found was:

CTD-BOT = 1.4931 DOX-CTD + 0.0564

It is noted that the differences were significantly higher for the first few casts, 1, 3 and 6. In part this may be because they sampled shallower waters with a different range of DO from the later casts, but it seems likely that there is a more serious issue at play. The sensor sampled anoxic waters during cast #1; there may have been some temporary effect on the sensor which may have taken a few casts for the sensor to return to normal. The CTD touched bottom during cast #3 which may be causing further confusion. There is no suggestion of time-dependence after cast #6. (See 2005-01-dox-comp1.xls.)

A separate analysis was done for the first 3 rosette casts and a decision made to recalibrate them individually. The following calibrations will be used:

CTD-BOT = 1.3559 DOX-CTD - 0.0402 (For cast #1)

CTD-BOT = 1.1112 DOX-CTD + 0.9646 (For casts #2 and 3)

CTD-BOT = 1.3643 DOX-CTD + 0.0905 (For casts #5, 10 and 12)

CTD-BOT = 1.4931 DOX-CTD + 0.0564 (For casts #13 to 102)

There are no bottles for casts #2, 5, 10, 12, 13 and 14, but comparing surface DO suggests 10 and 12 may be slightly high while #13 and #14 look much like cast #15 which was quite close to the general fit.

During 2004-37 when this sensor was used the result was

CTD-BOT = 1.359 DOX-CTD + 0.069

That cruise was one from Juan de Fuca/Strait of Georgia, in which the range of DO was limited with no really deep casts; it was noted then that differences were quite flat with time

Sample 441 which had been flagged by the analyst as “c”, showed up as a severe outlier so the flag was changed to “d”.

##### 14. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was 0.9 to 2.3s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Conductivity

During 2004-37 when these sensors were also used, the primary and secondary conductivity channels were shifted by -0.8 and -0.5 records, respectively. Tests were run on casts #1, 38 and 101 using a variety of settings near those to see what works best at reducing instabilities without oversmoothing. The best choice was a shift of +1.0 records for the primary and +0.2 records for the secondary.

All casts were put through SHIFT using +1.0 records for the primary conductivity and +0.2 records for the secondary. (Output *.SHFC1 and SHFC2).

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Values from +80 to +180 were tried and the best overall match of features was with a choice of +140 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerned cast #11 which was an upcast only.

##### DETAILED EDITING

The secondary sensors were chosen for the archive. Page plots were produced using (T1,S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing, Casts requiring heavy editing were: 2-6,12-13,15-18,23-26.

Cast #29 required special handling. The top 81.2db of downcast data were acquired with the pumps turned off. Cast #34 was from the same site, 6 hours later and 0.5km away, so program FRACTURE was used to create a file containing only the top 81.2db of data. Program JOIN was used to patch the edited downcast file of #29 to the top 81.2db of the edited file for cast #34. The joined file was named 2005-01-0029.new2. The order of joining was chosen so the headers are from the deep cast, but this meant the 0-81db data came after the deeper data. A text editor was used to rearrange the data. CLEAN was run to ensure the headers are correct. The output was named 2005-01-0029.NEW.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors –These sensors were used during 2004-37 when the primary sensors were found to be high by 0.0035psu and the secondary were high by 0.0019psu.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The salinity was low between 10 and 40m for cast #10, but otherwise all data fell within the climatology.

Comparison with ARGO float – An ARGO float with dissolved oxygen sensor was launched near cast #85. The data from 10 days later was compared with that of cast #85 after the latter was recalibrated, metre-averaged and thinned to the same depths as the ARGO file. The DO traces were very similar in shape. The SBE DO was higher, but never differed by more than 1ml/l, the largest differences being in the thermocline. Below 1500m the SBE is lower, as is expected. For temperature and salinity the ARGO float was always lower than the CTD, with temperature differing by an average of 0.025Cº and salinity by 0.022psu between 800 and 2000db.

##### Initial Recalibration

An initial recalibration was run on the MRG and SAM files using file 2005-01-recal.ccf to raise the secondary salinity by 0.0044psu and to apply the corrections to the dissolved oxygen as noted in section 12. COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the calibration was applied to the EDT files as well. The variable calibration of the DO channel worked very well. However there remains more scatter in the salinity than usual, on the order of ±0.002psu.  (See 2005-01-sal-comp2.xls and 2005-01-dox-comp2.xls.)

##### Special Fluorometer Processing

The COR files were clipped to 100db. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

The metre-averaged downcast files were thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences were fairly flat when plotted versus pressure or DO values when only values above 1300m were included. On average, the downcast DO data appears to be high by about 0.1ml/l. A second recalibration was applied to the downcast files only (AVG, THN1 and CLIP) to subtract 0.1ml/l. (Output: COR2, THN2 and CLIPCOR)

A final run of COMPARE was done to ensure the recalibration was done properly and it looks fine. (See 2005-01-dox-comp3.xls and 2005-01-comp4.xls.)

The clipped files were put through REMOVE and HEADEDIT and named as *.FCTD and saved for Angelica Peña.

The SAMCOR2 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag. PAR was removed from all casts except 1-10, 34, 44, 60. (Output: *.REM)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

There were problems with the calibration of the salinity. The

salinometer malfunctioned and data run before and after repair

look notably different. The later data has been used for

recalibration, but there remain some doubts about it.

The SBE dissolved oxygen sensor data were recalibrated differently

for the earlier casts. Anoxic conditions during cast #1 may have

affected the sensor up to cast #15. All casts from #15 onward were

treated the same. The DO comparison with bottles had more scatter than

usual.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.6ml/l from 0 - 200m

•	±0.4ml/l from 200 – 400m

•	±0.1ml/l from 400 - 1300m

•	data was removed from 1300m down since it is considered

unreliable there.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data. Profile plots were made displaying PAR, DO, fluorescence and transmissivity profiles.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag. PAR was removed from all casts except 1-10, 34, 44, 60. (Output: *.MRGREM)

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The quality flags for salinity from the March analyses were changed to “d”. (Output: CHE)

The chlorophyll data became available after the CHE files had been created. The spreadsheet was converted to CHL files; those files were put through HEADEDIT to fix the units and named CHL1. The CHE files were renamed CHE1, sorted on sample numbers to create CHE2, then merged with the CHL1 files and the output files were named CHE.

##### 24. Thermosalinograph Data

a.) Checking calibrations

There were 4 files containing TSG data, but one was very small. All con files were identical. A report was printed for the first con file and the calibrations were checked for temperature and conductivity. No errors were found. The primary sensors were recalibrated in Dec. 2004, the secondary temperature in August 2004. When last calibrated the drift in salinity was found to be +0.0008psu per month and the temperature drift about -0.0012Cº.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

In the first two files, the latitude and longitude channels are bad – apparently the acquisition system was set to add positions to the headers only. The positions are ok in the 3rd and 4th files. Marie Robert prepared spreadsheets with the data from the first two files to which she added positions taken from the SCS to match the times in the TSG files (determined from the start time and interval.) This should work well if the timing is good, which it has been in recent use.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels to files 3 & 4; those files were named *.ATC.

Time-series plots were produced and no severe problems noted. There were no suspect spikes in the T and S data. The salinity is suspiciously flat for most of time, but shows reasonable variations for the last day. A quick check of CTD data shows that there was in fact very little variation in surface salinity, so the data is probably fine. There are a few points in the spreadsheet files that are suspicious.

c.)  Checking Time Channel

The CTD data after editing and metre-averaging was thinned to reduce the files to a single point at or within 1db of 4.5db and exported to a spreadsheet. The TSG files (either the spreadsheets prepared by Marie, or the ATC files after opening in EXCEL) were reduced to the times when CTDs were run. Those files were combined in a spreadsheet and the various temperature, salinity and fluorescence data compared. The first cast (Saanich Inlet) was excluded because the vertical gradients were large, so failure to match the intake pressure would lead to large differences.

Since times were used to match the records, comparing positions should be a good test of timing for the 4th file, and in the case of the 2 files prepared by Marie, they are a test of her patching of positions. The differences between the latitudes and longitudes each had an average of 0.0001º, so there are no apparent problems. This spreadsheet will also be used in step (e). (See 2005-01CTD-tsg-comp.xls)

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are no obvious problems this time, so this step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 The differences between the two temperature channels were plotted for the 4th file. They are very noisy with a few quiet patches that probably correspond to stops. If all data is included the temperature from the lab is higher than the remote temperature by an average of 0.1983Cº. When differences were examined during 5 CTD casts (65, 68, 78, 92 & 102) the average differences were -0.1937, -0.1741, -0.2041, -0.2099 & -0.1916 respectively, for an average of -0.1947 which is slightly lower than the overall average but remarkably close. The cast for which the local temperature variations were lowest was #65, and that is very close to the average. When the differences were sorted by value and the first and last 5000 values excluded, the average was -0.1969Cº. So we can feel some confidence in saying that the TSG temperature is higher than the intake temperature by about 0.197Cº.  What we cannot be so confident about is how reliable the intake temperature is.

During 2004-29 the differences were found to be -0.13Cº and during 2005-02 they were -0.21Cº. We expect some drift in the differences with time and there is some cause to expect these numbers to vary with average water temperature; surface temperatures for this cruise are lower than those of 2004-29 and a little lower than for 2005-02.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. Averages were calculated when outliers were excluded – there were 3 records for which all variables were outliers and each of the variables had one other record that was an outlier. (See 2005-01CTD-tsg-comp.xls)

The TSG primary temperature was higher than the CTD by about 0.22Cº and the remote TSG temperature was higher by about 0.017Cº. The TSG salinity was lower than the CTD salinity by about 0.03psu. When the same equipment was used during 2005-02 the intake temperature was found to be higher than the CTD by about 0.06Cº, but that may reflect higher temperature gradients during the later cruise.

The same data was used to compare the fluorescence from the two instruments, and the fit was                                                                                                                                                                                                                      CTD FL = 6.36 * TSG Fl – 0.36 with the CTD Fluorescence being an average of 4.6 times the TSG Fluorescence.

- Loop Bottle Comparisons There were 12 loop bottles. The salinity values were compared with the TSG files and indicate that the TSG is low by an average of 0.039psu. There is a lot of noise with values from +0.04 to -0.08psu but no suggestion of time dependence. (See 2005-01-loop-sal.xls.)

f.) Calibration History

The TSG was recalibrated in December 2004. It was used for 2005-02, but there were some concerns about the performance of the Portasal when the loop samples were processed for that cruise. The TSG was lower than the loop samples by about 0.02psu, but those samples are believed to be low by at least 0.01psu which would imply the TSG was low by at least 0.03psu.

Conclusions

Since the primary temperature sensor was recently calibrated, we expect good data. The temperature in the TSG is 0.197Cº higher than the intake temperature and higher than the CTD by about 0.22Cº. The 0.20Cº difference is presumed to be due to heating in the ship. During 2005-02 the difference was found to be a little less, but the surface water temperatures were higher so we might expect less heating then.

The salinity is lower than the loop bottles by about 0.04psu and lower than the CTD records between 4 and 5db by about 0.03psu. There was a lot of noise in the salinity calibration data from this Portasal, and a lot of noise in the differences in the comparison. During 2005-02 the TSG was found to be low by about 0.02psu, but it has since been discovered that some data that was analyzed on the Portasal at the same time gave salinities low by >0.01psu. So a correction of 0.03psu appears to be a wise choice for both cruises with uncertainty of ±0.01psu.

The TSG Fluorescence appears to be low by a factor of about 5.

f.) Editing

There were a few suspicious records in the 2 spreadsheet files.

File 2005-01-0001.cnv was prepared with -99 entered for temperature and salinity in one record. Salinity jumped by 1psu and the primary temperature by 2Cº, while the intake temperature varied little. The first 62 records were removed since the salinity values were all 0 and the temperature values high. There were some records with no positions (record #6084-6094 and 6385-6442 in the original file); pad values were entered for the time being. If it is impossible to find positions these records should be removed. Otherwise positions should be substituted for the pad values.

For file 2005-01-0002.cnv one salinity value was replaced with -99 because salinity jumped by 0.4 with little change in temperature.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.20 Cº and +0.03psu to Temperature:Primary and Salinity:T0:C0. After calibration the two temperature channels were compared and are closer.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary, Flag, UPloy0 (flow rate)

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, add the depth of sampling to the header and change the fluorescence channel name to FLUORESCENCE and the corresponding units to VOLTS.

The sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars:

3. Touch down. PAR sensor cap may not have been removed.

10/11 downcast/upcast split because of error in con file for downcast

29. Pumps not turned on until 60m

37. 23 bottles at 2000m

48. Cast went to 100m with pumps off, returned to surface and rerun. Original file overwritten.

60. Wrong con file used so no altimeter

66. Computer problem – rotated SeaBird time/position headers

79. Hit bottom, mud on weights.

TSG

1-2. No positions were recorded; lat/long information taken from SCS files.

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-01 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 11 February 2005 End: 26 February 2005 |  |  |  |  |  |  |
| Location: Gulf of Alaska |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Robert M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 18/11/04 | Factory “ |  |  |
| Conductivity | 2399 | 19/11/04 | “ |  |  |
| Secondary Temp. | 2106 | 18/11/04 | “ |  |  |
| Secondary Cond. | 1763 | 19/11/04 | “ |  |  |
| Transmissometer | 333DR | 29/11/04 | Factory |  |  |
| Biospherical PAR | 4656 | 11/02/03 |  |  |  |
| SBE 43 DO sensor | 43 | 16/11/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| pH | 180293 | 21/06/02 |  |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2005-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 4/12/04 | Factory |  |  |
| Conductivity | 2487 | 4/12/04 | “ |  |  |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab WetstarFluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? | ? |  |  |