## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.3 | Run at 2005/04/01 12:11:07 | Compare Excel Add-in Version 2.1
| Primary | Secondary |  | Comparison Name
Ref | Sample # | Sample # |  | Sample # - Sample #
Comp 1 | Pressure | Pressure |  | Pressure - Pressure
Comp 2 | Bottle DOX | CTD DOX |  | CTD DOX - Bottle DOX
Average Size: | 0.5
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Sample # | Sample # | Sample # | (S-P) | Pressure | Pressure | Pressure | (S-P) | Bottle DOX | CTD DOX | CTD DOX | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2005-01-0001.add | 2005-01-0001.sam | 121 | 4 | 4 | 0 | 0 | 150 | 150.116 | 0.031536 | 0.11583 | 0.044 | 0.048 | 0 | 0.004
2005-01-0001.add | 2005-01-0001.sam | 121 | 5 | 5 | 0 | 0 | 124 | 123.98 | 0.034371 | -0.01989 | 0.1 | 0.04676 | 0.0014834 | -0.05324
2005-01-0001.add | 2005-01-0001.sam | 121 | 6 | 6 | 0 | 0 | 100 | 100.069 | 0.033628 | 0.06913 | 0.502 | 0.50074 | 0.0087174 | -0.0012562
2005-01-0001.add | 2005-01-0001.sam | 121 | 7 | 7 | 0 | 0 | 75 | 74.9138 | 0.027111 | -0.086235 | 3.897 | 2.8522 | 0.0043797 | -1.0448
2005-01-0001.add | 2005-01-0001.sam | 121 | 8 | 8 | 0 | 0 | 50 | 51.3413 | 0.027671 | 1.3413 | 4.361 | 3.2522 | 0.002912 | -1.1088
2005-01-0001.add | 2005-01-0001.sam | 121 | 9 | 9 | 0 | 0 | 25 | 25.1009 | 0.027476 | 0.10091 | 4.951 | 3.687 | 0.0021289 | -1.264
2005-01-0001.add | 2005-01-0001.sam | 121 | 10 | 10 | 0 | 0 | 10 | 9.9867 | 0.029191 | -0.013256 | 5.288 | 3.877 | 0.013065 | -1.411
2005-01-0001.add | 2005-01-0001.sam | 121 | 11 | 11 | 0 | 0 | 5 | 4.7859 | 0.036172 | -0.21414 | 5.596 | 4.2193 | 0.0032873 | -1.3767
2005-01-0003.add | 2005-01-0003.sam | 121 | 13 | 13 | 0 | 0 | 113 | 112.872 | 0.36499 | -0.12807 | 4.84 | 3.4827 | 0.0042317 | -1.3573
2005-01-0003.add | 2005-01-0003.sam | 121 | 14 | 14 | 0 | 0 | 100 | 99.5313 | 0.83782 | -0.46873 | 5.704 | 4.2358 | 0.0015184 | -1.4682
2005-01-0003.add | 2005-01-0003.sam | 121 | 15 | 15 | 0 | 0 | 75 | 75.7888 | 0.27621 | 0.78878 | 5.836 | 4.3923 | 0.0015558 | -1.4437
2005-01-0003.add | 2005-01-0003.sam | 121 | 16 | 16 | 0 | 0 | 75 | 75.6808 | 0.77145 | 0.68083 | 5.815 | 4.3895 | 0.0016078 | -1.4255
2005-01-0003.add | 2005-01-0003.sam | 121 | 17 | 17 | 0 | 0 | 50 | 49.7631 | 0.43775 | -0.23688 | 5.819 | 4.3972 | 0.0035285 | -1.4218
2005-01-0003.add | 2005-01-0003.sam | 121 | 18 | 18 | 0 | 0 | 25 | 24.9832 | 0.39083 | -0.016769 | 5.979 | 4.5156 | 0.0024659 | -1.4634
2005-01-0003.add | 2005-01-0003.sam | 121 | 19 | 19 | 0 | 0 | 10 | 9.3925 | 0.71696 | -0.60752 | 6.15 | 4.6644 | 0.012831 | -1.4856
2005-01-0003.add | 2005-01-0003.sam | 121 | 20 | 20 | 0 | 0 | 5 | 5.0346 | 0.27685 | 0.034562 | 6.212 | 4.6928 | 0.0017865 | -1.5192
2005-01-0006.add | 2005-01-0006.sam | 121 | 22 | 22 | 0 | 0 | 1300 | 1303.76 | 0.35288 | 3.762 | 0.567 | 0.3693 | 0.0017447 | -0.1977
2005-01-0006.add | 2005-01-0006.sam | 121 | 23 | 23 | 0 | 0 | 1000 | 999.212 | 1.6259 | -0.78778 | 0.309 | 0.21 | 0 | -0.099
2005-01-0006.add | 2005-01-0006.sam | 121 | 24 | 24 | 0 | 0 | 1000 | 997.807 | 1.2004 | -2.1928 | 0.332 | 0.20993 | 0.00057973 | -0.12207
2005-01-0006.add | 2005-01-0006.sam | 121 | 25 | 25 | 0 | 0 | 800 | 801.543 | 0.70444 | 1.5434 | 0.334 | 0.16153 | 0.0013606 | -0.17247
2005-01-0006.add | 2005-01-0006.sam | 121 | 26 | 26 | 0 | 0 | 800 | 802.404 | 0.71701 | 2.4039 | 0.371 | 0.1613 | 0.0010539 | -0.2097
2005-01-0006.add | 2005-01-0006.sam | 121 | 27 | 27 | 0 | 0 | 600 | 599.013 | 0.91443 | -0.98651 | 0.658 | 0.4346 | 0.0012006 | -0.2234
2005-01-0006.add | 2005-01-0006.sam | 121 | 28 | 28 | 0 | 0 | 400 | 398.581 | 0.17745 | -1.4193 | 1.332 | 0.92265 | 0.0023447 | -0.40936
2005-01-0006.add | 2005-01-0006.sam | 121 | 29 | 29 | 0 | 0 | 300 | 299.775 | 0.14397 | -0.22507 | 2.005 | 1.3029 | 0.0011328 | -0.70209
2005-01-0006.add | 2005-01-0006.sam | 121 | 31 | 31 | 0 | 0 | 200 | 200.878 | 0.85795 | 0.87797 | 2.713 | 1.9264 | 0.0010691 | -0.7866
2005-01-0006.add | 2005-01-0006.sam | 121 | 32 | 32 | 0 | 0 | 175 | 176.117 | 0.4941 | 1.1168 | 2.583 | 1.8416 | 0.0012006 | -0.7414
2005-01-0006.add | 2005-01-0006.sam | 121 | 33 | 33 | 0 | 0 | 150 | 153.872 | 1.0608 | 3.8722 | 2.808 | 2.0064 | 0.0013354 | -0.80164
2005-01-0006.add | 2005-01-0006.sam | 121 | 34 | 34 | 0 | 0 | 125 | 122.952 | 1.2095 | -2.0484 | 3.281 | 2.3332 | 0.0016834 | -0.94776
2005-01-0006.add | 2005-01-0006.sam | 121 | 35 | 35 | 0 | 0 | 100 | 103.636 | 0.47119 | 3.6364 | 4.613 | 3.2813 | 0.0027519 | -1.3317
2005-01-0006.add | 2005-01-0006.sam | 121 | 36 | 36 | 0 | 0 | 75 | 77.44 | 1.0386 | 2.44 | 5.689 | 4.0995 | 0.0030114 | -1.5895
2005-01-0006.add | 2005-01-0006.sam | 121 | 37 | 37 | 0 | 0 | 50 | 49.7815 | 0.76289 | -0.21853 | 6.204 | 4.4806 | 0.0012583 | -1.7234
2005-01-0006.add | 2005-01-0006.sam | 121 | 38 | 38 | 0 | 0 | 25 | 24.3834 | 0.71988 | -0.61664 | 6.167 | 4.4338 | 0.0029481 | -1.7332
2005-01-0006.add | 2005-01-0006.sam | 121 | 39 | 39 | 0 | 0 | 10 | 10.7183 | 0.73741 | 0.71827 | 6.31 | 4.5799 | 0.0018105 | -1.7302
2005-01-0006.add | 2005-01-0006.sam | 121 | 40 | 40 | 0 | 0 | 5 | 5.4036 | 0.55856 | 0.40359 | 6.294 | 4.5822 | 0.001127 | -1.7118
2005-01-0015.add | 2005-01-0015.sam | 121 | 58 | 58 | 0 | 0 | 2533 | 2532.73 | 1.2074 | -0.26587 | 1.951 | 1.3029 | 0.0022558 | -0.64814
2005-01-0015.add | 2005-01-0015.sam | 121 | 59 | 59 | 0 | 0 | 2000 | 2000.38 | 0.3458 | 0.37964 | 1.489 | 0.95083 | 0.0012803 | -0.53817
2005-01-0015.add | 2005-01-0015.sam | 121 | 60 | 60 | 0 | 0 | 1500 | 1503.05 | 0.4089 | 3.0458 | 0.696 | 0.478 | 0 | -0.218
2005-01-0015.add | 2005-01-0015.sam | 121 | 61 | 61 | 0 | 0 | 1250 | 1251.74 | 0.60806 | 1.7356 | 0.412 | 0.2704 | 0.0012005 | -0.1416
2005-01-0015.add | 2005-01-0015.sam | 121 | 62 | 62 | 0 | 0 | 1000 | 998.815 | 0.48144 | -1.1854 | 0.301 | 0.18318 | 0.0023805 | -0.11782
2005-01-0015.add | 2005-01-0015.sam | 121 | 63 | 63 | 0 | 0 | 800 | 800.15 | 0.12118 | 0.15008 | 0.275 | 0.17374 | 0.00099807 | -0.10126
2005-01-0015.add | 2005-01-0015.sam | 63 | 64 | 64 | 0 | 0 | 600 | 599.996 | 0.23264 | -0.0036011 | 0.577 | 0.33051 | 0.001674 | -0.24649
2005-01-0015.add | 2005-01-0015.sam | 120 | 65 | 65 | 0 | 0 | 600 | 601.442 | 0.47932 | 1.4417 | 0.599 | 0.33226 | 0.0014866 | -0.26674
2005-01-0015.add | 2005-01-0015.sam | 121 | 66 | 66 | 0 | 0 | 400 | 399.031 | 0.75268 | -0.96942 | 1.545 | 0.99696 | 0.0016852 | -0.54804
2005-01-0015.add | 2005-01-0015.sam | 121 | 67 | 67 | 0 | 0 | 300 | 300.776 | 0.97634 | 0.77582 | 2.559 | 1.6957 | 0.002548 | -0.86334
2005-01-0015.add | 2005-01-0015.sam | 121 | 68 | 68 | 0 | 0 | 250 | 250.259 | 0.53014 | 0.25896 | 3.049 | 2.0467 | 0.0023078 | -1.0023
2005-01-0015.add | 2005-01-0015.sam | 121 | 69 | 69 | 0 | 0 | 200 | 199.679 | 0.23447 | -0.32088 | 3.775 | 2.5483 | 0.0018423 | -1.2267
2005-01-0015.add | 2005-01-0015.sam | 121 | 70 | 70 | 0 | 0 | 175 | 175.212 | 0.13585 | 0.2122 | 4.189 | 2.8469 | 0.0011569 | -1.3421
2005-01-0015.add | 2005-01-0015.sam | 121 | 71 | 71 | 0 | 0 | 150 | 150.493 | 0.59836 | 0.49258 | 4.505 | 3.0726 | 0.0018841 | -1.4324
2005-01-0015.add | 2005-01-0015.sam | 121 | 72 | 72 | 0 | 0 | 125 | 125.563 | 0.46488 | 0.56315 | 6.439 | 3.1943 | 0.0016615 | -3.2447
2005-01-0015.add | 2005-01-0015.sam | 121 | 73 | 73 | 0 | 0 | 100 | 100.487 | 0.78314 | 0.48656 | 4.644 | 3.1975 | 0.0013041 | -1.4465
2005-01-0015.add | 2005-01-0015.sam | 121 | 74 | 74 | 0 | 0 | 75 | 75.8665 | 0.90257 | 0.86648 | 6.554 | 4.4961 | 0.0029635 | -2.0579
2005-01-0015.add | 2005-01-0015.sam | 121 | 75 | 75 | 0 | 0 | 50 | 49.2447 | 0.31268 | -0.75528 | 7.184 | 4.5756 | 0.0016338 | -2.6084
2005-01-0015.add | 2005-01-0015.sam | 121 | 76 | 76 | 0 | 0 | 25 | 24.716 | 1.0151 | -0.28402 | 6.629 | 4.5983 | 0.0018676 | -2.0307
2005-01-0015.add | 2005-01-0015.sam | 121 | 77 | 77 | 0 | 0 | 10 | 9.0068 | 0.67514 | -0.99325 | 6.638 | 4.5999 | 0.0019402 | -2.0381
2005-01-0015.add | 2005-01-0015.sam | 121 | 78 | 78 | 0 | 0 | 5 | 5.0642 | 0.2947 | 0.06419 | 6.588 | 4.6054 | 0.0016601 | -1.9826
2005-01-0021.add | 2005-01-0021.sam | 121 | 82 | 82 | 0 | 0 | 3236 | 3236.62 | 0.64341 | 0.62231 | 2.49 | 1.5379 | 0.0019487 | -0.95205
2005-01-0021.add | 2005-01-0021.sam | 121 | 83 | 83 | 0 | 0 | 3000 | 2998.97 | 0.21126 | -1.0339 | 2.343 | 1.5143 | 0.0013079 | -0.8287
2005-01-0021.add | 2005-01-0021.sam | 121 | 84 | 84 | 0 | 0 | 2500 | 2497 | 1.1161 | -2.9988 | 2.017 | 1.2764 | 0.0014111 | -0.7406
2005-01-0021.add | 2005-01-0021.sam | 121 | 85 | 85 | 0 | 0 | 2000 | 2002.86 | 0.98787 | 2.8608 | 1.4 | 0.89347 | 0.0028754 | -0.50653
2005-01-0021.add | 2005-01-0021.sam | 121 | 86 | 86 | 0 | 0 | 1500 | 1501.1 | 0.63182 | 1.1025 | 0.748 | 0.43764 | 0.0024118 | -0.31036
2005-01-0021.add | 2005-01-0021.sam | 121 | 87 | 87 | 0 | 0 | 1250 | 1251.58 | 0.73213 | 1.5765 | 0.441 | 0.25452 | 0.0027931 | -0.18648
2005-01-0021.add | 2005-01-0021.sam | 121 | 88 | 88 | 0 | 0 | 1000 | 1001.11 | 0.73318 | 1.1101 | 0.324 | 0.1806 | 0.0012006 | -0.1434
2005-01-0021.add | 2005-01-0021.sam | 121 | 89 | 89 | 0 | 0 | 1000 | 1001.51 | 0.61447 | 1.5127 | 0.311 | 0.18014 | 0.0016498 | -0.13086
2005-01-0021.add | 2005-01-0021.sam | 121 | 90 | 90 | 0 | 0 | 800 | 800.83 | 0.17902 | 0.82983 | 0.629 | 0.18255 | 0.0014434 | -0.44645
2005-01-0021.add | 2005-01-0021.sam | 121 | 91 | 91 | 0 | 0 | 600 | 600.214 | 0.29665 | 0.21435 | 0.477 | 0.28279 | 0.0016017 | -0.19421
2005-01-0021.add | 2005-01-0021.sam | 121 | 92 | 92 | 0 | 0 | 400 | 401.523 | 0.63925 | 1.5229 | 1.156 | 0.74521 | 0.0019745 | -0.41079
2005-01-0021.add | 2005-01-0021.sam | 121 | 93 | 93 | 0 | 0 | 300 | 300.986 | 0.34858 | 0.98602 | 1.681 | 1.0969 | 0.0017547 | -0.58407
2005-01-0021.add | 2005-01-0021.sam | 121 | 94 | 94 | 0 | 0 | 250 | 248.325 | 0.7745 | -1.6754 | 2.138 | 1.3515 | 0.0014383 | -0.78649
2005-01-0021.add | 2005-01-0021.sam | 121 | 95 | 95 | 0 | 0 | 200 | 199.613 | 0.39718 | -0.38728 | 2.786 | 1.8271 | 0.0020829 | -0.95894
2005-01-0021.add | 2005-01-0021.sam | 121 | 96 | 96 | 0 | 0 | 175 | 175.175 | 0.10193 | 0.1747 | 2.824 | 1.9475 | 0.0013421 | -0.87647
2005-01-0021.add | 2005-01-0021.sam | 121 | 97 | 97 | 0 | 0 | 150 | 148.3 | 0.46924 | -1.7003 | 3.775 | 2.4851 | 0.0016243 | -1.2899
2005-01-0021.add | 2005-01-0021.sam | 121 | 98 | 98 | 0 | 0 | 125 | 123.554 | 0.55859 | -1.4458 | 4.361 | 2.9115 | 0.014906 | -1.4495
2005-01-0021.add | 2005-01-0021.sam | 121 | 99 | 99 | 0 | 0 | 100 | 100.957 | 0.94996 | 0.95711 | 5.89 | 3.944 | 0.0056295 | -1.946
2005-01-0021.add | 2005-01-0021.sam | 121 | 100 | 100 | 0 | 0 | 75 | 75.0847 | 0.34425 | 0.084717 | 6.629 | 4.4955 | 0.0013044 | -2.1335
2005-01-0021.add | 2005-01-0021.sam | 121 | 101 | 101 | 0 | 0 | 50 | 46.6954 | 0.38691 | -3.3046 | 6.615 | 4.5092 | 0.0018394 | -2.1058
2005-01-0021.add | 2005-01-0021.sam | 121 | 102 | 102 | 0 | 0 | 25 | 23.1122 | 0.85177 | -1.8878 | 6.582 | 4.5181 | 0.0014259 | -2.0639
2005-01-0021.add | 2005-01-0021.sam | 121 | 103 | 103 | 0 | 0 | 10 | 9.6326 | 0.55622 | -0.36743 | 6.616 | 4.5223 | 0.0015575 | -2.0937
2005-01-0021.add | 2005-01-0021.sam | 121 | 104 | 104 | 0 | 0 | 5 | 4.3149 | 0.52709 | -0.68509 | 6.635 | 4.524 | 0.0012706 | -2.1109
2005-01-0029.add | 2005-01-0029.sam | 121 | 122 | 122 | 0 | 0 | 3686 | 3687.26 | 0.36633 | 1.259 | 2.905 | 1.8111 | 0.00089619 | -1.0939
2005-01-0029.add | 2005-01-0029.sam | 121 | 123 | 123 | 0 | 0 | 3500 | 3500.58 | 0.1461 | 0.58008 | 2.717 | 1.7655 | 0.0015809 | -0.95145
2005-01-0029.add | 2005-01-0029.sam | 121 | 124 | 124 | 0 | 0 | 3000 | 3000.69 | 0.44202 | 0.69385 | 2.321 | 1.5257 | 0.0012493 | -0.7953
2005-01-0029.add | 2005-01-0029.sam | 121 | 125 | 125 | 0 | 0 | 2500 | 2501 | 0.61918 | 1.0012 | 1.829 | 1.2279 | 0.00076817 | -0.6011
2005-01-0029.add | 2005-01-0029.sam | 121 | 126 | 126 | 0 | 0 | 2000 | 2001.35 | 0.55271 | 1.3527 | 1.284 | 0.8526 | 0.0026692 | -0.4314
2005-01-0029.add | 2005-01-0029.sam | 121 | 127 | 127 | 0 | 0 | 1500 | 1499.45 | 0.144 | -0.55054 | 0.54 | 0.36298 | 0.0024562 | -0.17702
2005-01-0029.add | 2005-01-0029.sam | 121 | 128 | 128 | 0 | 0 | 1250 | 1251.36 | 0.29601 | 1.3634 | 0.352 | 0.2222 | 0.0021433 | -0.1298
2005-01-0029.add | 2005-01-0029.sam | 121 | 129 | 129 | 0 | 0 | 1250 | 1250.75 | 0.11873 | 0.74548 | 0.313 | 0.22228 | 0.0021917 | -0.090719
2005-01-0029.add | 2005-01-0029.sam | 121 | 130 | 130 | 0 | 0 | 1000 | 1000.08 | 0.34883 | 0.084961 | 0.256 | 0.186 | 0 | -0.07
2005-01-0029.add | 2005-01-0029.sam | 121 | 131 | 131 | 0 | 0 | 800 | 798.738 | 0.43931 | -1.2622 | 0.303 | 0.22246 | 0.0019322 | -0.080546
2005-01-0029.add | 2005-01-0029.sam | 121 | 132 | 132 | 0 | 0 | 600 | 599.267 | 0.85277 | -0.73328 | 0.713 | 0.47631 | 0.0014436 | -0.23669
2005-01-0029.add | 2005-01-0029.sam | 121 | 133 | 133 | 0 | 0 | 400 | 401.662 | 0.38839 | 1.6621 | 1.693 | 1.1175 | 0.0017084 | -0.57549
2005-01-0029.add | 2005-01-0029.sam | 121 | 134 | 134 | 0 | 0 | 300 | 300.21 | 0.23693 | 0.21014 | 2.417 | 1.569 | 0.0039453 | -0.84797
2005-01-0029.add | 2005-01-0029.sam | 121 | 135 | 135 | 0 | 0 | 250 | 248.47 | 0.45633 | -1.5302 | 2.896 | 1.9172 | 0.0018221 | -0.97879
2005-01-0029.add | 2005-01-0029.sam | 121 | 136 | 136 | 0 | 0 | 200 | 200.389 | 0.89934 | 0.38884 | 3.898 | 2.5827 | 0.0021456 | -1.3153
2005-01-0029.add | 2005-01-0029.sam | 121 | 137 | 137 | 0 | 0 | 175 | 173.467 | 0.53286 | -1.5333 | 4.546 | 3.1111 | 0.0014921 | -1.4349
2005-01-0029.add | 2005-01-0029.sam | 121 | 138 | 138 | 0 | 0 | 150 | 149.968 | 0.64972 | -0.031616 | 5.172 | 3.4436 | 0.0019536 | -1.7284
2005-01-0029.add | 2005-01-0029.sam | 121 | 139 | 139 | 0 | 0 | 125 | 124.238 | 0.25195 | -0.76163 | 5.513 | 3.6641 | 0.001536 | -1.8489
2005-01-0029.add | 2005-01-0029.sam | 121 | 140 | 140 | 0 | 0 | 100 | 101.299 | 0.11158 | 1.2993 | 6.305 | 4.197 | 0.0031303 | -2.108
2005-01-0029.add | 2005-01-0029.sam | 121 | 141 | 141 | 0 | 0 | 75 | 75.0294 | 0.43513 | 0.029358 | 6.779 | 4.5605 | 0.0015763 | -2.2185
2005-01-0029.add | 2005-01-0029.sam | 121 | 142 | 142 | 0 | 0 | 50 | 49.558 | 0.55966 | -0.44204 | 6.772 | 4.5715 | 0.0010493 | -2.2005
2005-01-0029.add | 2005-01-0029.sam | 121 | 143 | 143 | 0 | 0 | 25 | 25.6428 | 0.63129 | 0.64284 | 6.762 | 4.5812 | 0.0013225 | -2.1808
2005-01-0029.add | 2005-01-0029.sam | 121 | 144 | 144 | 0 | 0 | 10 | 10.5445 | 0.34888 | 0.5445 | 6.764 | 4.5865 | 0.0018528 | -2.1774
2005-01-0029.add | 2005-01-0029.sam | 121 | 145 | 145 | 0 | 0 | 5 | 4.9337 | 0.23253 | -0.066248 | 6.755 | 4.5886 | 0.0013785 | -2.1664
2005-01-0039.add | 2005-01-0039.sam | 121 | 186 | 186 | 0 | 0 | 4029 | 4028.8 | 0.32073 | -0.20166 | 3.039 | 1.9311 | 0.0012286 | -1.1079
2005-01-0039.add | 2005-01-0039.sam | 121 | 187 | 187 | 0 | 0 | 3500 | 3500.27 | 0.6253 | 0.26562 | 2.857 | 1.8347 | 0.002743 | -1.0223
2005-01-0039.add | 2005-01-0039.sam | 79 | 188 | 188 | 0 | 0 | 3000 | 2999.04 | 0.22154 | -0.96387 | 2.454 | 1.5906 | 0.0026761 | -0.86338
2005-01-0039.add | 2005-01-0039.sam | 120 | 189 | 189 | 0 | 0 | 3000 | 2999.58 | 0.28575 | -0.41626 | 2.443 | 1.5897 | 0.002941 | -0.85335
2005-01-0039.add | 2005-01-0039.sam | 121 | 190 | 190 | 0 | 0 | 2500 | 2501.35 | 0.66389 | 1.3479 | 1.91 | 1.2279 | 0.00078071 | -0.68212
2005-01-0039.add | 2005-01-0039.sam | 121 | 191 | 191 | 0 | 0 | 2000 | 1999.44 | 0.19311 | -0.55872 | 1.216 | 0.79214 | 0.0074926 | -0.42386
2005-01-0039.add | 2005-01-0039.sam | 121 | 192 | 192 | 0 | 0 | 1500 | 1500.66 | 0.76893 | 0.65515 | 0.509 | 0.34511 | 0.0023587 | -0.16389
2005-01-0039.add | 2005-01-0039.sam | 121 | 193 | 193 | 0 | 0 | 1250 | 1250.62 | 0.45739 | 0.62097 | 0.308 | 0.22352 | 0.0025103 | -0.084479
2005-01-0039.add | 2005-01-0039.sam | 121 | 194 | 194 | 0 | 0 | 1000 | 1001.12 | 0.88904 | 1.1205 | 0.267 | 0.1986 | 0.0012006 | -0.068397
2005-01-0039.add | 2005-01-0039.sam | 121 | 195 | 195 | 0 | 0 | 800 | 801.782 | 0.14669 | 1.7822 | 0.33 | 0.22269 | 0.0019834 | -0.10731
2005-01-0039.add | 2005-01-0039.sam | 121 | 196 | 196 | 0 | 0 | 600 | 600.598 | 0.78588 | 0.59802 | 0.472 | 0.32331 | 0.0021089 | -0.14869
2005-01-0039.add | 2005-01-0039.sam | 121 | 197 | 197 | 0 | 0 | 400 | 399.57 | 0.083243 | -0.42956 | 1.3 | 0.80869 | 0.013204 | -0.49131
2005-01-0039.add | 2005-01-0039.sam | 121 | 198 | 198 | 0 | 0 | 300 | 302.34 | 0.38527 | 2.3397 | 1.939 | 1.2606 | 0.0018827 | -0.67841
2005-01-0039.add | 2005-01-0039.sam | 121 | 199 | 199 | 0 | 0 | 250 | 250.15 | 0.38326 | 0.15016 | 2.562 | 1.6433 | 0.0019382 | -0.91867
2005-01-0039.add | 2005-01-0039.sam | 121 | 200 | 200 | 0 | 0 | 200 | 202.2 | 0.40927 | 2.2003 | 3.639 | 2.3511 | 0.0008104 | -1.2879
2005-01-0039.add | 2005-01-0039.sam | 121 | 201 | 201 | 0 | 0 | 175 | 175.838 | 0.54872 | 0.83826 | 3.815 | 2.604 | 0.0016973 | -1.211
2005-01-0039.add | 2005-01-0039.sam | 121 | 202 | 202 | 0 | 0 | 150 | 150.782 | 0.82894 | 0.78172 | 4.392 | 2.871 | 0.0032544 | -1.521
2005-01-0039.add | 2005-01-0039.sam | 121 | 203 | 203 | 0 | 0 | 125 | 124.476 | 0.32172 | -0.52409 | 5.028 | 3.3187 | 0.0023088 | -1.7093
2005-01-0039.add | 2005-01-0039.sam | 121 | 204 | 204 | 0 | 0 | 100 | 100.209 | 0.54753 | 0.20907 | 6.763 | 4.4733 | 0.0024342 | -2.2897
2005-01-0039.add | 2005-01-0039.sam | 121 | 205 | 205 | 0 | 0 | 75 | 73.7679 | 0.2861 | -1.2322 | 6.784 | 4.5131 | 0.001646 | -2.2709
2005-01-0039.add | 2005-01-0039.sam | 121 | 206 | 206 | 0 | 0 | 50 | 48.2123 | 0.56667 | -1.7876 | 6.817 | 4.5501 | 0.0016361 | -2.2669
2005-01-0039.add | 2005-01-0039.sam | 121 | 207 | 207 | 0 | 0 | 25 | 24.2342 | 0.38071 | -0.76578 | 6.786 | 4.5613 | 0.001885 | -2.2247
2005-01-0039.add | 2005-01-0039.sam | 121 | 208 | 208 | 0 | 0 | 10 | 10.0164 | 0.076046 | 0.016413 | 6.817 | 4.5683 | 0.0016469 | -2.2487
2005-01-0039.add | 2005-01-0039.sam | 121 | 209 | 209 | 0 | 0 | 1 | 5.2329 | 0.082847 | 4.2329 | 6.834 | 4.5684 | 0.0017776 | -2.2656
2005-01-0057.add | 2005-01-0057.sam | 121 | 258 | 258 | 0 | 0 | 4313 | 4310.57 | 0.39195 | -2.4321 | 3.354 | 2.0893 | 0.0030343 | -1.2647
2005-01-0057.add | 2005-01-0057.sam | 121 | 259 | 259 | 0 | 0 | 4000 | 3999.69 | 0.63828 | -0.30737 | 3.153 | 2.0264 | 0.0015592 | -1.1266
2005-01-0057.add | 2005-01-0057.sam | 121 | 260 | 260 | 0 | 0 | 3500 | 3501.07 | 0.30709 | 1.0686 | 2.89 | 1.8585 | 0.0029751 | -1.0315
2005-01-0057.add | 2005-01-0057.sam | 121 | 261 | 261 | 0 | 0 | 3000 | 2998.49 | 0.15812 | -1.5078 | 2.505 | 1.6089 | 0.0018999 | -0.89612
2005-01-0057.add | 2005-01-0057.sam | 121 | 262 | 262 | 0 | 0 | 2500 | 2499.86 | 0.21967 | -0.13965 | 1.899 | 1.2276 | 0.0013806 | -0.67139
2005-01-0057.add | 2005-01-0057.sam | 121 | 263 | 263 | 0 | 0 | 2000 | 2001.45 | 0.63236 | 1.4532 | 1.293 | 0.8383 | 0.0023897 | -0.4547
2005-01-0057.add | 2005-01-0057.sam | 121 | 264 | 264 | 0 | 0 | 1500 | 1500.75 | 0.29274 | 0.74805 | 0.623 | 0.42333 | 0.001106 | -0.19967
2005-01-0057.add | 2005-01-0057.sam | 121 | 265 | 265 | 0 | 0 | 1250 | 1248.49 | 0.28315 | -1.5132 | 0.434 | 0.2852 | 0.00087197 | -0.1488
2005-01-0057.add | 2005-01-0057.sam | 121 | 266 | 266 | 0 | 0 | 1000 | 999.947 | 1.3105 | -0.053162 | 0.378 | 0.24974 | 0.0018643 | -0.12826
2005-01-0057.add | 2005-01-0057.sam | 121 | 267 | 267 | 0 | 0 | 820 | 828.006 | 2.0944 | 8.0057 | 0.375 | 0.25752 | 0.0021451 | -0.11748
2005-01-0057.add | 2005-01-0057.sam | 121 | 268 | 268 | 0 | 0 | 600 | 598.926 | 0.68876 | -1.0737 | 0.455 | 0.30445 | 0.0013724 | -0.15055
2005-01-0057.add | 2005-01-0057.sam | 121 | 269 | 269 | 0 | 0 | 400 | 402.571 | 0.61032 | 2.571 | 0.929 | 0.57207 | 0.0015692 | -0.35693
2005-01-0057.add | 2005-01-0057.sam | 121 | 270 | 270 | 0 | 0 | 300 | 300.608 | 0.75269 | 0.60803 | 1.406 | 0.87831 | 0.002141 | -0.52769
2005-01-0057.add | 2005-01-0057.sam | 121 | 271 | 271 | 0 | 0 | 250 | 250.852 | 0.22462 | 0.85242 | 1.949 | 1.2268 | 0.00079044 | -0.72222
2005-01-0057.add | 2005-01-0057.sam | 121 | 272 | 272 | 0 | 0 | 200 | 202.783 | 1.25 | 2.783 | 2.57 | 1.7213 | 0.005514 | -0.84868
2005-01-0057.add | 2005-01-0057.sam | 121 | 273 | 273 | 0 | 0 | 175 | 175.01 | 1.3186 | 0.010483 | 3.29 | 2.0763 | 0.0055284 | -1.2137
2005-01-0057.add | 2005-01-0057.sam | 121 | 274 | 274 | 0 | 0 | 150 | 149.167 | 0.46681 | -0.83347 | 3.888 | 2.5045 | 0.0017083 | -1.3835
2005-01-0057.add | 2005-01-0057.sam | 121 | 275 | 275 | 0 | 0 | 125 | 123.53 | 0.7145 | -1.4701 | 5.292 | 3.4371 | 0.0025153 | -1.8549
2005-01-0057.add | 2005-01-0057.sam | 121 | 276 | 276 | 0 | 0 | 100 | 98.4833 | 0.23182 | -1.5167 | 6.901 | 4.5215 | 0.0025791 | -2.3795
2005-01-0057.add | 2005-01-0057.sam | 121 | 277 | 277 | 0 | 0 | 75 | 74.6983 | 0.76241 | -0.30167 | 6.975 | 4.5762 | 0.0024665 | -2.3988
2005-01-0057.add | 2005-01-0057.sam | 121 | 278 | 278 | 0 | 0 | 50 | 50.6391 | 0.4675 | 0.63911 | 6.912 | 4.593 | 0.0017391 | -2.319
2005-01-0057.add | 2005-01-0057.sam | 121 | 279 | 279 | 0 | 0 | 25 | 25.6736 | 0.4654 | 0.67356 | 6.954 | 4.6045 | 0.0015811 | -2.3495
2005-01-0057.add | 2005-01-0057.sam | 121 | 280 | 280 | 0 | 0 | 10 | 11.6197 | 0.3429 | 1.6197 | 6.966 | 4.609 | 0.0016146 | -2.357
2005-01-0057.add | 2005-01-0057.sam | 121 | 281 | 281 | 0 | 0 | 5 | 4.6998 | 0.14191 | -0.30022 | 6.954 | 4.6136 | 0.0015638 | -2.3403
2005-01-0062.add | 2005-01-0062.sam | 121 | 296 | 296 | 0 | 0 | 2000 | 2001.13 | 1.0251 | 1.1251 | 1.267 | 0.78446 | 0.0014434 | -0.48255
2005-01-0062.add | 2005-01-0062.sam | 121 | 297 | 297 | 0 | 0 | 1500 | 1500.21 | 0.34241 | 0.20837 | 0.529 | 0.34697 | 0.0020328 | -0.18203
2005-01-0062.add | 2005-01-0062.sam | 121 | 298 | 298 | 0 | 0 | 1250 | 1248.27 | 0.76804 | -1.7306 | 0.43 | 0.26829 | 0.0018772 | -0.16171
2005-01-0062.add | 2005-01-0062.sam | 121 | 299 | 299 | 0 | 0 | 1000 | 998.363 | 0.99065 | -1.6374 | 0.39 | 0.25443 | 0.0023303 | -0.13557
2005-01-0062.add | 2005-01-0062.sam | 121 | 300 | 300 | 0 | 0 | 800 | 799.043 | 0.57849 | -0.95666 | 0.558 | 0.32993 | 0.00072727 | -0.22807
2005-01-0062.add | 2005-01-0062.sam | 121 | 301 | 301 | 0 | 0 | 600 | 599.199 | 0.57864 | -0.80145 | 0.772 | 0.45436 | 0.0007746 | -0.31764
2005-01-0062.add | 2005-01-0062.sam | 121 | 302 | 302 | 0 | 0 | 400 | 399.813 | 0.79705 | -0.18658 | 1.194 | 0.74686 | 0.0017287 | -0.44714
2005-01-0062.add | 2005-01-0062.sam | 121 | 303 | 303 | 0 | 0 | 300 | 302.534 | 1.5837 | 2.5336 | 1.945 | 1.2119 | 0.001771 | -0.73312
2005-01-0062.add | 2005-01-0062.sam | 121 | 304 | 304 | 0 | 0 | 250 | 250.467 | 0.38517 | 0.46663 | 2.809 | 1.807 | 0.003248 | -1.002
2005-01-0062.add | 2005-01-0062.sam | 121 | 305 | 305 | 0 | 0 | 200 | 201.606 | 0.55856 | 1.6057 | 4.283 | 2.7778 | 0.0023032 | -1.5052
2005-01-0062.add | 2005-01-0062.sam | 121 | 306 | 306 | 0 | 0 | 175 | 176.595 | 1.0024 | 1.595 | 4.887 | 3.1532 | 0.0023219 | -1.7338
2005-01-0062.add | 2005-01-0062.sam | 121 | 307 | 307 | 0 | 0 | 150 | 149.272 | 0.62809 | -0.72841 | 5.666 | 3.6766 | 0.0046223 | -1.9894
2005-01-0062.add | 2005-01-0062.sam | 121 | 308 | 308 | 0 | 0 | 125 | 124.915 | 0.28729 | -0.085297 | 6.629 | 4.3459 | 0.0022517 | -2.2831
2005-01-0062.add | 2005-01-0062.sam | 121 | 309 | 309 | 0 | 0 | 100 | 101.107 | 0.70513 | 1.107 | 6.729 | 4.4501 | 0.013204 | -2.2789
2005-01-0062.add | 2005-01-0062.sam | 121 | 310 | 310 | 0 | 0 | 75 | 72.8961 | 1.519 | -2.1039 | 6.651 | 4.5979 | 0.0017311 | -2.0531
2005-01-0062.add | 2005-01-0062.sam | 121 | 311 | 311 | 0 | 0 | 50 | 50.1174 | 0.30115 | 0.11736 | 6.886 | 4.6012 | 0.0015276 | -2.2848
2005-01-0062.add | 2005-01-0062.sam | 121 | 312 | 312 | 0 | 0 | 25 | 25.4089 | 1.2266 | 0.40886 | 6.893 | 4.6052 | 0.0019398 | -2.2878
2005-01-0062.add | 2005-01-0062.sam | 121 | 313 | 313 | 0 | 0 | 10 | 10.7292 | 0.4704 | 0.7292 | 6.885 | 4.6101 | 0.001881 | -2.2749
2005-01-0062.add | 2005-01-0062.sam | 121 | 314 | 314 | 0 | 0 | 5 | 5.3663 | 0.18718 | 0.36628 | 6.884 | 4.6127 | 0.001021 | -2.2713
2005-01-0065.add | 2005-01-0065.sam | 121 | 315 | 315 | 0 | 0 | 4880 | 4880.81 | 0.20966 | 0.81055 | 3.185 | 2.0817 | 0.0024677 | -1.1033
2005-01-0065.add | 2005-01-0065.sam | 121 | 316 | 316 | 0 | 0 | 4000 | 4001.56 | 0.40515 | 1.5637 | 3.22 | 2.0068 | 0.0014182 | -1.2132
2005-01-0065.add | 2005-01-0065.sam | 121 | 317 | 317 | 0 | 0 | 3500 | 3499.19 | 0.50342 | -0.80615 | 2.99 | 1.8758 | 0.0020761 | -1.1142
2005-01-0065.add | 2005-01-0065.sam | 121 | 318 | 318 | 0 | 0 | 3000 | 2998.99 | 0.25781 | -1.0093 | 2.608 | 1.5924 | 0.0025964 | -1.0156
2005-01-0065.add | 2005-01-0065.sam | 121 | 319 | 319 | 0 | 0 | 2500 | 2501.33 | 0.088744 | 1.3301 | 1.921 | 1.2051 | 0.002893 | -0.71593
2005-01-0065.add | 2005-01-0065.sam | 121 | 320 | 320 | 0 | 0 | 2000 | 2000.25 | 0.13442 | 0.25146 | 1.221 | 0.7649 | 0.0026344 | -0.4561
2005-01-0065.add | 2005-01-0065.sam | 121 | 321 | 321 | 0 | 0 | 1500 | 1498.98 | 1.1157 | -1.0192 | 0.511 | 0.329 | 0 | -0.182
2005-01-0065.add | 2005-01-0065.sam | 121 | 322 | 322 | 0 | 0 | 1250 | 1251.03 | 0.60304 | 1.0312 | 0.401 | 0.23434 | 0.0014919 | -0.16666
2005-01-0065.add | 2005-01-0065.sam | 121 | 323 | 323 | 0 | 0 | 1000 | 1001.73 | 0.70181 | 1.7333 | 0.408 | 0.22027 | 0.0015492 | -0.18773
2005-01-0065.add | 2005-01-0065.sam | 121 | 324 | 324 | 0 | 0 | 800 | 798.45 | 0.87808 | -1.5501 | 0.51 | 0.29217 | 0.00079945 | -0.21784
2005-01-0065.add | 2005-01-0065.sam | 121 | 325 | 325 | 0 | 0 | 600 | 603.792 | 0.26731 | 3.7917 | 0.735 | 0.45698 | 0.0017534 | -0.27803
2005-01-0065.add | 2005-01-0065.sam | 121 | 326 | 326 | 0 | 0 | 400 | 400.409 | 0.29807 | 0.40942 | 1.324 | 0.8295 | 0.00059335 | -0.4945
2005-01-0065.add | 2005-01-0065.sam | 121 | 327 | 327 | 0 | 0 | 300 | 299.06 | 0.15771 | -0.94 | 2.01 | 1.2576 | 0.0043092 | -0.75237
2005-01-0065.add | 2005-01-0065.sam | 121 | 328 | 328 | 0 | 0 | 250 | 250.306 | 0.54952 | 0.30617 | 2.802 | 1.7763 | 0.0030435 | -1.0257
2005-01-0065.add | 2005-01-0065.sam | 121 | 329 | 329 | 0 | 0 | 200 | 202.134 | 0.4043 | 2.1338 | 4.558 | 3.0289 | 0.0020587 | -1.5291
2005-01-0065.add | 2005-01-0065.sam | 121 | 330 | 330 | 0 | 0 | 175 | 176.465 | 0.42054 | 1.4651 | 5.718 | 3.6728 | 0.0047478 | -2.0452
2005-01-0065.add | 2005-01-0065.sam | 121 | 331 | 331 | 0 | 0 | 150 | 151.844 | 0.53805 | 1.844 | 6.421 | 4.1619 | 0.0033134 | -2.2591
2005-01-0065.add | 2005-01-0065.sam | 121 | 332 | 332 | 0 | 0 | 125 | 123.934 | 0.52989 | -1.0661 | 6.596 | 4.3362 | 0.0065157 | -2.2598
2005-01-0065.add | 2005-01-0065.sam | 121 | 333 | 333 | 0 | 0 | 100 | 100.078 | 0.54369 | 0.078339 | 6.761 | 4.4581 | 0.00087248 | -2.3028
2005-01-0065.add | 2005-01-0065.sam | 121 | 334 | 334 | 0 | 0 | 75 | 75.1951 | 0.32323 | 0.19512 | 6.768 | 4.5042 | 0.0016126 | -2.2638
2005-01-0065.add | 2005-01-0065.sam | 121 | 335 | 335 | 0 | 0 | 50 | 50.1043 | 0.31573 | 0.10427 | 6.8 | 4.5157 | 0.0013758 | -2.2843
2005-01-0065.add | 2005-01-0065.sam | 121 | 336 | 336 | 0 | 0 | 25 | 26.2796 | 0.65328 | 1.2796 | 6.822 | 4.5275 | 0.0016689 | -2.2945
2005-01-0065.add | 2005-01-0065.sam | 121 | 337 | 337 | 0 | 0 | 10 | 10.422 | 0.4802 | 0.42198 | 6.822 | 4.5323 | 0.0015943 | -2.2897
2005-01-0065.add | 2005-01-0065.sam | 121 | 338 | 338 | 0 | 0 | 5 | 5.39 | 0.29041 | 0.39003 | 6.827 | 4.5289 | 0.0016854 | -2.2981
2005-01-0068.add | 2005-01-0068.sam | 121 | 340 | 340 | 0 | 0 | 4569 | 4569.91 | 0.23267 | 0.91113 | 3.219 | 2.0236 | 0.0043433 | -1.1954
2005-01-0068.add | 2005-01-0068.sam | 0 | 341 | 0 | 0 | -341 | 4000 | 0 | 0 | -4000 | 3.175 | 0 | 0 | -3.175
2005-01-0068.add | 2005-01-0068.sam | 121 | 342 | 342 | 0 | 0 | 3500 | 3501.73 | 0.17905 | 1.7283 | 2.917 | 1.8446 | 0.001982 | -1.0723
2005-01-0068.add | 2005-01-0068.sam | 121 | 343 | 343 | 0 | 0 | 3000 | 2990.45 | 0.1111 | -9.5525 | 2.502 | 1.5831 | 0.002152 | -0.91895
2005-01-0068.add | 2005-01-0068.sam | 121 | 344 | 344 | 0 | 0 | 2500 | 2488.34 | 1.0528 | -11.6555 | 1.866 | 1.1926 | 0.0022141 | -0.67337
2005-01-0068.add | 2005-01-0068.sam | 121 | 345 | 345 | 0 | 0 | 2000 | 2009.89 | 0.69746 | 9.8904 | 1.238 | 0.78359 | 0.0022718 | -0.45441
2005-01-0068.add | 2005-01-0068.sam | 121 | 346 | 346 | 0 | 0 | 1500 | 1507.32 | 0.46572 | 7.318 | 0.462 | 0.30934 | 0.0017008 | -0.15266
2005-01-0068.add | 2005-01-0068.sam | 121 | 347 | 347 | 0 | 0 | 1250 | 1260.5 | 0.19854 | 10.4989 | 0.303 | 0.19877 | 0.00093773 | -0.10423
2005-01-0068.add | 2005-01-0068.sam | 121 | 348 | 348 | 0 | 0 | 1000 | 1002.15 | 0.53413 | 2.152 | 0.312 | 0.21537 | 0.0014613 | -0.096628
2005-01-0068.add | 2005-01-0068.sam | 121 | 349 | 349 | 0 | 0 | 800 | 798.81 | 0.81482 | -1.1902 | 0.542 | 0.34909 | 0.00249 | -0.19291
2005-01-0068.add | 2005-01-0068.sam | 121 | 350 | 350 | 0 | 0 | 600 | 593.42 | 0.70698 | -6.5796 | 1.021 | 0.62656 | 0.0032375 | -0.39444
2005-01-0068.add | 2005-01-0068.sam | 121 | 351 | 351 | 0 | 0 | 400 | 399.468 | 0.43363 | -0.53192 | 2.105 | 1.3231 | 0.0017168 | -0.78195
2005-01-0068.add | 2005-01-0068.sam | 121 | 352 | 352 | 0 | 0 | 300 | 297.774 | 0.29935 | -2.2262 | 3.786 | 2.4287 | 0.0014037 | -1.3573
2005-01-0068.add | 2005-01-0068.sam | 121 | 353 | 353 | 0 | 0 | 250 | 249.757 | 0.42768 | -0.24347 | 4.593 | 3.0405 | 0.003679 | -1.5525
2005-01-0068.add | 2005-01-0068.sam | 121 | 354 | 354 | 0 | 0 | 200 | 201.071 | 0.14409 | 1.0714 | 5.291 | 3.4124 | 0.0028862 | -1.8786
2005-01-0068.add | 2005-01-0068.sam | 121 | 355 | 355 | 0 | 0 | 175 | 175.289 | 0.3426 | 0.28914 | 5.823 | 3.7705 | 0.0031573 | -2.0525
2005-01-0068.add | 2005-01-0068.sam | 121 | 356 | 356 | 0 | 0 | 150 | 149.378 | 0.2337 | -0.62225 | 6.256 | 4.0585 | 0.0032993 | -2.1975
2005-01-0068.add | 2005-01-0068.sam | 121 | 357 | 357 | 0 | 0 | 125 | 125.288 | 0.36369 | 0.28801 | 6.346 | 4.1637 | 0.0017177 | -2.1823
2005-01-0068.add | 2005-01-0068.sam | 121 | 358 | 358 | 0 | 0 | 100 | 98.579 | 0.59052 | -1.421 | 6.716 | 4.4143 | 0.0018895 | -2.3017
2005-01-0068.add | 2005-01-0068.sam | 121 | 359 | 359 | 0 | 0 | 75 | 75.6397 | 0.31522 | 0.63966 | 6.582 | 4.3857 | 0.0012572 | -2.1963
2005-01-0068.add | 2005-01-0068.sam | 121 | 360 | 360 | 0 | 0 | 50 | 49.8857 | 0.14984 | -0.11427 | 6.61 | 4.3946 | 0.0015036 | -2.2154
2005-01-0068.add | 2005-01-0068.sam | 121 | 361 | 361 | 0 | 0 | 25 | 24.7805 | 0.18714 | -0.21947 | 6.629 | 4.4029 | 0.0022066 | -2.2261
2005-01-0068.add | 2005-01-0068.sam | 121 | 362 | 362 | 0 | 0 | 10 | 10.087 | 0.11165 | 0.086975 | 6.6 | 4.403 | 0.0018797 | -2.197
2005-01-0068.add | 2005-01-0068.sam | 121 | 363 | 363 | 0 | 0 | 5 | 4.5509 | 0.27574 | -0.44913 | 6.61 | 4.4072 | 0.001711 | -2.2028
2005-01-0073.add | 2005-01-0073.sam | 121 | 366 | 366 | 0 | 0 | 4297 | 4296.11 | 0.85002 | -0.89062 | 3.264 | 1.9875 | 0.0056304 | -1.2765
2005-01-0073.add | 2005-01-0073.sam | 121 | 367 | 367 | 0 | 0 | 4000 | 3998.81 | 0.82125 | -1.1899 | 3.125 | 1.977 | 0.0031318 | -1.148
2005-01-0073.add | 2005-01-0073.sam | 121 | 368 | 368 | 0 | 0 | 3500 | 3500.42 | 0.28291 | 0.41797 | 2.923 | 1.844 | 0 | -1.079
2005-01-0073.add | 2005-01-0073.sam | 121 | 369 | 369 | 0 | 0 | 3000 | 3001.6 | 0.44799 | 1.603 | 2.497 | 1.5765 | 0.0026016 | -0.92051
2005-01-0073.add | 2005-01-0073.sam | 121 | 370 | 370 | 0 | 0 | 2500 | 2501.5 | 0.84328 | 1.5024 | 1.899 | 1.2116 | 0.0013283 | -0.68736
2005-01-0073.add | 2005-01-0073.sam | 121 | 371 | 371 | 0 | 0 | 2000 | 2001.51 | 0.47385 | 1.5132 | 1.244 | 0.79838 | 0.0016546 | -0.44562
2005-01-0073.add | 2005-01-0073.sam | 121 | 372 | 372 | 0 | 0 | 1500 | 1500.79 | 0.67265 | 0.78565 | 0.489 | 0.32808 | 0.00064014 | -0.16092
2005-01-0073.add | 2005-01-0073.sam | 121 | 373 | 373 | 0 | 0 | 1250 | 1250.28 | 0.12335 | 0.27564 | 0.317 | 0.21626 | 0.0017867 | -0.10074
2005-01-0073.add | 2005-01-0073.sam | 121 | 374 | 374 | 0 | 0 | 1000 | 1000.49 | 0.12977 | 0.49457 | 0.314 | 0.21597 | 0.00036364 | -0.098033
2005-01-0073.add | 2005-01-0073.sam | 121 | 375 | 375 | 0 | 0 | 800 | 801.742 | 0.52064 | 1.7421 | 0.499 | 0.32793 | 0.00051211 | -0.17107
2005-01-0073.add | 2005-01-0073.sam | 121 | 376 | 376 | 0 | 0 | 600 | 600.251 | 0.23163 | 0.25092 | 0.962 | 0.61294 | 0.0017715 | -0.34906
2005-01-0073.add | 2005-01-0073.sam | 121 | 377 | 377 | 0 | 0 | 400 | 398.905 | 0.58172 | -1.0949 | 1.906 | 1.2254 | 0.0030304 | -0.68064
2005-01-0073.add | 2005-01-0073.sam | 121 | 378 | 378 | 0 | 0 | 300 | 301.608 | 0.37111 | 1.6081 | 2.907 | 1.8644 | 0.0019371 | -1.0426
2005-01-0073.add | 2005-01-0073.sam | 121 | 379 | 379 | 0 | 0 | 250 | 251.725 | 0.50634 | 1.7254 | 3.898 | 2.4838 | 0.0092709 | -1.4142
2005-01-0073.add | 2005-01-0073.sam | 121 | 380 | 380 | 0 | 0 | 200 | 200.658 | 0.83374 | 0.65791 | 4.581 | 3.0897 | 0.0023932 | -1.4913
2005-01-0073.add | 2005-01-0073.sam | 121 | 381 | 381 | 0 | 0 | 175 | 175.629 | 0.2906 | 0.62889 | 5.052 | 3.2801 | 0.0019439 | -1.7719
2005-01-0073.add | 2005-01-0073.sam | 121 | 382 | 382 | 0 | 0 | 150 | 148.758 | 0.44328 | -1.2416 | 5.332 | 3.4803 | 0.0025891 | -1.8517
2005-01-0073.add | 2005-01-0073.sam | 121 | 383 | 383 | 0 | 0 | 125 | 125.149 | 0.78861 | 0.14897 | 5.92 | 3.857 | 0.009024 | -2.063
2005-01-0073.add | 2005-01-0073.sam | 121 | 384 | 384 | 0 | 0 | 100 | 100.411 | 0.39308 | 0.41145 | 6.368 | 4.1701 | 0.0022345 | -2.1979
2005-01-0073.add | 2005-01-0073.sam | 121 | 385 | 385 | 0 | 0 | 75 | 77.0474 | 0.48443 | 2.0474 | 6.632 | 4.3789 | 0.002038 | -2.2531
2005-01-0073.add | 2005-01-0073.sam | 121 | 386 | 386 | 0 | 0 | 50 | 50.7156 | 0.28242 | 0.71559 | 6.702 | 4.4604 | 0.0012024 | -2.2416
2005-01-0073.add | 2005-01-0073.sam | 121 | 387 | 387 | 0 | 0 | 25 | 25.2391 | 0.17237 | 0.23913 | 6.746 | 4.462 | 0.0016602 | -2.284
2005-01-0073.add | 2005-01-0073.sam | 121 | 388 | 388 | 0 | 0 | 10 | 10.1999 | 0.17594 | 0.19987 | 6.725 | 4.464 | 0.0003873 | -2.261
2005-01-0073.add | 2005-01-0073.sam | 121 | 389 | 389 | 0 | 0 | 5 | 4.4886 | 0.24641 | -0.5114 | 6.726 | 4.4673 | 0.0012418 | -2.2587
2005-01-0077.add | 2005-01-0077.sam | 121 | 392 | 392 | 0 | 0 | 4201 | 4201.11 | 0.2505 | 0.10693 | 3.062 | 1.9142 | 0.0022484 | -1.1478
2005-01-0077.add | 2005-01-0077.sam | 121 | 393 | 393 | 0 | 0 | 4000 | 3998.58 | 0.26882 | -1.4228 | 3.008 | 1.8954 | 0.0052708 | -1.1126
2005-01-0077.add | 2005-01-0077.sam | 121 | 394 | 394 | 0 | 0 | 3500 | 3498.99 | 0.63928 | -1.0127 | 2.834 | 1.7689 | 0.0029451 | -1.0651
2005-01-0077.add | 2005-01-0077.sam | 121 | 395 | 395 | 0 | 0 | 3000 | 3000.35 | 0.12517 | 0.3457 | 2.407 | 1.5056 | 0.0026573 | -0.90141
2005-01-0077.add | 2005-01-0077.sam | 121 | 396 | 396 | 0 | 0 | 2500 | 2500.04 | 0.38573 | 0.035645 | 1.904 | 1.1867 | 0.0025 | -0.71727
2005-01-0077.add | 2005-01-0077.sam | 121 | 397 | 397 | 0 | 0 | 2000 | 2000.94 | 0.31785 | 0.93713 | 1.293 | 0.82421 | 0.00099931 | -0.46879
2005-01-0077.add | 2005-01-0077.sam | 121 | 398 | 398 | 0 | 0 | 1500 | 1499.71 | 0.35183 | -0.28955 | 0.506 | 0.34704 | 0.00045454 | -0.15896
2005-01-0077.add | 2005-01-0077.sam | 121 | 399 | 399 | 0 | 0 | 1250 | 1249.27 | 0.21818 | -0.73413 | 0.3 | 0.21146 | 0.0015548 | -0.088537
2005-01-0077.add | 2005-01-0077.sam | 121 | 400 | 400 | 0 | 0 | 1000 | 1000.34 | 0.21223 | 0.34375 | 0.242 | 0.17125 | 0.0019928 | -0.070752
2005-01-0077.add | 2005-01-0077.sam | 121 | 401 | 401 | 0 | 0 | 800 | 801.259 | 0.8392 | 1.2589 | 0.499 | 0.24225 | 0.0024334 | -0.25675
2005-01-0077.add | 2005-01-0077.sam | 121 | 402 | 402 | 0 | 0 | 600 | 598.737 | 0.83297 | -1.2629 | 0.852 | 0.55191 | 0.0017888 | -0.30009
2005-01-0077.add | 2005-01-0077.sam | 121 | 403 | 403 | 0 | 0 | 400 | 399.773 | 0.37799 | -0.22678 | 2.335 | 1.3935 | 0.0015659 | -0.94149
2005-01-0077.add | 2005-01-0077.sam | 121 | 404 | 404 | 0 | 0 | 300 | 301.302 | 0.28849 | 1.3024 | 3.397 | 2.1879 | 0.0017008 | -1.2091
2005-01-0077.add | 2005-01-0077.sam | 121 | 405 | 405 | 0 | 0 | 250 | 250.568 | 0.27467 | 0.56839 | 4.027 | 2.6173 | 0.0019649 | -1.4097
2005-01-0077.add | 2005-01-0077.sam | 121 | 406 | 406 | 0 | 0 | 200 | 199.803 | 0.4091 | -0.19714 | 4.617 | 3.102 | 0.001981 | -1.515
2005-01-0077.add | 2005-01-0077.sam | 121 | 407 | 407 | 0 | 0 | 175 | 175.51 | 0.3087 | 0.50995 | 4.884 | 3.1672 | 0.0012974 | -1.7168
2005-01-0077.add | 2005-01-0077.sam | 121 | 408 | 408 | 0 | 0 | 150 | 149.628 | 0.32269 | -0.37155 | 5.215 | 3.4144 | 0.0037168 | -1.8006
2005-01-0077.add | 2005-01-0077.sam | 121 | 409 | 409 | 0 | 0 | 125 | 125.132 | 0.26626 | 0.13247 | 5.811 | 3.7742 | 0.005372 | -2.0368
2005-01-0077.add | 2005-01-0077.sam | 121 | 410 | 410 | 0 | 0 | 100 | 99.9763 | 0.35596 | -0.023651 | 6.629 | 4.3573 | 0.0032074 | -2.2717
2005-01-0077.add | 2005-01-0077.sam | 121 | 411 | 411 | 0 | 0 | 75 | 74.4035 | 0.088755 | -0.59649 | 6.722 | 4.4359 | 0.00062452 | -2.2861
2005-01-0077.add | 2005-01-0077.sam | 121 | 412 | 412 | 0 | 0 | 50 | 50.5914 | 0.38225 | 0.59137 | 6.642 | 4.4173 | 0.0018245 | -2.2247
2005-01-0077.add | 2005-01-0077.sam | 121 | 413 | 413 | 0 | 0 | 25 | 25.2013 | 0.1047 | 0.20133 | 6.699 | 4.4224 | 0.0011819 | -2.2766
2005-01-0077.add | 2005-01-0077.sam | 121 | 414 | 414 | 0 | 0 | 10 | 9.7596 | 0.2665 | -0.24044 | 6.689 | 4.4264 | 0.0011142 | -2.2626
2005-01-0077.add | 2005-01-0077.sam | 121 | 415 | 415 | 0 | 0 | 5 | 5.6067 | 0.28208 | 0.60674 | 6.678 | 4.4256 | 0.001617 | -2.2524
2005-01-0082.add | 2005-01-0082.sam | 121 | 418 | 418 | 0 | 0 | 3342 | 3342.31 | 0.35376 | 0.30786 | 2.731 | 1.6874 | 0.0018858 | -1.0436
2005-01-0082.add | 2005-01-0082.sam | 121 | 419 | 419 | 0 | 0 | 3000 | 3001.04 | 0.49995 | 1.0378 | 2.299 | 1.4561 | 0.00054546 | -0.84295
2005-01-0082.add | 2005-01-0082.sam | 121 | 420 | 420 | 0 | 0 | 2500 | 2497.99 | 0.84452 | -2.0073 | 1.908 | 1.2162 | 0.0012845 | -0.69182
2005-01-0082.add | 2005-01-0082.sam | 121 | 421 | 421 | 0 | 0 | 2000 | 1995.74 | 0.36664 | -4.2631 | 1.351 | 0.87144 | 0.0031699 | -0.47956
2005-01-0082.add | 2005-01-0082.sam | 121 | 422 | 422 | 0 | 0 | 1500 | 1496.83 | 0.094953 | -3.1746 | 0.64 | 0.4072 | 0.00087197 | -0.2328
2005-01-0082.add | 2005-01-0082.sam | 121 | 423 | 423 | 0 | 0 | 1250 | 1247.79 | 0.12816 | -2.2123 | 0.324 | 0.22221 | 0.0026106 | -0.10179
2005-01-0082.add | 2005-01-0082.sam | 121 | 424 | 424 | 0 | 0 | 1000 | 996.482 | 0.15313 | -3.5178 | 0.214 | 0.15563 | 0.0017895 | -0.058372
2005-01-0082.add | 2005-01-0082.sam | 121 | 425 | 425 | 0 | 0 | 800 | 806.108 | 0.41673 | 6.1075 | 0.267 | 0.15604 | 0.0017146 | -0.11096
2005-01-0082.add | 2005-01-0082.sam | 121 | 426 | 426 | 0 | 0 | 600 | 604.872 | 0.31697 | 4.8721 | 0.529 | 0.31112 | 0.0017088 | -0.21788
2005-01-0082.add | 2005-01-0082.sam | 121 | 427 | 427 | 0 | 0 | 400 | 407.077 | 0.34465 | 7.0774 | 1.713 | 1.0677 | 0.0033176 | -0.64533
2005-01-0082.add | 2005-01-0082.sam | 121 | 428 | 428 | 0 | 0 | 300 | 303.923 | 0.13549 | 3.9229 | 3.016 | 1.9182 | 0.0032042 | -1.0978
2005-01-0082.add | 2005-01-0082.sam | 121 | 429 | 429 | 0 | 0 | 250 | 253.186 | 0.20356 | 3.1861 | 3.587 | 2.2928 | 0.00092633 | -1.2942
2005-01-0082.add | 2005-01-0082.sam | 121 | 430 | 430 | 0 | 0 | 200 | 201.27 | 0.51275 | 1.27 | 4.454 | 2.8803 | 0.002273 | -1.5737
2005-01-0082.add | 2005-01-0082.sam | 121 | 431 | 431 | 0 | 0 | 175 | 173.963 | 0.49482 | -1.0366 | 4.623 | 3.1247 | 0.0020064 | -1.4983
2005-01-0082.add | 2005-01-0082.sam | 121 | 432 | 432 | 0 | 0 | 150 | 151.536 | 0.14839 | 1.5359 | 4.464 | 2.9449 | 0.0018055 | -1.5191
2005-01-0082.add | 2005-01-0082.sam | 121 | 433 | 433 | 0 | 0 | 125 | 126.586 | 0.29554 | 1.5856 | 5.484 | 3.5627 | 0.002504 | -1.9213
2005-01-0082.add | 2005-01-0082.sam | 121 | 434 | 434 | 0 | 0 | 100 | 100.402 | 0.29562 | 0.40169 | 6.458 | 4.2226 | 0.0023223 | -2.2354
2005-01-0082.add | 2005-01-0082.sam | 121 | 435 | 435 | 0 | 0 | 75 | 74.0483 | 0.22137 | -0.95171 | 6.518 | 4.3244 | 0.0017313 | -2.1936
2005-01-0082.add | 2005-01-0082.sam | 121 | 436 | 436 | 0 | 0 | 50 | 49.93 | 0.39272 | -0.070042 | 6.609 | 4.3228 | 0.0020977 | -2.2862
2005-01-0082.add | 2005-01-0082.sam | 121 | 437 | 437 | 0 | 0 | 25 | 24.5799 | 0.27341 | -0.42008 | 7.002 | 4.3443 | 0.002411 | -2.6577
2005-01-0082.add | 2005-01-0082.sam | 121 | 438 | 438 | 0 | 0 | 10 | 11.1119 | 0.17292 | 1.1119 | 6.539 | 4.3595 | 0.0029918 | -2.1795
2005-01-0082.add | 2005-01-0082.sam | 121 | 439 | 439 | 0 | 0 | 5 | 3.6794 | 0.1964 | -1.3206 | 6.575 | 4.3625 | 0.0029554 | -2.2125
2005-01-0085.add | 2005-01-0085.sam | 121 | 441 | 441 | 0 | 0 | 2000 | 1999.55 | 0.22737 | -0.44653 | 2.13 | 0.84788 | 0.0020904 | -1.2821
2005-01-0085.add | 2005-01-0085.sam | 121 | 442 | 442 | 0 | 0 | 1500 | 1500.86 | 0.45987 | 0.85706 | 0.604 | 0.38846 | 0.0015548 | -0.21554
2005-01-0085.add | 2005-01-0085.sam | 121 | 443 | 443 | 0 | 0 | 1250 | 1251.43 | 0.82942 | 1.4276 | 0.356 | 0.22507 | 0.00072727 | -0.13093
2005-01-0085.add | 2005-01-0085.sam | 121 | 444 | 444 | 0 | 0 | 1000 | 1001.33 | 0.15632 | 1.3288 | 0.273 | 0.16955 | 0.0019577 | -0.10345
2005-01-0085.add | 2005-01-0085.sam | 121 | 445 | 445 | 0 | 0 | 800 | 803.855 | 0.11578 | 3.855 | 0.315 | 0.19381 | 0.0018363 | -0.12119
2005-01-0085.add | 2005-01-0085.sam | 121 | 446 | 446 | 0 | 0 | 600 | 601.676 | 0.59367 | 1.6758 | 0.548 | 0.34427 | 0.0015492 | -0.20373
2005-01-0085.add | 2005-01-0085.sam | 121 | 447 | 447 | 0 | 0 | 400 | 401.528 | 0.58964 | 1.5281 | 1.83 | 1.1684 | 0.0015963 | -0.66156
2005-01-0085.add | 2005-01-0085.sam | 121 | 448 | 448 | 0 | 0 | 300 | 299.696 | 0.31095 | -0.30435 | 2.453 | 1.5772 | 0.0039386 | -0.87577
2005-01-0085.add | 2005-01-0085.sam | 121 | 449 | 449 | 0 | 0 | 250 | 251.149 | 0.12937 | 1.1492 | 3.151 | 2.0046 | 0.002128 | -1.1463
2005-01-0085.add | 2005-01-0085.sam | 121 | 450 | 450 | 0 | 0 | 200 | 200.618 | 1.0845 | 0.61838 | 4.161 | 2.6787 | 0.0081011 | -1.4824
2005-01-0085.add | 2005-01-0085.sam | 121 | 451 | 451 | 0 | 0 | 175 | 175.661 | 0.14709 | 0.66069 | 4.682 | 3.0428 | 0.0013396 | -1.6392
2005-01-0085.add | 2005-01-0085.sam | 121 | 452 | 452 | 0 | 0 | 150 | 150.244 | 0.082279 | 0.24446 | 4.968 | 3.2226 | 0.0019595 | -1.7454
2005-01-0085.add | 2005-01-0085.sam | 121 | 453 | 453 | 0 | 0 | 125 | 125.428 | 0.21249 | 0.42841 | 5.175 | 3.4016 | 0.0017677 | -1.7734
2005-01-0085.add | 2005-01-0085.sam | 121 | 454 | 454 | 0 | 0 | 100 | 100.394 | 0.52829 | 0.39412 | 6.578 | 4.3486 | 0.0013157 | -2.2294
2005-01-0085.add | 2005-01-0085.sam | 121 | 455 | 455 | 0 | 0 | 75 | 74.9905 | 0.39506 | -0.0095062 | 6.647 | 4.4121 | 0.00081207 | -2.2349
2005-01-0085.add | 2005-01-0085.sam | 121 | 456 | 456 | 0 | 0 | 50 | 50.6495 | 0.22069 | 0.64949 | 6.697 | 4.428 | 0.0012676 | -2.269
2005-01-0085.add | 2005-01-0085.sam | 121 | 457 | 457 | 0 | 0 | 25 | 25.1855 | 0.23854 | 0.18554 | 6.682 | 4.4291 | 0.0013817 | -2.2529
2005-01-0085.add | 2005-01-0085.sam | 121 | 458 | 458 | 0 | 0 | 10 | 9.6049 | 0.19968 | -0.39507 | 6.667 | 4.4308 | 0.0022261 | -2.2362
2005-01-0085.add | 2005-01-0085.sam | 121 | 459 | 459 | 0 | 0 | 5 | 5.2854 | 0.077476 | 0.28541 | 6.796 | 4.4309 | 0.00048819 | -2.3651
2005-01-0090.add | 2005-01-0090.sam | 121 | 461 | 461 | 0 | 0 | 2572 | 2572.39 | 0.48374 | 0.3938 | 1.933 | 1.1895 | 0.0015007 | -0.74349
2005-01-0090.add | 2005-01-0090.sam | 121 | 462 | 462 | 0 | 0 | 2000 | 2000.27 | 0.69889 | 0.27039 | 1.423 | 0.90289 | 0.0024793 | -0.52011
2005-01-0090.add | 2005-01-0090.sam | 121 | 463 | 463 | 0 | 0 | 1500 | 1501.18 | 0.3593 | 1.1809 | 0.752 | 0.43963 | 0.0013174 | -0.31237
2005-01-0090.add | 2005-01-0090.sam | 121 | 464 | 464 | 0 | 0 | 1250 | 1251.23 | 0.5052 | 1.2319 | 0.454 | 0.2794 | 0.0012005 | -0.1746
2005-01-0090.add | 2005-01-0090.sam | 121 | 465 | 465 | 0 | 0 | 1000 | 998.646 | 0.23251 | -1.3539 | 0.243 | 0.16799 | 0.00058446 | -0.075008
2005-01-0090.add | 2005-01-0090.sam | 0 | 9465 | 0 | 0 | -9465 | 1000 | 0 | 0 | -1000 | 1.383 | 0 | 0 | -1.383
2005-01-0090.add | 2005-01-0090.sam | 121 | 466 | 466 | 0 | 0 | 800 | 802.636 | 0.22548 | 2.6359 | 0.277 | 0.16296 | 0.0017146 | -0.11404
2005-01-0090.add | 2005-01-0090.sam | 121 | 467 | 467 | 0 | 0 | 600 | 599.335 | 0.87627 | -0.6651 | 0.541 | 0.35181 | 0.0019075 | -0.18919
2005-01-0090.add | 2005-01-0090.sam | 121 | 468 | 468 | 0 | 0 | 400 | 399.909 | 0.12728 | -0.090973 | 1.552 | 0.98201 | 0.0017344 | -0.56999
2005-01-0090.add | 2005-01-0090.sam | 121 | 469 | 469 | 0 | 0 | 300 | 300.394 | 0.44982 | 0.39441 | 2.416 | 1.5411 | 0.0016861 | -0.87488
2005-01-0090.add | 2005-01-0090.sam | 121 | 470 | 470 | 0 | 0 | 250 | 249.659 | 0.06262 | -0.34068 | 2.957 | 1.8523 | 0.0019439 | -1.1046
2005-01-0090.add | 2005-01-0090.sam | 121 | 471 | 471 | 0 | 0 | 200 | 198.585 | 0.54379 | -1.4151 | 3.854 | 2.4955 | 0.0034907 | -1.3585
2005-01-0090.add | 2005-01-0090.sam | 121 | 472 | 472 | 0 | 0 | 175 | 174.145 | 0.18582 | -0.85545 | 4.205 | 2.749 | 0.001668 | -1.456
2005-01-0090.add | 2005-01-0090.sam | 121 | 473 | 473 | 0 | 0 | 150 | 151.187 | 0.86085 | 1.1866 | 4.518 | 2.8798 | 0.0018226 | -1.6382
2005-01-0090.add | 2005-01-0090.sam | 121 | 474 | 474 | 0 | 0 | 125 | 125.718 | 0.69051 | 0.718 | 4.776 | 3.1219 | 0.0013291 | -1.6541
2005-01-0090.add | 2005-01-0090.sam | 121 | 475 | 475 | 0 | 0 | 100 | 98.6873 | 0.55424 | -1.3127 | 5.58 | 3.5725 | 0.0039436 | -2.0075
2005-01-0090.add | 2005-01-0090.sam | 121 | 476 | 476 | 0 | 0 | 75 | 73.0154 | 0.16895 | -1.9846 | 6.595 | 4.3425 | 0.0017893 | -2.2525
2005-01-0090.add | 2005-01-0090.sam | 121 | 477 | 477 | 0 | 0 | 50 | 50.4418 | 0.18342 | 0.44178 | 6.586 | 4.3829 | 0.0013355 | -2.2031
2005-01-0090.add | 2005-01-0090.sam | 121 | 478 | 478 | 0 | 0 | 25 | 24.57 | 0.30696 | -0.43002 | 6.627 | 4.4097 | 0.0021378 | -2.2173
2005-01-0090.add | 2005-01-0090.sam | 121 | 479 | 479 | 0 | 0 | 10 | 11.6758 | 0.17642 | 1.6758 | 6.633 | 4.4159 | 0.0015077 | -2.2171
2005-01-0090.add | 2005-01-0090.sam | 121 | 480 | 480 | 0 | 0 | 5 | 6.0236 | 0.19825 | 1.0235 | 6.675 | 4.4178 | 0.0017766 | -2.2572
2005-01-0095.add | 2005-01-0095.sam | 121 | 483 | 483 | 0 | 0 | 2625 | 2633.92 | 0.2719 | 8.9204 | 1.843 | 1.1427 | 0.0025483 | -0.7003
2005-01-0095.add | 2005-01-0095.sam | 121 | 484 | 484 | 0 | 0 | 2500 | 2500.54 | 0.23963 | 0.53565 | 1.835 | 1.1621 | 0.00090911 | -0.67292
2005-01-0095.add | 2005-01-0095.sam | 121 | 485 | 485 | 0 | 0 | 2000 | 2002.19 | 0.21459 | 2.1924 | 1.4 | 0.8981 | 0.0024372 | -0.5019
2005-01-0095.add | 2005-01-0095.sam | 121 | 486 | 486 | 0 | 0 | 1500 | 1500.99 | 0.097577 | 0.98535 | 0.7 | 0.45113 | 0.0021831 | -0.24887
2005-01-0095.add | 2005-01-0095.sam | 121 | 487 | 487 | 0 | 0 | 1250 | 1249.63 | 0.21728 | -0.3717 | 0.398 | 0.26439 | 0.0015512 | -0.13361
2005-01-0095.add | 2005-01-0095.sam | 121 | 488 | 488 | 0 | 0 | 1000 | 999.412 | 0.2711 | -0.58826 | 0.314 | 0.1841 | 0.00062456 | -0.1299
2005-01-0095.add | 2005-01-0095.sam | 121 | 489 | 489 | 0 | 0 | 800 | 799.052 | 0.38068 | -0.94788 | 0.242 | 0.14844 | 0.0013958 | -0.093562
2005-01-0095.add | 2005-01-0095.sam | 121 | 490 | 490 | 0 | 0 | 600 | 598.112 | 0.26103 | -1.8876 | 0.379 | 0.24249 | 0.0019413 | -0.13651
2005-01-0095.add | 2005-01-0095.sam | 121 | 491 | 491 | 0 | 0 | 400 | 399.795 | 0.37405 | -0.20538 | 1.316 | 0.84912 | 0.0018054 | -0.46688
2005-01-0095.add | 2005-01-0095.sam | 121 | 492 | 492 | 0 | 0 | 300 | 297.419 | 0.17995 | -2.5813 | 2.115 | 1.3692 | 0.0009489 | -0.74576
2005-01-0095.add | 2005-01-0095.sam | 121 | 493 | 493 | 0 | 0 | 250 | 253.01 | 0.25307 | 3.0104 | 2.8 | 1.8068 | 0.0018557 | -0.9932
2005-01-0095.add | 2005-01-0095.sam | 121 | 494 | 494 | 0 | 0 | 200 | 204.763 | 0.060919 | 4.7634 | 3.504 | 2.2439 | 0.0033862 | -1.2601
2005-01-0095.add | 2005-01-0095.sam | 121 | 495 | 495 | 0 | 0 | 175 | 174.843 | 0.61067 | -0.15747 | 3.387 | 2.2014 | 0.0032068 | -1.1856
2005-01-0095.add | 2005-01-0095.sam | 121 | 496 | 496 | 0 | 0 | 150 | 151.252 | 0.54264 | 1.2519 | 3.39 | 2.2263 | 0.0069041 | -1.1637
2005-01-0095.add | 2005-01-0095.sam | 121 | 497 | 497 | 0 | 0 | 125 | 127.108 | 0.3692 | 2.1076 | 3.869 | 2.521 | 0.0025478 | -1.348
2005-01-0095.add | 2005-01-0095.sam | 121 | 498 | 498 | 0 | 0 | 100 | 99.6608 | 0.29939 | -0.33924 | 4.575 | 3.0116 | 0.0016894 | -1.5634
2005-01-0095.add | 2005-01-0095.sam | 121 | 499 | 499 | 0 | 0 | 75 | 76.8172 | 0.26404 | 1.8172 | 5.491 | 3.5302 | 0.0068678 | -1.9608
2005-01-0095.add | 2005-01-0095.sam | 121 | 500 | 500 | 0 | 0 | 50 | 49.8571 | 0.44803 | -0.14287 | 6.567 | 4.3769 | 0.0019192 | -2.1901
2005-01-0095.add | 2005-01-0095.sam | 121 | 501 | 501 | 0 | 0 | 25 | 24.4847 | 0.15707 | -0.51528 | 6.639 | 4.414 | 0.001449 | -2.225
2005-01-0095.add | 2005-01-0095.sam | 121 | 502 | 502 | 0 | 0 | 10 | 11.8118 | 0.35219 | 1.8117 | 6.612 | 4.4292 | 0.0009602 | -2.1828
2005-01-0095.add | 2005-01-0095.sam | 121 | 503 | 503 | 0 | 0 | 5 | 6.2732 | 0.15298 | 1.2732 | 6.618 | 4.4269 | 0.0010634 | -2.1911
2005-01-0100.add | 2005-01-0100.sam | 121 | 506 | 506 | 0 | 0 | 1756 | 1756.28 | 0.58251 | 0.28455 | 1.279 | 0.69494 | 0.00077779 | -0.58406
2005-01-0100.add | 2005-01-0100.sam | 121 | 507 | 507 | 0 | 0 | 1500 | 1501.19 | 0.21378 | 1.1948 | 0.868 | 0.44155 | 0.0014434 | -0.42645
2005-01-0100.add | 2005-01-0100.sam | 121 | 508 | 508 | 0 | 0 | 1250 | 1251 | 0.5765 | 0.99658 | 0.384 | 0.26006 | 0.0017715 | -0.12394
2005-01-0100.add | 2005-01-0100.sam | 121 | 509 | 509 | 0 | 0 | 1000 | 1001.76 | 0.37376 | 1.7635 | 0.218 | 0.15902 | 0.0018529 | -0.058984
2005-01-0100.add | 2005-01-0100.sam | 121 | 510 | 510 | 0 | 0 | 800 | 800.357 | 0.26156 | 0.35693 | 0.302 | 0.18711 | 0.0016723 | -0.11489
2005-01-0100.add | 2005-01-0100.sam | 121 | 511 | 511 | 0 | 0 | 600 | 599.476 | 0.47336 | -0.52405 | 0.805 | 0.50397 | 0.00036364 | -0.30103
2005-01-0100.add | 2005-01-0100.sam | 121 | 512 | 512 | 0 | 0 | 400 | 403.584 | 0.092177 | 3.5844 | 1.704 | 1.0882 | 0.0018506 | -0.61578
2005-01-0100.add | 2005-01-0100.sam | 121 | 513 | 513 | 0 | 0 | 300 | 300.032 | 0.15906 | 0.032105 | 1.898 | 1.2397 | 0.00099809 | -0.65826
2005-01-0100.add | 2005-01-0100.sam | 121 | 514 | 514 | 0 | 0 | 250 | 248.451 | 0.52123 | -1.5489 | 2.014 | 1.3259 | 0.00064942 | -0.68814
2005-01-0100.add | 2005-01-0100.sam | 121 | 515 | 515 | 0 | 0 | 200 | 199.203 | 0.38477 | -0.79654 | 2.29 | 1.4877 | 0.0026054 | -0.80225
2005-01-0100.add | 2005-01-0100.sam | 121 | 516 | 516 | 0 | 0 | 175 | 175.318 | 0.31151 | 0.3177 | 2.346 | 1.5429 | 0.0016723 | -0.80311
2005-01-0100.add | 2005-01-0100.sam | 121 | 517 | 517 | 0 | 0 | 150 | 151.213 | 0.067558 | 1.2126 | 2.455 | 1.6147 | 0.0017061 | -0.8403
2005-01-0100.add | 2005-01-0100.sam | 121 | 518 | 518 | 0 | 0 | 125 | 124.946 | 0.52959 | -0.054436 | 2.644 | 1.7423 | 0.0015614 | -0.90175
2005-01-0100.add | 2005-01-0100.sam | 121 | 519 | 519 | 0 | 0 | 100 | 100.14 | 0.53719 | 0.13982 | 3.005 | 1.9836 | 0.0018574 | -1.0214
2005-01-0100.add | 2005-01-0100.sam | 121 | 520 | 520 | 0 | 0 | 75 | 75.0758 | 0.22863 | 0.075821 | 3.565 | 2.3543 | 0.0017361 | -1.2107
2005-01-0100.add | 2005-01-0100.sam | 121 | 521 | 521 | 0 | 0 | 50 | 49.3962 | 0.14441 | -0.6038 | 4.301 | 2.8475 | 0.0021951 | -1.4535
2005-01-0100.add | 2005-01-0100.sam | 121 | 522 | 522 | 0 | 0 | 25 | 25.0857 | 0.10492 | 0.085743 | 6.661 | 4.4148 | 0.0016083 | -2.2462
2005-01-0100.add | 2005-01-0100.sam | 121 | 523 | 523 | 0 | 0 | 10 | 9.8932 | 0.25334 | -0.10683 | 6.815 | 4.4787 | 0.0017787 | -2.3363
2005-01-0100.add | 2005-01-0100.sam | 121 | 524 | 524 | 0 | 0 | 5 | 5.4559 | 0.1619 | 0.4559 | 6.825 | 4.5417 | 0.0016444 | -2.2833
Create Fit Sheet

## File List
Number of Files= |  | 18
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2005-01-0001.add | 2005-01-0001.sam | 12 | 19 | 2005-01-0001.add & 2005-01-0001.sam
2 | 2005-01-0003.add | 2005-01-0003.sam | 21 | 28 | 2005-01-0003.add & 2005-01-0003.sam
3 | 2005-01-0006.add | 2005-01-0006.sam | 30 | 47 | 2005-01-0006.add & 2005-01-0006.sam
4 | 2005-01-0015.add | 2005-01-0015.sam | 49 | 69 | 2005-01-0015.add & 2005-01-0015.sam
5 | 2005-01-0021.add | 2005-01-0021.sam | 71 | 93 | 2005-01-0021.add & 2005-01-0021.sam
6 | 2005-01-0029.add | 2005-01-0029.sam | 95 | 118 | 2005-01-0029.add & 2005-01-0029.sam
7 | 2005-01-0039.add | 2005-01-0039.sam | 120 | 143 | 2005-01-0039.add & 2005-01-0039.sam
8 | 2005-01-0057.add | 2005-01-0057.sam | 145 | 168 | 2005-01-0057.add & 2005-01-0057.sam
9 | 2005-01-0062.add | 2005-01-0062.sam | 170 | 188 | 2005-01-0062.add & 2005-01-0062.sam
10 | 2005-01-0065.add | 2005-01-0065.sam | 190 | 213 | 2005-01-0065.add & 2005-01-0065.sam
11 | 2005-01-0068.add | 2005-01-0068.sam | 215 | 238 | 2005-01-0068.add & 2005-01-0068.sam
12 | 2005-01-0073.add | 2005-01-0073.sam | 240 | 263 | 2005-01-0073.add & 2005-01-0073.sam
13 | 2005-01-0077.add | 2005-01-0077.sam | 265 | 288 | 2005-01-0077.add & 2005-01-0077.sam
14 | 2005-01-0082.add | 2005-01-0082.sam | 290 | 311 | 2005-01-0082.add & 2005-01-0082.sam
15 | 2005-01-0085.add | 2005-01-0085.sam | 313 | 331 | 2005-01-0085.add & 2005-01-0085.sam
16 | 2005-01-0090.add | 2005-01-0090.sam | 333 | 353 | 2005-01-0090.add & 2005-01-0090.sam
17 | 2005-01-0095.add | 2005-01-0095.sam | 355 | 375 | 2005-01-0095.add & 2005-01-0095.sam
18 | 2005-01-0100.add | 2005-01-0100.sam | 377 | 395 | 2005-01-0100.add & 2005-01-0100.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (CTD DOX - Bottle DOX) | CTD DOX | 18
2 | (CTD DOX - Bottle DOX) | Pressure | 18
3 | (CTD DOX - Bottle DOX) | File Pair # | 18
4 | (CTD DOX - Bottle DOX) | CTD DOX | 18
5 | (CTD DOX - Bottle DOX) | CTD DOX | 1
6 | (CTD DOX - Bottle DOX) | CTD DOX | 1
7 | (CTD DOX - Bottle DOX) | CTD DOX | 1
8 | (CTD DOX - Bottle DOX) | CTD DOX | 1

## Fit_1
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=CTD DOX
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.99358 |  | X | CALC Y
m= | -0.49311 | 0.0022965 |  | Std Dev of Resid |  | 0.064266 |  | 0.14844 | -0.12964
b= | -0.056447 | 0.0060204 |  | Resid Sum of Squares |  | 1.2308 |  | 4.6136 | -2.3315
# of Values Extracted = |  |  | 367
# of Values Used in Fit = |  |  | 300
File | Use | Sample # | CTD DOX | CTD DOX | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | CTD DOX
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 4 | 0.048 | 0 | 0.004 | -0.080116 | -0.084116 |  | 0.3693 | -0.1977
1 | NO | 5 | 0.04676 | 0.0014834 | -0.05324 | -0.079505 | -0.026265 |  | 0.21 | -0.099
1 | NO | 6 | 0.50074 | 0.0087174 | -0.0012562 | -0.30337 | -0.30211 |  | 0.20993 | -0.12207
1 | NO | 7 | 2.8522 | 0.0043797 | -1.0448 | -1.4629 | -0.41812 |  | 0.16153 | -0.17247
1 | NO | 8 | 3.2522 | 0.002912 | -1.1088 | -1.6602 | -0.55139 |  | 0.1613 | -0.2097
1 | NO | 9 | 3.687 | 0.0021289 | -1.264 | -1.8745 | -0.61051 |  | 0.4346 | -0.2234
1 | NO | 10 | 3.877 | 0.013065 | -1.411 | -1.9682 | -0.55721 |  | 0.92265 | -0.40936
1 | NO | 11 | 4.2193 | 0.0032873 | -1.3767 | -2.1371 | -0.76038 |  | 1.3029 | -0.70209
2 | NO | 13 | 3.4827 | 0.0042317 | -1.3573 | -1.7738 | -0.41653 |  | 1.3029 | -0.64814
2 | NO | 14 | 4.2358 | 0.0015184 | -1.4682 | -2.1452 | -0.67702 |  | 0.95083 | -0.53817
2 | NO | 15 | 4.3923 | 0.0015558 | -1.4437 | -2.2224 | -0.77868 |  | 0.478 | -0.218
2 | NO | 16 | 4.3895 | 0.0016078 | -1.4255 | -2.221 | -0.79546 |  | 0.2704 | -0.1416
2 | NO | 17 | 4.3972 | 0.0035285 | -1.4218 | -2.2248 | -0.80302 |  | 0.18318 | -0.11782
2 | NO | 18 | 4.5156 | 0.0024659 | -1.4634 | -2.2831 | -0.8197 |  | 0.17374 | -0.10126
2 | NO | 19 | 4.6644 | 0.012831 | -1.4856 | -2.3565 | -0.87089 |  | 0.33051 | -0.24649
2 | NO | 20 | 4.6928 | 0.0017865 | -1.5192 | -2.3705 | -0.8513 |  | 0.33226 | -0.26674
3 | YES | 22 | 0.3693 | 0.0017447 | -0.1977 | -0.23855 | -0.04085 |  | 0.99696 | -0.54804
3 | YES | 23 | 0.21 | 0 | -0.099 | -0.16 | -0.061001 |  | 1.6957 | -0.86334
3 | YES | 24 | 0.20993 | 0.00057973 | -0.12207 | -0.15996 | -0.03789 |  | 2.0467 | -1.0023
3 | YES | 25 | 0.16153 | 0.0013606 | -0.17247 | -0.1361 | 0.036372 |  | 2.5483 | -1.2267
3 | YES | 26 | 0.1613 | 0.0010539 | -0.2097 | -0.13598 | 0.073717 |  | 2.8469 | -1.3421
3 | YES | 27 | 0.4346 | 0.0012006 | -0.2234 | -0.27076 | -0.047358 |  | 3.0726 | -1.4324
3 | YES | 28 | 0.92265 | 0.0023447 | -0.40936 | -0.51141 | -0.10206 |  | 3.1975 | -1.4465
3 | YES | 29 | 1.3029 | 0.0011328 | -0.70209 | -0.69893 | 0.0031628 |  | 4.4961 | -2.0579
3 | NO | 31 | 1.9264 | 0.0010691 | -0.7866 | -1.0064 | -0.21978 |  | 1.5379 | -0.95205
3 | NO | 32 | 1.8416 | 0.0012006 | -0.7414 | -0.96456 | -0.22317 |  | 1.5143 | -0.8287
3 | NO | 33 | 2.0064 | 0.0013354 | -0.80164 | -1.0458 | -0.24417 |  | 1.2764 | -0.7406
3 | NO | 34 | 2.3332 | 0.0016834 | -0.94776 | -1.207 | -0.25924 |  | 0.89347 | -0.50653
3 | NO | 35 | 3.2813 | 0.0027519 | -1.3317 | -1.6745 | -0.34284 |  | 0.43764 | -0.31036
3 | NO | 36 | 4.0995 | 0.0030114 | -1.5895 | -2.0779 | -0.48842 |  | 0.25452 | -0.18648
3 | NO | 37 | 4.4806 | 0.0012583 | -1.7234 | -2.2659 | -0.54255 |  | 0.1806 | -0.1434
3 | NO | 38 | 4.4338 | 0.0029481 | -1.7332 | -2.2428 | -0.50958 |  | 0.18014 | -0.13086
3 | NO | 39 | 4.5799 | 0.0018105 | -1.7302 | -2.3148 | -0.58468 |  | 0.28279 | -0.19421
3 | NO | 40 | 4.5822 | 0.001127 | -1.7118 | -2.316 | -0.6042 |  | 0.74521 | -0.41079
4 | YES | 58 | 1.3029 | 0.0022558 | -0.64814 | -0.6989 | -0.050764 |  | 1.0969 | -0.58407
4 | YES | 59 | 0.95083 | 0.0012803 | -0.53817 | -0.52532 | 0.012849 |  | 1.3515 | -0.78649
4 | YES | 60 | 0.478 | 0 | -0.218 | -0.29215 | -0.074155 |  | 1.8271 | -0.95894
4 | YES | 61 | 0.2704 | 0.0012005 | -0.1416 | -0.18978 | -0.04818 |  | 1.9475 | -0.87647
4 | YES | 62 | 0.18318 | 0.0023805 | -0.11782 | -0.14678 | -0.028958 |  | 2.4851 | -1.2899
4 | YES | 63 | 0.17374 | 0.00099807 | -0.10126 | -0.14212 | -0.040854 |  | 2.9115 | -1.4495
4 | YES | 64 | 0.33051 | 0.001674 | -0.24649 | -0.21942 | 0.027067 |  | 3.944 | -1.946
4 | YES | 65 | 0.33226 | 0.0014866 | -0.26674 | -0.22029 | 0.046455 |  | 4.4955 | -2.1335
4 | YES | 66 | 0.99696 | 0.0016852 | -0.54804 | -0.54806 | -1.8925e-05 |  | 4.5092 | -2.1058
4 | YES | 67 | 1.6957 | 0.002548 | -0.86334 | -0.8926 | -0.029259 |  | 4.5223 | -2.0937
4 | YES | 68 | 2.0467 | 0.0023078 | -1.0023 | -1.0657 | -0.063341 |  | 4.524 | -2.1109
4 | YES | 69 | 2.5483 | 0.0018423 | -1.2267 | -1.313 | -0.086346 |  | 1.7655 | -0.95145
4 | YES | 70 | 2.8469 | 0.0011569 | -1.3421 | -1.4603 | -0.11813 |  | 1.5257 | -0.7953
4 | YES | 71 | 3.0726 | 0.0018841 | -1.4324 | -1.5716 | -0.13924 |  | 1.2279 | -0.6011
4 | NO | 72 | 3.1943 | 0.0016615 | -3.2447 | -1.6316 | 1.6131 |  | 0.8526 | -0.4314
4 | YES | 73 | 3.1975 | 0.0013041 | -1.4465 | -1.6332 | -0.18673 |  | 0.36298 | -0.17702
4 | YES | 74 | 4.4961 | 0.0029635 | -2.0579 | -2.2735 | -0.21568 |  | 0.2222 | -0.1298
4 | NO | 75 | 4.5756 | 0.0016338 | -2.6084 | -2.3127 | 0.29562 |  | 0.22228 | -0.090719
4 | NO | 76 | 4.5983 | 0.0018676 | -2.0307 | -2.3239 | -0.29315 |  | 0.186 | -0.07
4 | NO | 77 | 4.5999 | 0.0019402 | -2.0381 | -2.3247 | -0.28669 |  | 0.22246 | -0.080546
4 | NO | 78 | 4.6054 | 0.0016601 | -1.9826 | -2.3274 | -0.34481 |  | 0.47631 | -0.23669
5 | YES | 82 | 1.5379 | 0.0019487 | -0.95205 | -0.81483 | 0.13722 |  | 1.1175 | -0.57549
5 | YES | 83 | 1.5143 | 0.0013079 | -0.8287 | -0.80317 | 0.025535 |  | 1.569 | -0.84797
5 | YES | 84 | 1.2764 | 0.0014111 | -0.7406 | -0.68586 | 0.054747 |  | 1.9172 | -0.97879
5 | YES | 85 | 0.89347 | 0.0028754 | -0.50653 | -0.49703 | 0.0095003 |  | 2.5827 | -1.3153
5 | YES | 86 | 0.43764 | 0.0024118 | -0.31036 | -0.27225 | 0.038113 |  | 3.1111 | -1.4349
5 | YES | 87 | 0.25452 | 0.0027931 | -0.18648 | -0.18195 | 0.0045246 |  | 3.4436 | -1.7284
5 | YES | 88 | 0.1806 | 0.0012006 | -0.1434 | -0.1455 | -0.0021075 |  | 3.6641 | -1.8489
5 | YES | 89 | 0.18014 | 0.0016498 | -0.13086 | -0.14528 | -0.014417 |  | 4.197 | -2.108
5 | NO | 90 | 0.18255 | 0.0014434 | -0.44645 | -0.14646 | 0.29999 |  | 4.5605 | -2.2185
5 | YES | 91 | 0.28279 | 0.0016017 | -0.19421 | -0.1959 | -0.0016887 |  | 4.5715 | -2.2005
5 | YES | 92 | 0.74521 | 0.0019745 | -0.41079 | -0.42392 | -0.013125 |  | 4.5812 | -2.1808
5 | YES | 93 | 1.0969 | 0.0017547 | -0.58407 | -0.59736 | -0.013291 |  | 4.5865 | -2.1774
5 | YES | 94 | 1.3515 | 0.0014383 | -0.78649 | -0.72289 | 0.063595 |  | 4.5886 | -2.1664
5 | YES | 95 | 1.8271 | 0.0020829 | -0.95894 | -0.95739 | 0.0015488 |  | 1.9311 | -1.1079
5 | YES | 96 | 1.9475 | 0.0013421 | -0.87647 | -1.0168 | -0.14033 |  | 1.8347 | -1.0223
5 | YES | 97 | 2.4851 | 0.0016243 | -1.2899 | -1.2819 | 0.0080787 |  | 1.5906 | -0.86338
5 | YES | 98 | 2.9115 | 0.014906 | -1.4495 | -1.4921 | -0.042629 |  | 1.5897 | -0.85335
5 | YES | 99 | 3.944 | 0.0056295 | -1.946 | -2.0013 | -0.055313 |  | 1.2279 | -0.68212
5 | YES | 100 | 4.4955 | 0.0013044 | -2.1335 | -2.2732 | -0.1397 |  | 0.79214 | -0.42386
5 | YES | 101 | 4.5092 | 0.0018394 | -2.1058 | -2.28 | -0.17416 |  | 0.34511 | -0.16389
5 | NO | 102 | 4.5181 | 0.0014259 | -2.0639 | -2.2844 | -0.22046 |  | 0.22352 | -0.084479
5 | YES | 103 | 4.5223 | 0.0015575 | -2.0937 | -2.2865 | -0.19281 |  | 0.1986 | -0.068397
5 | YES | 104 | 4.524 | 0.0012706 | -2.1109 | -2.2873 | -0.17636 |  | 0.22269 | -0.10731
6 | NO | 122 | 1.8111 | 0.00089619 | -1.0939 | -0.94953 | 0.14435 |  | 0.32331 | -0.14869
6 | YES | 123 | 1.7655 | 0.0015809 | -0.95145 | -0.92706 | 0.024384 |  | 0.80869 | -0.49131
6 | YES | 124 | 1.5257 | 0.0012493 | -0.7953 | -0.80879 | -0.013491 |  | 1.2606 | -0.67841
6 | YES | 125 | 1.2279 | 0.00076817 | -0.6011 | -0.66194 | -0.060841 |  | 1.6433 | -0.91867
6 | YES | 126 | 0.8526 | 0.0026692 | -0.4314 | -0.47688 | -0.045479 |  | 2.3511 | -1.2879
6 | YES | 127 | 0.36298 | 0.0024562 | -0.17702 | -0.23544 | -0.058421 |  | 2.604 | -1.211
6 | YES | 128 | 0.2222 | 0.0021433 | -0.1298 | -0.16602 | -0.036213 |  | 2.871 | -1.521
6 | YES | 129 | 0.22228 | 0.0021917 | -0.090719 | -0.16606 | -0.075337 |  | 3.3187 | -1.7093
6 | YES | 130 | 0.186 | 0 | -0.07 | -0.14817 | -0.078166 |  | 4.4733 | -2.2897
6 | YES | 131 | 0.22246 | 0.0019322 | -0.080546 | -0.16614 | -0.085597 |  | 4.5131 | -2.2709
6 | YES | 132 | 0.47631 | 0.0014436 | -0.23669 | -0.29132 | -0.054637 |  | 4.5501 | -2.2669
6 | YES | 133 | 1.1175 | 0.0017084 | -0.57549 | -0.60751 | -0.032018 |  | 4.5613 | -2.2247
6 | YES | 134 | 1.569 | 0.0039453 | -0.84797 | -0.83016 | 0.017812 |  | 4.5683 | -2.2487
6 | YES | 135 | 1.9172 | 0.0018221 | -0.97879 | -1.0018 | -0.023062 |  | 4.5684 | -2.2656
6 | YES | 136 | 2.5827 | 0.0021456 | -1.3153 | -1.33 | -0.014679 |  | 2.0264 | -1.1266
6 | YES | 137 | 3.1111 | 0.0014921 | -1.4349 | -1.5906 | -0.15564 |  | 1.8585 | -1.0315
6 | YES | 138 | 3.4436 | 0.0019536 | -1.7284 | -1.7545 | -0.026189 |  | 1.6089 | -0.89612
6 | YES | 139 | 3.6641 | 0.001536 | -1.8489 | -1.8633 | -0.014391 |  | 1.2276 | -0.67139
6 | YES | 140 | 4.197 | 0.0031303 | -2.108 | -2.126 | -0.017996 |  | 0.8383 | -0.4547
6 | YES | 141 | 4.5605 | 0.0015763 | -2.2185 | -2.3053 | -0.086757 |  | 0.42333 | -0.19967
6 | YES | 142 | 4.5715 | 0.0010493 | -2.2005 | -2.3107 | -0.11017 |  | 0.2852 | -0.1488
6 | YES | 143 | 4.5812 | 0.0013225 | -2.1808 | -2.3155 | -0.13471 |  | 0.24974 | -0.12826
6 | YES | 144 | 4.5865 | 0.0018528 | -2.1774 | -2.3181 | -0.14068 |  | 0.25752 | -0.11748
6 | YES | 145 | 4.5886 | 0.0013785 | -2.1664 | -2.3192 | -0.1528 |  | 0.30445 | -0.15055
7 | YES | 186 | 1.9311 | 0.0012286 | -1.1079 | -1.0087 | 0.099173 |  | 0.57207 | -0.35693
7 | YES | 187 | 1.8347 | 0.002743 | -1.0223 | -0.96117 | 0.061124 |  | 0.87831 | -0.52769
7 | YES | 188 | 1.5906 | 0.0026761 | -0.86338 | -0.8408 | 0.022578 |  | 1.2268 | -0.72222
7 | YES | 189 | 1.5897 | 0.002941 | -0.85335 | -0.84032 | 0.013027 |  | 1.7213 | -0.84868
7 | YES | 190 | 1.2279 | 0.00078071 | -0.68212 | -0.66193 | 0.020194 |  | 2.0763 | -1.2137
7 | YES | 191 | 0.79214 | 0.0074926 | -0.42386 | -0.44706 | -0.023202 |  | 2.5045 | -1.3835
7 | YES | 192 | 0.34511 | 0.0023587 | -0.16389 | -0.22662 | -0.06273 |  | 3.4371 | -1.8549
7 | YES | 193 | 0.22352 | 0.0025103 | -0.084479 | -0.16667 | -0.082189 |  | 4.593 | -2.319
7 | YES | 194 | 0.1986 | 0.0012006 | -0.068397 | -0.15438 | -0.085984 |  | 4.6045 | -2.3495
7 | YES | 195 | 0.22269 | 0.0019834 | -0.10731 | -0.16626 | -0.058942 |  | 4.609 | -2.357
7 | YES | 196 | 0.32331 | 0.0021089 | -0.14869 | -0.21587 | -0.067179 |  | 4.6136 | -2.3403
7 | YES | 197 | 0.80869 | 0.013204 | -0.49131 | -0.45522 | 0.036082 |  | 0.78446 | -0.48255
7 | YES | 198 | 1.2606 | 0.0018827 | -0.67841 | -0.67806 | 0.0003533 |  | 0.34697 | -0.18203
7 | YES | 199 | 1.6433 | 0.0019382 | -0.91867 | -0.86679 | 0.051875 |  | 0.26829 | -0.16171
7 | YES | 200 | 2.3511 | 0.0008104 | -1.2879 | -1.2158 | 0.072096 |  | 0.25443 | -0.13557
7 | YES | 201 | 2.604 | 0.0016973 | -1.211 | -1.3405 | -0.12944 |  | 0.32993 | -0.22807
7 | YES | 202 | 2.871 | 0.0032544 | -1.521 | -1.4722 | 0.048797 |  | 0.45436 | -0.31764
7 | YES | 203 | 3.3187 | 0.0023088 | -1.7093 | -1.6929 | 0.016375 |  | 0.74686 | -0.44714
7 | YES | 204 | 4.4733 | 0.0024342 | -2.2897 | -2.2623 | 0.027472 |  | 1.2119 | -0.73312
7 | YES | 205 | 4.5131 | 0.001646 | -2.2709 | -2.2819 | -0.010984 |  | 1.807 | -1.002
7 | YES | 206 | 4.5501 | 0.0016361 | -2.2669 | -2.3001 | -0.033229 |  | 2.7778 | -1.5052
7 | YES | 207 | 4.5613 | 0.001885 | -2.2247 | -2.3057 | -0.081011 |  | 4.3459 | -2.2831
7 | YES | 208 | 4.5683 | 0.0016469 | -2.2487 | -2.3092 | -0.060508 |  | 4.4501 | -2.2789
7 | YES | 209 | 4.5684 | 0.0017776 | -2.2656 | -2.3092 | -0.043583 |  | 4.6012 | -2.2848
8 | NO | 258 | 2.0893 | 0.0030343 | -1.2647 | -1.0867 | 0.17801 |  | 4.6052 | -2.2878
8 | YES | 259 | 2.0264 | 0.0015592 | -1.1266 | -1.0557 | 0.070969 |  | 4.6101 | -2.2749
8 | YES | 260 | 1.8585 | 0.0029751 | -1.0315 | -0.97291 | 0.058559 |  | 4.6127 | -2.2713
8 | YES | 261 | 1.6089 | 0.0018999 | -0.89612 | -0.84981 | 0.046318 |  | 2.0817 | -1.1033
8 | YES | 262 | 1.2276 | 0.0013806 | -0.67139 | -0.6618 | 0.0095912 |  | 1.8758 | -1.1142
8 | YES | 263 | 0.8383 | 0.0023897 | -0.4547 | -0.46982 | -0.01512 |  | 1.2051 | -0.71593
8 | YES | 264 | 0.42333 | 0.001106 | -0.19967 | -0.2652 | -0.065528 |  | 0.7649 | -0.4561
8 | YES | 265 | 0.2852 | 0.00087197 | -0.1488 | -0.19708 | -0.04828 |  | 0.329 | -0.182
8 | YES | 266 | 0.24974 | 0.0018643 | -0.12826 | -0.1796 | -0.051343 |  | 0.23434 | -0.16666
8 | YES | 267 | 0.25752 | 0.0021451 | -0.11748 | -0.18343 | -0.065955 |  | 0.22027 | -0.18773
8 | YES | 268 | 0.30445 | 0.0013724 | -0.15055 | -0.20658 | -0.056032 |  | 0.29217 | -0.21784
8 | YES | 269 | 0.57207 | 0.0015692 | -0.35693 | -0.33854 | 0.018394 |  | 0.45698 | -0.27803
8 | YES | 270 | 0.87831 | 0.002141 | -0.52769 | -0.48955 | 0.038131 |  | 0.8295 | -0.4945
8 | YES | 271 | 1.2268 | 0.00079044 | -0.72222 | -0.66139 | 0.060835 |  | 1.2576 | -0.75237
8 | YES | 272 | 1.7213 | 0.005514 | -0.84868 | -0.90525 | -0.056573 |  | 1.7763 | -1.0257
8 | YES | 273 | 2.0763 | 0.0055284 | -1.2137 | -1.0803 | 0.13346 |  | 3.0289 | -1.5291
8 | YES | 274 | 2.5045 | 0.0017083 | -1.3835 | -1.2915 | 0.092023 |  | 4.3362 | -2.2598
8 | YES | 275 | 3.4371 | 0.0025153 | -1.8549 | -1.7513 | 0.10361 |  | 4.4581 | -2.3028
8 | NO | 276 | 4.5215 | 0.0025791 | -2.3795 | -2.286 | 0.093474 |  | 4.5042 | -2.2638
8 | NO | 277 | 4.5762 | 0.0024665 | -2.3988 | -2.313 | 0.085831 |  | 4.5157 | -2.2843
8 | YES | 278 | 4.593 | 0.0017391 | -2.319 | -2.3213 | -0.0022733 |  | 4.5275 | -2.2945
8 | YES | 279 | 4.6045 | 0.0015811 | -2.3495 | -2.327 | 0.022441 |  | 4.5323 | -2.2897
8 | YES | 280 | 4.609 | 0.0016146 | -2.357 | -2.3292 | 0.027737 |  | 4.5289 | -2.2981
8 | YES | 281 | 4.6136 | 0.0015638 | -2.3403 | -2.3315 | 0.0088541 |  | 1.8446 | -1.0723
9 | YES | 296 | 0.78446 | 0.0014434 | -0.48255 | -0.44327 | 0.039273 |  | 1.5831 | -0.91895
9 | YES | 297 | 0.34697 | 0.0020328 | -0.18203 | -0.22754 | -0.045508 |  | 1.1926 | -0.67337
9 | YES | 298 | 0.26829 | 0.0018772 | -0.16171 | -0.18874 | -0.027033 |  | 0.78359 | -0.45441
9 | YES | 299 | 0.25443 | 0.0023303 | -0.13557 | -0.18191 | -0.04634 |  | 0.30934 | -0.15266
9 | YES | 300 | 0.32993 | 0.00072727 | -0.22807 | -0.21914 | 0.0089245 |  | 0.19877 | -0.10423
9 | YES | 301 | 0.45436 | 0.0007746 | -0.31764 | -0.2805 | 0.037137 |  | 0.21537 | -0.096628
9 | YES | 302 | 0.74686 | 0.0017287 | -0.44714 | -0.42473 | 0.022409 |  | 0.34909 | -0.19291
9 | YES | 303 | 1.2119 | 0.001771 | -0.73312 | -0.65404 | 0.079076 |  | 0.62656 | -0.39444
9 | YES | 304 | 1.807 | 0.003248 | -1.002 | -0.94749 | 0.054528 |  | 1.3231 | -0.78195
9 | YES | 305 | 2.7778 | 0.0023032 | -1.5052 | -1.4262 | 0.07906 |  | 2.4287 | -1.3573
9 | NO | 306 | 3.1532 | 0.0023219 | -1.7338 | -1.6113 | 0.12244 |  | 3.0405 | -1.5525
9 | NO | 307 | 3.6766 | 0.0046223 | -1.9894 | -1.8694 | 0.12005 |  | 4.1637 | -2.1823
9 | YES | 308 | 4.3459 | 0.0022517 | -2.2831 | -2.1995 | 0.083665 |  | 4.4143 | -2.3017
9 | YES | 309 | 4.4501 | 0.013204 | -2.2789 | -2.2508 | 0.028068 |  | 4.3857 | -2.1963
9 | NO | 310 | 4.5979 | 0.0017311 | -2.0531 | -2.3237 | -0.27061 |  | 4.3946 | -2.2154
9 | YES | 311 | 4.6012 | 0.0015276 | -2.2848 | -2.3253 | -0.040527 |  | 4.4029 | -2.2261
9 | YES | 312 | 4.6052 | 0.0019398 | -2.2878 | -2.3273 | -0.039574 |  | 4.403 | -2.197
9 | YES | 313 | 4.6101 | 0.001881 | -2.2749 | -2.3297 | -0.054786 |  | 4.4072 | -2.2028
9 | YES | 314 | 4.6127 | 0.001021 | -2.2713 | -2.331 | -0.059668 |  | 1.977 | -1.148
10 | YES | 315 | 2.0817 | 0.0024677 | -1.1033 | -1.0829 | 0.020385 |  | 1.844 | -1.079
10 | NO | 316 | 2.0068 | 0.0014182 | -1.2132 | -1.046 | 0.16713 |  | 1.5765 | -0.92051
10 | YES | 317 | 1.8758 | 0.0020761 | -1.1142 | -0.98143 | 0.13277 |  | 1.2116 | -0.68736
10 | NO | 318 | 1.5924 | 0.0025964 | -1.0156 | -0.84168 | 0.17392 |  | 0.79838 | -0.44562
10 | YES | 319 | 1.2051 | 0.002893 | -0.71593 | -0.65068 | 0.065244 |  | 0.32808 | -0.16092
10 | YES | 320 | 0.7649 | 0.0026344 | -0.4561 | -0.43363 | 0.02247 |  | 0.21626 | -0.10074
10 | YES | 321 | 0.329 | 0 | -0.182 | -0.21868 | -0.036681 |  | 0.21597 | -0.098033
10 | YES | 322 | 0.23434 | 0.0014919 | -0.16666 | -0.172 | -0.0053414 |  | 0.32793 | -0.17107
10 | YES | 323 | 0.22027 | 0.0015492 | -0.18773 | -0.16507 | 0.022661 |  | 0.61294 | -0.34906
10 | YES | 324 | 0.29217 | 0.00079945 | -0.21784 | -0.20052 | 0.017318 |  | 1.2254 | -0.68064
10 | YES | 325 | 0.45698 | 0.0017534 | -0.27803 | -0.28179 | -0.003762 |  | 1.8644 | -1.0426
10 | YES | 326 | 0.8295 | 0.00059335 | -0.4945 | -0.46549 | 0.02901 |  | 3.0897 | -1.4913
10 | YES | 327 | 1.2576 | 0.0043092 | -0.75237 | -0.6766 | 0.075772 |  | 3.2801 | -1.7719
10 | YES | 328 | 1.7763 | 0.0030435 | -1.0257 | -0.93234 | 0.093397 |  | 3.4803 | -1.8517
10 | YES | 329 | 3.0289 | 0.0020587 | -1.5291 | -1.5501 | -0.020995 |  | 4.1701 | -2.1979
10 | NO | 330 | 3.6728 | 0.0047478 | -2.0452 | -1.8675 | 0.17768 |  | 4.3789 | -2.2531
10 | NO | 331 | 4.1619 | 0.0033134 | -2.2591 | -2.1087 | 0.15032 |  | 4.4604 | -2.2416
10 | YES | 332 | 4.3362 | 0.0065157 | -2.2598 | -2.1947 | 0.065098 |  | 4.462 | -2.284
10 | YES | 333 | 4.4581 | 0.00087248 | -2.3028 | -2.2548 | 0.048033 |  | 4.464 | -2.261
10 | YES | 334 | 4.5042 | 0.0016126 | -2.2638 | -2.2775 | -0.013695 |  | 4.4673 | -2.2587
10 | YES | 335 | 4.5157 | 0.0013758 | -2.2843 | -2.2832 | 0.0011642 |  | 1.8954 | -1.1126
10 | YES | 336 | 4.5275 | 0.0016689 | -2.2945 | -2.289 | 0.0055156 |  | 1.7689 | -1.0651
10 | YES | 337 | 4.5323 | 0.0015943 | -2.2897 | -2.2914 | -0.0016215 |  | 1.5056 | -0.90141
10 | YES | 338 | 4.5289 | 0.0016854 | -2.2981 | -2.2897 | 0.0083954 |  | 1.1867 | -0.71727
11 | NO | 340 | 2.0236 | 0.0043433 | -1.1954 | -1.0543 | 0.14103 |  | 0.82421 | -0.46879
11 | NO | 0 | 0 | 0 | -3.175 | -0.056447 | 3.1186 |  | 0.34704 | -0.15896
11 | YES | 342 | 1.8446 | 0.001982 | -1.0723 | -0.96607 | 0.10628 |  | 0.21146 | -0.088537
11 | YES | 343 | 1.5831 | 0.002152 | -0.91895 | -0.83707 | 0.081882 |  | 0.17125 | -0.070752
11 | YES | 344 | 1.1926 | 0.0022141 | -0.67337 | -0.64455 | 0.028824 |  | 0.24225 | -0.25675
11 | YES | 345 | 0.78359 | 0.0022718 | -0.45441 | -0.44284 | 0.01157 |  | 0.55191 | -0.30009
11 | YES | 346 | 0.30934 | 0.0017008 | -0.15266 | -0.20899 | -0.056325 |  | 2.1879 | -1.2091
11 | YES | 347 | 0.19877 | 0.00093773 | -0.10423 | -0.15446 | -0.050231 |  | 2.6173 | -1.4097
11 | YES | 348 | 0.21537 | 0.0014613 | -0.096628 | -0.16265 | -0.066021 |  | 3.102 | -1.515
11 | YES | 349 | 0.34909 | 0.00249 | -0.19291 | -0.22859 | -0.035679 |  | 3.1672 | -1.7168
11 | YES | 350 | 0.62656 | 0.0032375 | -0.39444 | -0.36541 | 0.029026 |  | 3.4144 | -1.8006
11 | YES | 351 | 1.3231 | 0.0017168 | -0.78195 | -0.70886 | 0.07309 |  | 4.3573 | -2.2717
11 | YES | 352 | 2.4287 | 0.0014037 | -1.3573 | -1.2541 | 0.1032 |  | 4.4359 | -2.2861
11 | YES | 353 | 3.0405 | 0.003679 | -1.5525 | -1.5558 | -0.0032409 |  | 4.4173 | -2.2247
11 | NO | 354 | 3.4124 | 0.0028862 | -1.8786 | -1.7392 | 0.13941 |  | 4.4224 | -2.2766
11 | NO | 355 | 3.7705 | 0.0031573 | -2.0525 | -1.9157 | 0.13677 |  | 4.4264 | -2.2626
11 | NO | 356 | 4.0585 | 0.0032993 | -2.1975 | -2.0577 | 0.13979 |  | 4.4256 | -2.2524
11 | YES | 357 | 4.1637 | 0.0017177 | -2.1823 | -2.1096 | 0.072695 |  | 1.4561 | -0.84295
11 | YES | 358 | 4.4143 | 0.0018895 | -2.3017 | -2.2332 | 0.068536 |  | 1.2162 | -0.69182
11 | YES | 359 | 4.3857 | 0.0012572 | -2.1963 | -2.2191 | -0.022776 |  | 0.87144 | -0.47956
11 | YES | 360 | 4.3946 | 0.0015036 | -2.2154 | -2.2235 | -0.0080647 |  | 0.4072 | -0.2328
11 | YES | 361 | 4.4029 | 0.0022066 | -2.2261 | -2.2276 | -0.0015172 |  | 0.22221 | -0.10179
11 | YES | 362 | 4.403 | 0.0018797 | -2.197 | -2.2276 | -0.030652 |  | 0.15563 | -0.058372
11 | YES | 363 | 4.4072 | 0.001711 | -2.2028 | -2.2297 | -0.026848 |  | 0.15604 | -0.11096
12 | NO | 366 | 1.9875 | 0.0056304 | -1.2765 | -1.0365 | 0.23999 |  | 0.31112 | -0.21788
12 | YES | 367 | 1.977 | 0.0031318 | -1.148 | -1.0313 | 0.11668 |  | 1.0677 | -0.64533
12 | YES | 368 | 1.844 | 0 | -1.079 | -0.96575 | 0.11325 |  | 1.9182 | -1.0978
12 | YES | 369 | 1.5765 | 0.0026016 | -0.92051 | -0.83383 | 0.086678 |  | 2.2928 | -1.2942
12 | YES | 370 | 1.2116 | 0.0013283 | -0.68736 | -0.65392 | 0.033433 |  | 2.8803 | -1.5737
12 | YES | 371 | 0.79838 | 0.0016546 | -0.44562 | -0.45014 | -0.0045181 |  | 3.1247 | -1.4983
12 | YES | 372 | 0.32808 | 0.00064014 | -0.16092 | -0.21823 | -0.057312 |  | 2.9449 | -1.5191
12 | YES | 373 | 0.21626 | 0.0017867 | -0.10074 | -0.16309 | -0.062341 |  | 4.3244 | -2.1936
12 | YES | 374 | 0.21597 | 0.00036364 | -0.098033 | -0.16294 | -0.06491 |  | 4.3595 | -2.1795
12 | YES | 375 | 0.32793 | 0.00051211 | -0.17107 | -0.21816 | -0.047089 |  | 4.3625 | -2.2125
12 | YES | 376 | 0.61294 | 0.0017715 | -0.34906 | -0.3587 | -0.0096383 |  | 0.38846 | -0.21554
12 | YES | 377 | 1.2254 | 0.0030304 | -0.68064 | -0.66069 | 0.019949 |  | 0.22507 | -0.13093
12 | YES | 378 | 1.8644 | 0.0019371 | -1.0426 | -0.97579 | 0.066839 |  | 0.16955 | -0.10345
12 | NO | 379 | 2.4838 | 0.0092709 | -1.4142 | -1.2812 | 0.13297 |  | 0.19381 | -0.12119
12 | YES | 380 | 3.0897 | 0.0023932 | -1.4913 | -1.58 | -0.088717 |  | 0.34427 | -0.20373
12 | YES | 381 | 3.2801 | 0.0019439 | -1.7719 | -1.6739 | 0.098039 |  | 1.1684 | -0.66156
12 | YES | 382 | 3.4803 | 0.0025891 | -1.8517 | -1.7726 | 0.079103 |  | 1.5772 | -0.87577
12 | NO | 383 | 3.857 | 0.009024 | -2.063 | -1.9584 | 0.10457 |  | 2.0046 | -1.1463
12 | YES | 384 | 4.1701 | 0.0022345 | -2.1979 | -2.1128 | 0.085154 |  | 2.6787 | -1.4824
12 | YES | 385 | 4.3789 | 0.002038 | -2.2531 | -2.2157 | 0.037392 |  | 3.0428 | -1.6392
12 | YES | 386 | 4.4604 | 0.0012024 | -2.2416 | -2.2559 | -0.014356 |  | 3.2226 | -1.7454
12 | YES | 387 | 4.462 | 0.0016602 | -2.284 | -2.2567 | 0.027344 |  | 3.4016 | -1.7734
12 | YES | 388 | 4.464 | 0.0003873 | -2.261 | -2.2577 | 0.0032984 |  | 4.3486 | -2.2294
12 | YES | 389 | 4.4673 | 0.0012418 | -2.2587 | -2.2593 | -0.00056915 |  | 4.4121 | -2.2349
13 | NO | 392 | 1.9142 | 0.0022484 | -1.1478 | -1.0004 | 0.14745 |  | 4.428 | -2.269
13 | YES | 393 | 1.8954 | 0.0052708 | -1.1126 | -0.99111 | 0.12145 |  | 4.4291 | -2.2529
13 | YES | 394 | 1.7689 | 0.0029451 | -1.0651 | -0.92871 | 0.13639 |  | 4.4308 | -2.2362
13 | YES | 395 | 1.5056 | 0.0026573 | -0.90141 | -0.79887 | 0.10254 |  | 1.1895 | -0.74349
13 | YES | 396 | 1.1867 | 0.0025 | -0.71727 | -0.64164 | 0.075635 |  | 0.90289 | -0.52011
13 | YES | 397 | 0.82421 | 0.00099931 | -0.46879 | -0.46287 | 0.0059193 |  | 0.43963 | -0.31237
13 | YES | 398 | 0.34704 | 0.00045454 | -0.15896 | -0.22758 | -0.068618 |  | 0.2794 | -0.1746
13 | YES | 399 | 0.21146 | 0.0015548 | -0.088537 | -0.16072 | -0.072185 |  | 0.16799 | -0.075008
13 | YES | 400 | 0.17125 | 0.0019928 | -0.070752 | -0.14089 | -0.070139 |  | 0.16296 | -0.11404
13 | YES | 401 | 0.24225 | 0.0024334 | -0.25675 | -0.1759 | 0.08085 |  | 0.35181 | -0.18919
13 | YES | 402 | 0.55191 | 0.0017888 | -0.30009 | -0.3286 | -0.028509 |  | 0.98201 | -0.56999
13 | NO | 403 | 1.3935 | 0.0015659 | -0.94149 | -0.7436 | 0.19788 |  | 1.5411 | -0.87488
13 | YES | 404 | 2.1879 | 0.0017008 | -1.2091 | -1.1353 | 0.073802 |  | 1.8523 | -1.1046
13 | YES | 405 | 2.6173 | 0.0019649 | -1.4097 | -1.3471 | 0.062629 |  | 2.4955 | -1.3585
13 | YES | 406 | 3.102 | 0.001981 | -1.515 | -1.5861 | -0.071052 |  | 2.749 | -1.456
13 | YES | 407 | 3.1672 | 0.0012974 | -1.7168 | -1.6182 | 0.098597 |  | 3.1219 | -1.6541
13 | YES | 408 | 3.4144 | 0.0037168 | -1.8006 | -1.7402 | 0.06041 |  | 4.3425 | -2.2525
13 | NO | 409 | 3.7742 | 0.005372 | -2.0368 | -1.9176 | 0.11922 |  | 4.3829 | -2.2031
13 | YES | 410 | 4.3573 | 0.0032074 | -2.2717 | -2.2051 | 0.066643 |  | 4.4097 | -2.2173
13 | YES | 411 | 4.4359 | 0.00062452 | -2.2861 | -2.2438 | 0.042255 |  | 4.4159 | -2.2171
13 | YES | 412 | 4.4173 | 0.0018245 | -2.2247 | -2.2347 | -0.010048 |  | 4.4178 | -2.2572
13 | YES | 413 | 4.4224 | 0.0011819 | -2.2766 | -2.2372 | 0.039367 |  | 1.1427 | -0.7003
13 | YES | 414 | 4.4264 | 0.0011142 | -2.2626 | -2.2392 | 0.023439 |  | 1.1621 | -0.67292
13 | YES | 415 | 4.4256 | 0.001617 | -2.2524 | -2.2387 | 0.013694 |  | 0.8981 | -0.5019
14 | NO | 418 | 1.6874 | 0.0018858 | -1.0436 | -0.88852 | 0.15509 |  | 0.45113 | -0.24887
14 | YES | 419 | 1.4561 | 0.00054546 | -0.84295 | -0.77444 | 0.068507 |  | 0.26439 | -0.13361
14 | YES | 420 | 1.2162 | 0.0012845 | -0.69182 | -0.65616 | 0.035657 |  | 0.1841 | -0.1299
14 | YES | 421 | 0.87144 | 0.0031699 | -0.47956 | -0.48616 | -0.0066019 |  | 0.14844 | -0.093562
14 | YES | 422 | 0.4072 | 0.00087197 | -0.2328 | -0.25724 | -0.024439 |  | 0.24249 | -0.13651
14 | YES | 423 | 0.22221 | 0.0026106 | -0.10179 | -0.16602 | -0.064227 |  | 0.84912 | -0.46688
14 | YES | 424 | 0.15563 | 0.0017895 | -0.058372 | -0.13319 | -0.074817 |  | 1.3692 | -0.74576
14 | YES | 425 | 0.15604 | 0.0017146 | -0.11096 | -0.13339 | -0.022434 |  | 1.8068 | -0.9932
14 | YES | 426 | 0.31112 | 0.0017088 | -0.21788 | -0.20986 | 0.0080219 |  | 2.2439 | -1.2601
14 | YES | 427 | 1.0677 | 0.0033176 | -0.64533 | -0.58293 | 0.062403 |  | 2.2014 | -1.1856
14 | YES | 428 | 1.9182 | 0.0032042 | -1.0978 | -1.0023 | 0.095494 |  | 2.2263 | -1.1637
14 | YES | 429 | 2.2928 | 0.00092633 | -1.2942 | -1.187 | 0.10717 |  | 2.521 | -1.348
14 | YES | 430 | 2.8803 | 0.002273 | -1.5737 | -1.4767 | 0.096986 |  | 3.0116 | -1.5634
14 | YES | 431 | 3.1247 | 0.0020064 | -1.4983 | -1.5973 | -0.099036 |  | 4.3769 | -2.1901
14 | YES | 432 | 2.9449 | 0.0018055 | -1.5191 | -1.5086 | 0.010516 |  | 4.414 | -2.225
14 | NO | 433 | 3.5627 | 0.002504 | -1.9213 | -1.8132 | 0.10807 |  | 4.4292 | -2.1828
14 | NO | 434 | 4.2226 | 0.0023223 | -2.2354 | -2.1387 | 0.096736 |  | 4.4269 | -2.1911
14 | YES | 435 | 4.3244 | 0.0017313 | -2.1936 | -2.1889 | 0.0046921 |  | 0.44155 | -0.42645
14 | NO | 436 | 4.3228 | 0.0020977 | -2.2862 | -2.1881 | 0.098186 |  | 0.26006 | -0.12394
14 | NO | 437 | 4.3443 | 0.002411 | -2.6577 | -2.1987 | 0.45908 |  | 0.15902 | -0.058984
14 | YES | 438 | 4.3595 | 0.0029918 | -2.1795 | -2.2062 | -0.026627 |  | 0.18711 | -0.11489
14 | YES | 439 | 4.3625 | 0.0029554 | -2.2125 | -2.2076 | 0.0048941 |  | 0.50397 | -0.30103
15 | NO | 441 | 0.84788 | 0.0020904 | -1.2821 | -0.47455 | 0.80757 |  | 1.0882 | -0.61578
15 | YES | 442 | 0.38846 | 0.0015548 | -0.21554 | -0.248 | -0.032466 |  | 1.2397 | -0.65826
15 | YES | 443 | 0.22507 | 0.00072727 | -0.13093 | -0.16743 | -0.036496 |  | 1.3259 | -0.68814
15 | YES | 444 | 0.16955 | 0.0019577 | -0.10345 | -0.14006 | -0.03661 |  | 1.4877 | -0.80225
15 | YES | 445 | 0.19381 | 0.0018363 | -0.12119 | -0.15202 | -0.030827 |  | 1.5429 | -0.80311
15 | YES | 446 | 0.34427 | 0.0015492 | -0.20373 | -0.22621 | -0.022485 |  | 1.6147 | -0.8403
15 | YES | 447 | 1.1684 | 0.0015963 | -0.66156 | -0.63262 | 0.028943 |  | 1.7423 | -0.90175
15 | YES | 448 | 1.5772 | 0.0039386 | -0.87577 | -0.8342 | 0.04157 |  | 1.9836 | -1.0214
15 | YES | 449 | 2.0046 | 0.002128 | -1.1463 | -1.045 | 0.10138 |  | 2.3543 | -1.2107
15 | YES | 450 | 2.6787 | 0.0081011 | -1.4824 | -1.3773 | 0.10503 |  | 2.8475 | -1.4535
15 | YES | 451 | 3.0428 | 0.0013396 | -1.6392 | -1.5569 | 0.082265 |  | 4.4148 | -2.2462
15 | YES | 452 | 3.2226 | 0.0019595 | -1.7454 | -1.6456 | 0.099833 |  | 4.4787 | -2.3363
15 | YES | 453 | 3.4016 | 0.0017677 | -1.7734 | -1.7338 | 0.039581 |  | 4.5417 | -2.2833
15 | YES | 454 | 4.3486 | 0.0013157 | -2.2294 | -2.2008 | 0.028544
15 | YES | 455 | 4.4121 | 0.00081207 | -2.2349 | -2.2321 | 0.0027611
15 | YES | 456 | 4.428 | 0.0012676 | -2.269 | -2.24 | 0.028991
15 | YES | 457 | 4.4291 | 0.0013817 | -2.2529 | -2.2405 | 0.012378
15 | YES | 458 | 4.4308 | 0.0022261 | -2.2362 | -2.2413 | -0.0051452
15 | NO | 459 | 4.4309 | 0.00048819 | -2.3651 | -2.2414 | 0.12366
16 | YES | 461 | 1.1895 | 0.0015007 | -0.74349 | -0.64301 | 0.10048
16 | YES | 462 | 0.90289 | 0.0024793 | -0.52011 | -0.50167 | 0.018432
16 | YES | 463 | 0.43963 | 0.0013174 | -0.31237 | -0.27323 | 0.039139
16 | YES | 464 | 0.2794 | 0.0012005 | -0.1746 | -0.19422 | -0.019618
16 | YES | 465 | 0.16799 | 0.00058446 | -0.075008 | -0.13929 | -0.064278
16 | NO | 0 | 0 | 0 | -1.383 | -0.056447 | 1.3266
16 | YES | 466 | 0.16296 | 0.0017146 | -0.11404 | -0.1368 | -0.022763
16 | YES | 467 | 0.35181 | 0.0019075 | -0.18919 | -0.22993 | -0.040739
16 | YES | 468 | 0.98201 | 0.0017344 | -0.56999 | -0.54069 | 0.029305
16 | YES | 469 | 1.5411 | 0.0016861 | -0.87488 | -0.81639 | 0.058483
16 | YES | 470 | 1.8523 | 0.0019439 | -1.1046 | -0.96986 | 0.13479
16 | YES | 471 | 2.4955 | 0.0034907 | -1.3585 | -1.287 | 0.071506
16 | YES | 472 | 2.749 | 0.001668 | -1.456 | -1.412 | 0.043942
16 | NO | 473 | 2.8798 | 0.0018226 | -1.6382 | -1.4765 | 0.16167
16 | YES | 474 | 3.1219 | 0.0013291 | -1.6541 | -1.5959 | 0.05819
16 | NO | 475 | 3.5725 | 0.0039436 | -2.0075 | -1.8181 | 0.18942
16 | YES | 476 | 4.3425 | 0.0017893 | -2.2525 | -2.1978 | 0.054712
16 | YES | 477 | 4.3829 | 0.0013355 | -2.2031 | -2.2177 | -0.014625
16 | YES | 478 | 4.4097 | 0.0021378 | -2.2173 | -2.2309 | -0.013596
16 | YES | 479 | 4.4159 | 0.0015077 | -2.2171 | -2.234 | -0.016883
16 | YES | 480 | 4.4178 | 0.0017766 | -2.2572 | -2.2349 | 0.022235
17 | YES | 483 | 1.1427 | 0.0025483 | -0.7003 | -0.61993 | 0.080371
17 | YES | 484 | 1.1621 | 0.00090911 | -0.67292 | -0.62948 | 0.043434
17 | YES | 485 | 0.8981 | 0.0024372 | -0.5019 | -0.49931 | 0.0025902
17 | YES | 486 | 0.45113 | 0.0021831 | -0.24887 | -0.27891 | -0.030038
17 | YES | 487 | 0.26439 | 0.0015512 | -0.13361 | -0.18682 | -0.053208
17 | YES | 488 | 0.1841 | 0.00062456 | -0.1299 | -0.14723 | -0.017327
17 | YES | 489 | 0.14844 | 0.0013958 | -0.093562 | -0.12964 | -0.036082
17 | YES | 490 | 0.24249 | 0.0019413 | -0.13651 | -0.17602 | -0.039509
17 | YES | 491 | 0.84912 | 0.0018054 | -0.46688 | -0.47516 | -0.0082846
17 | YES | 492 | 1.3692 | 0.0009489 | -0.74576 | -0.73164 | 0.014124
17 | YES | 493 | 1.8068 | 0.0018557 | -0.9932 | -0.9474 | 0.045795
17 | YES | 494 | 2.2439 | 0.0033862 | -1.2601 | -1.1629 | 0.097143
17 | YES | 495 | 2.2014 | 0.0032068 | -1.1856 | -1.142 | 0.043675
17 | YES | 496 | 2.2263 | 0.0069041 | -1.1637 | -1.1543 | 0.0094216
17 | YES | 497 | 2.521 | 0.0025478 | -1.348 | -1.2996 | 0.048446
17 | YES | 498 | 3.0116 | 0.0016894 | -1.5634 | -1.5415 | 0.021865
17 | NO | 499 | 3.5302 | 0.0068678 | -1.9608 | -1.7972 | 0.16363
17 | YES | 500 | 4.3769 | 0.0019192 | -2.1901 | -2.2148 | -0.024666
17 | YES | 501 | 4.414 | 0.001449 | -2.225 | -2.233 | -0.008046
17 | YES | 502 | 4.4292 | 0.0009602 | -2.1828 | -2.2405 | -0.057726
17 | YES | 503 | 4.4269 | 0.0010634 | -2.1911 | -2.2394 | -0.048382
18 | NO | 506 | 0.69494 | 0.00077779 | -0.58406 | -0.39913 | 0.18493
18 | YES | 507 | 0.44155 | 0.0014434 | -0.42645 | -0.27418 | 0.15228
18 | YES | 508 | 0.26006 | 0.0017715 | -0.12394 | -0.18468 | -0.060743
18 | YES | 509 | 0.15902 | 0.0018529 | -0.058984 | -0.13486 | -0.075877
18 | YES | 510 | 0.18711 | 0.0016723 | -0.11489 | -0.14871 | -0.033819
18 | YES | 511 | 0.50397 | 0.00036364 | -0.30103 | -0.30496 | -0.0039263
18 | YES | 512 | 1.0882 | 0.0018506 | -0.61578 | -0.59306 | 0.022715
18 | YES | 513 | 1.2397 | 0.00099809 | -0.65826 | -0.66778 | -0.0095133
18 | YES | 514 | 1.3259 | 0.00064942 | -0.68814 | -0.71025 | -0.022104
18 | YES | 515 | 1.4877 | 0.0026054 | -0.80225 | -0.79008 | 0.012173
18 | YES | 516 | 1.5429 | 0.0016723 | -0.80311 | -0.81727 | -0.014158
18 | YES | 517 | 1.6147 | 0.0017061 | -0.8403 | -0.85268 | -0.012379
18 | YES | 518 | 1.7423 | 0.0015614 | -0.90175 | -0.91557 | -0.01382
18 | YES | 519 | 1.9836 | 0.0018574 | -1.0214 | -1.0346 | -0.013245
18 | YES | 520 | 2.3543 | 0.0017361 | -1.2107 | -1.2174 | -0.0066968
18 | YES | 521 | 2.8475 | 0.0021951 | -1.4535 | -1.4606 | -0.0070701
18 | YES | 522 | 4.4148 | 0.0016083 | -2.2462 | -2.2334 | 0.012774
18 | YES | 523 | 4.4787 | 0.0017787 | -2.3363 | -2.2649 | 0.071365
18 | YES | 524 | 4.5417 | 0.0016444 | -2.2833 | -2.296 | -0.012687
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
NO
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. CTD DOX
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.66823 |  | X | CALC Y
m= | 0.0017553 | 7.3134e-05 |  | Std Dev of Resid |  | 0.48188 |  | 5.0346 | -1.8418
b= | -1.8507 | 0.038029 |  | Resid Sum of Squares |  | 66.4113 |  | 1499.71 | 0.78174
# of Values Extracted = |  |  | 367
# of Values Used in Fit = |  |  | 288
File | Use | Sample # | Pressure | Pressure | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 4 | 150.116 | 0.031536 | 0.004 | -1.5872 | -1.5912 |  | 123.98 | -0.05324
1 | YES | 5 | 123.98 | 0.034371 | -0.05324 | -1.6331 | -1.5798 |  | 100.069 | -0.0012562
1 | YES | 6 | 100.069 | 0.033628 | -0.0012562 | -1.675 | -1.6738 |  | 74.9138 | -1.0448
1 | YES | 7 | 74.9138 | 0.027111 | -1.0448 | -1.7192 | -0.67439 |  | 51.3413 | -1.1088
1 | YES | 8 | 51.3413 | 0.027671 | -1.1088 | -1.7606 | -0.65179 |  | 25.1009 | -1.264
1 | YES | 9 | 25.1009 | 0.027476 | -1.264 | -1.8066 | -0.54259 |  | 9.9867 | -1.411
1 | YES | 10 | 9.9867 | 0.029191 | -1.411 | -1.8332 | -0.42213 |  | 112.872 | -1.3573
1 | NO | 11 | 4.7859 | 0.036172 | -1.3767 | -1.8423 | -0.46561 |  | 99.5313 | -1.4682
2 | YES | 13 | 112.872 | 0.36499 | -1.3573 | -1.6526 | -0.29527 |  | 75.7888 | -1.4437
2 | YES | 14 | 99.5313 | 0.83782 | -1.4682 | -1.676 | -0.2078 |  | 75.6808 | -1.4255
2 | YES | 15 | 75.7888 | 0.27621 | -1.4437 | -1.7176 | -0.27397 |  | 49.7631 | -1.4218
2 | YES | 16 | 75.6808 | 0.77145 | -1.4255 | -1.7178 | -0.29234 |  | 24.9832 | -1.4634
2 | YES | 17 | 49.7631 | 0.43775 | -1.4218 | -1.7633 | -0.34157 |  | 9.3925 | -1.4856
2 | YES | 18 | 24.9832 | 0.39083 | -1.4634 | -1.8068 | -0.3434 |  | 5.0346 | -1.5192
2 | YES | 19 | 9.3925 | 0.71696 | -1.4856 | -1.8342 | -0.34857 |  | 1303.76 | -0.1977
2 | YES | 20 | 5.0346 | 0.27685 | -1.5192 | -1.8418 | -0.32262 |  | 999.212 | -0.099
3 | YES | 22 | 1303.76 | 0.35288 | -0.1977 | 0.43779 | 0.6355 |  | 997.807 | -0.12207
3 | YES | 23 | 999.212 | 1.6259 | -0.099 | -0.096776 | 0.0022241 |  | 801.543 | -0.17247
3 | YES | 24 | 997.807 | 1.2004 | -0.12207 | -0.099242 | 0.022832 |  | 802.404 | -0.2097
3 | YES | 25 | 801.543 | 0.70444 | -0.17247 | -0.44374 | -0.27127 |  | 599.013 | -0.2234
3 | YES | 26 | 802.404 | 0.71701 | -0.2097 | -0.44223 | -0.23253 |  | 398.581 | -0.40936
3 | YES | 27 | 599.013 | 0.91443 | -0.2234 | -0.79924 | -0.57584 |  | 299.775 | -0.70209
3 | YES | 28 | 398.581 | 0.17745 | -0.40936 | -1.1511 | -0.7417 |  | 200.878 | -0.7866
3 | YES | 29 | 299.775 | 0.14397 | -0.70209 | -1.3245 | -0.6224 |  | 176.117 | -0.7414
3 | YES | 31 | 200.878 | 0.85795 | -0.7866 | -1.4981 | -0.71149 |  | 153.872 | -0.80164
3 | YES | 32 | 176.117 | 0.4941 | -0.7414 | -1.5415 | -0.80015 |  | 122.952 | -0.94776
3 | YES | 33 | 153.872 | 1.0608 | -0.80164 | -1.5806 | -0.77895 |  | 103.636 | -1.3317
3 | YES | 34 | 122.952 | 1.2095 | -0.94776 | -1.6349 | -0.6871 |  | 77.44 | -1.5895
3 | YES | 35 | 103.636 | 0.47119 | -1.3317 | -1.6688 | -0.3371 |  | 49.7815 | -1.7234
3 | YES | 36 | 77.44 | 1.0386 | -1.5895 | -1.7148 | -0.12522 |  | 24.3834 | -1.7332
3 | YES | 37 | 49.7815 | 0.76289 | -1.7234 | -1.7633 | -0.039939 |  | 10.7183 | -1.7302
3 | YES | 38 | 24.3834 | 0.71988 | -1.7332 | -1.8079 | -0.07466 |  | 5.4036 | -1.7118
3 | YES | 39 | 10.7183 | 0.73741 | -1.7302 | -1.8319 | -0.10172 |  | 1251.74 | -0.1416
3 | YES | 40 | 5.4036 | 0.55856 | -1.7118 | -1.8412 | -0.12941 |  | 998.815 | -0.11782
4 | NO | 58 | 2532.73 | 1.2074 | -0.64814 | 2.595 | 3.2431 |  | 800.15 | -0.10126
4 | NO | 59 | 2000.38 | 0.3458 | -0.53817 | 1.6606 | 2.1987 |  | 599.996 | -0.24649
4 | NO | 60 | 1503.05 | 0.4089 | -0.218 | 0.7876 | 1.0056 |  | 601.442 | -0.26674
4 | YES | 61 | 1251.74 | 0.60806 | -0.1416 | 0.34648 | 0.48809 |  | 399.031 | -0.54804
4 | YES | 62 | 998.815 | 0.48144 | -0.11782 | -0.097473 | 0.020345 |  | 300.776 | -0.86334
4 | YES | 63 | 800.15 | 0.12118 | -0.10126 | -0.44619 | -0.34492 |  | 250.259 | -1.0023
4 | YES | 64 | 599.996 | 0.23264 | -0.24649 | -0.79751 | -0.55102 |  | 199.679 | -1.2267
4 | YES | 65 | 601.442 | 0.47932 | -0.26674 | -0.79498 | -0.52823 |  | 175.212 | -1.3421
4 | YES | 66 | 399.031 | 0.75268 | -0.54804 | -1.1503 | -0.60222 |  | 150.493 | -1.4324
4 | YES | 67 | 300.776 | 0.97634 | -0.86334 | -1.3227 | -0.45939 |  | 125.563 | -3.2447
4 | YES | 68 | 250.259 | 0.53014 | -1.0023 | -1.4114 | -0.40906 |  | 100.487 | -1.4465
4 | YES | 69 | 199.679 | 0.23447 | -1.2267 | -1.5002 | -0.27349 |  | 75.8665 | -2.0579
4 | YES | 70 | 175.212 | 0.13585 | -1.3421 | -1.5431 | -0.20099 |  | 49.2447 | -2.6084
4 | YES | 71 | 150.493 | 0.59836 | -1.4324 | -1.5865 | -0.15416 |  | 24.716 | -2.0307
4 | YES | 72 | 125.563 | 0.46488 | -3.2447 | -1.6303 | 1.6144 |  | 9.0068 | -2.0381
4 | YES | 73 | 100.487 | 0.78314 | -1.4465 | -1.6743 | -0.22784 |  | 5.0642 | -1.9826
4 | YES | 74 | 75.8665 | 0.90257 | -2.0579 | -1.7175 | 0.34036 |  | 1251.58 | -0.18648
4 | YES | 75 | 49.2447 | 0.31268 | -2.6084 | -1.7642 | 0.84413 |  | 1001.11 | -0.1434
4 | YES | 76 | 24.716 | 1.0151 | -2.0307 | -1.8073 | 0.22345 |  | 1001.51 | -0.13086
4 | YES | 77 | 9.0068 | 0.67514 | -2.0381 | -1.8349 | 0.20318 |  | 800.83 | -0.44645
4 | YES | 78 | 5.0642 | 0.2947 | -1.9826 | -1.8418 | 0.14082 |  | 600.214 | -0.19421
5 | NO | 82 | 3236.62 | 0.64341 | -0.95205 | 3.8305 | 4.7826 |  | 401.523 | -0.41079
5 | NO | 83 | 2998.97 | 0.21126 | -0.8287 | 3.4134 | 4.2421 |  | 300.986 | -0.58407
5 | NO | 84 | 2497 | 1.1161 | -0.7406 | 2.5323 | 3.2729 |  | 248.325 | -0.78649
5 | NO | 85 | 2002.86 | 0.98787 | -0.50653 | 1.6649 | 2.1714 |  | 199.613 | -0.95894
5 | NO | 86 | 1501.1 | 0.63182 | -0.31036 | 0.78418 | 1.0945 |  | 175.175 | -0.87647
5 | YES | 87 | 1251.58 | 0.73213 | -0.18648 | 0.3462 | 0.53268 |  | 148.3 | -1.2899
5 | YES | 88 | 1001.11 | 0.73318 | -0.1434 | -0.093444 | 0.049953 |  | 123.554 | -1.4495
5 | YES | 89 | 1001.51 | 0.61447 | -0.13086 | -0.092742 | 0.038117 |  | 100.957 | -1.946
5 | YES | 90 | 800.83 | 0.17902 | -0.44645 | -0.44499 | 0.0014617 |  | 75.0847 | -2.1335
5 | YES | 91 | 600.214 | 0.29665 | -0.19421 | -0.79713 | -0.60292 |  | 46.6954 | -2.1058
5 | YES | 92 | 401.523 | 0.63925 | -0.41079 | -1.1459 | -0.7351 |  | 23.1122 | -2.0639
5 | YES | 93 | 300.986 | 0.34858 | -0.58407 | -1.3224 | -0.7383 |  | 9.6326 | -2.0937
5 | YES | 94 | 248.325 | 0.7745 | -0.78649 | -1.4148 | -0.62831 |  | 1499.45 | -0.17702
5 | YES | 95 | 199.613 | 0.39718 | -0.95894 | -1.5003 | -0.54136 |  | 1251.36 | -0.1298
5 | YES | 96 | 175.175 | 0.10193 | -0.87647 | -1.5432 | -0.66673 |  | 1250.75 | -0.090719
5 | YES | 97 | 148.3 | 0.46924 | -1.2899 | -1.5904 | -0.30043 |  | 1000.08 | -0.07
5 | YES | 98 | 123.554 | 0.55859 | -1.4495 | -1.6338 | -0.1843 |  | 798.738 | -0.080546
5 | YES | 99 | 100.957 | 0.94996 | -1.946 | -1.6735 | 0.27251 |  | 599.267 | -0.23669
5 | YES | 100 | 75.0847 | 0.34425 | -2.1335 | -1.7189 | 0.41464 |  | 401.662 | -0.57549
5 | YES | 101 | 46.6954 | 0.38691 | -2.1058 | -1.7687 | 0.3371 |  | 300.21 | -0.84797
5 | YES | 102 | 23.1122 | 0.85177 | -2.0639 | -1.8101 | 0.2538 |  | 248.47 | -0.97879
5 | YES | 103 | 9.6326 | 0.55622 | -2.0937 | -1.8338 | 0.25989 |  | 200.389 | -1.3153
5 | NO | 104 | 4.3149 | 0.52709 | -2.1109 | -1.8431 | 0.26784 |  | 173.467 | -1.4349
6 | NO | 122 | 3687.26 | 0.36633 | -1.0939 | 4.6215 | 5.7154 |  | 149.968 | -1.7284
6 | NO | 123 | 3500.58 | 0.1461 | -0.95145 | 4.2938 | 5.2453 |  | 124.238 | -1.8489
6 | NO | 124 | 3000.69 | 0.44202 | -0.7953 | 3.4164 | 4.2117 |  | 101.299 | -2.108
6 | NO | 125 | 2501 | 0.61918 | -0.6011 | 2.5393 | 3.1404 |  | 75.0294 | -2.2185
6 | NO | 126 | 2001.35 | 0.55271 | -0.4314 | 1.6623 | 2.0937 |  | 49.558 | -2.2005
6 | YES | 127 | 1499.45 | 0.144 | -0.17702 | 0.78129 | 0.9583 |  | 25.6428 | -2.1808
6 | YES | 128 | 1251.36 | 0.29601 | -0.1298 | 0.34582 | 0.47562 |  | 10.5445 | -2.1774
6 | YES | 129 | 1250.75 | 0.11873 | -0.090719 | 0.34475 | 0.43546 |  | 1250.62 | -0.084479
6 | YES | 130 | 1000.08 | 0.34883 | -0.07 | -0.095252 | -0.025252 |  | 1001.12 | -0.068397
6 | YES | 131 | 798.738 | 0.43931 | -0.080546 | -0.44867 | -0.36812 |  | 801.782 | -0.10731
6 | YES | 132 | 599.267 | 0.85277 | -0.23669 | -0.79879 | -0.56211 |  | 600.598 | -0.14869
6 | YES | 133 | 401.662 | 0.38839 | -0.57549 | -1.1456 | -0.57016 |  | 399.57 | -0.49131
6 | YES | 134 | 300.21 | 0.23693 | -0.84797 | -1.3237 | -0.47576 |  | 302.34 | -0.67841
6 | YES | 135 | 248.47 | 0.45633 | -0.97879 | -1.4145 | -0.43576 |  | 250.15 | -0.91867
6 | YES | 136 | 200.389 | 0.89934 | -1.3153 | -1.4989 | -0.18362 |  | 202.2 | -1.2879
6 | YES | 137 | 173.467 | 0.53286 | -1.4349 | -1.5462 | -0.11128 |  | 175.838 | -1.211
6 | YES | 138 | 149.968 | 0.64972 | -1.7284 | -1.5874 | 0.14092 |  | 150.782 | -1.521
6 | YES | 139 | 124.238 | 0.25195 | -1.8489 | -1.6326 | 0.21627 |  | 124.476 | -1.7093
6 | YES | 140 | 101.299 | 0.11158 | -2.108 | -1.6729 | 0.43516 |  | 100.209 | -2.2897
6 | YES | 141 | 75.0294 | 0.43513 | -2.2185 | -1.719 | 0.49954 |  | 73.7679 | -2.2709
6 | YES | 142 | 49.558 | 0.55966 | -2.2005 | -1.7637 | 0.43684 |  | 48.2123 | -2.2669
6 | YES | 143 | 25.6428 | 0.63129 | -2.1808 | -1.8057 | 0.37512 |  | 24.2342 | -2.2247
6 | YES | 144 | 10.5445 | 0.34888 | -2.1774 | -1.8322 | 0.34528 |  | 10.0164 | -2.2487
6 | NO | 145 | 4.9337 | 0.23253 | -2.1664 | -1.842 | 0.32434 |  | 5.2329 | -2.2656
7 | NO | 186 | 4028.8 | 0.32073 | -1.1079 | 5.221 | 6.3289 |  | 1248.49 | -0.1488
7 | NO | 187 | 3500.27 | 0.6253 | -1.0223 | 4.2933 | 5.3156 |  | 999.947 | -0.12826
7 | NO | 188 | 2999.04 | 0.22154 | -0.86338 | 3.4135 | 4.2769 |  | 828.006 | -0.11748
7 | NO | 189 | 2999.58 | 0.28575 | -0.85335 | 3.4144 | 4.2678 |  | 598.926 | -0.15055
7 | NO | 190 | 2501.35 | 0.66389 | -0.68212 | 2.5399 | 3.222 |  | 402.571 | -0.35693
7 | NO | 191 | 1999.44 | 0.19311 | -0.42386 | 1.6589 | 2.0828 |  | 300.608 | -0.52769
7 | NO | 192 | 1500.66 | 0.76893 | -0.16389 | 0.78341 | 0.9473 |  | 250.852 | -0.72222
7 | YES | 193 | 1250.62 | 0.45739 | -0.084479 | 0.34452 | 0.429 |  | 202.783 | -0.84868
7 | YES | 194 | 1001.12 | 0.88904 | -0.068397 | -0.093427 | -0.02503 |  | 175.01 | -1.2137
7 | YES | 195 | 801.782 | 0.14669 | -0.10731 | -0.44332 | -0.33601 |  | 149.167 | -1.3835
7 | YES | 196 | 600.598 | 0.78588 | -0.14869 | -0.79646 | -0.64776 |  | 123.53 | -1.8549
7 | YES | 197 | 399.57 | 0.083243 | -0.49131 | -1.1493 | -0.65801 |  | 98.4833 | -2.3795
7 | YES | 198 | 302.34 | 0.38527 | -0.67841 | -1.32 | -0.64157 |  | 74.6983 | -2.3988
7 | YES | 199 | 250.15 | 0.38326 | -0.91867 | -1.4116 | -0.49293 |  | 50.6391 | -2.319
7 | YES | 200 | 202.2 | 0.40927 | -1.2879 | -1.4958 | -0.20786 |  | 25.6736 | -2.3495
7 | YES | 201 | 175.838 | 0.54872 | -1.211 | -1.542 | -0.33098 |  | 11.6197 | -2.357
7 | YES | 202 | 150.782 | 0.82894 | -1.521 | -1.586 | -0.065034 |  | 1248.27 | -0.16171
7 | YES | 203 | 124.476 | 0.32172 | -1.7093 | -1.6322 | 0.077121 |  | 998.363 | -0.13557
7 | YES | 204 | 100.209 | 0.54753 | -2.2897 | -1.6748 | 0.61496 |  | 799.043 | -0.22807
7 | YES | 205 | 73.7679 | 0.2861 | -2.2709 | -1.7212 | 0.54972 |  | 599.199 | -0.31764
7 | YES | 206 | 48.2123 | 0.56667 | -2.2669 | -1.7661 | 0.50087 |  | 399.813 | -0.44714
7 | YES | 207 | 24.2342 | 0.38071 | -2.2247 | -1.8081 | 0.41654 |  | 302.534 | -0.73312
7 | YES | 208 | 10.0164 | 0.076046 | -2.2487 | -1.8331 | 0.41555 |  | 250.467 | -1.002
7 | YES | 209 | 5.2329 | 0.082847 | -2.2656 | -1.8415 | 0.42411 |  | 201.606 | -1.5052
8 | NO | 258 | 4310.57 | 0.39195 | -1.2647 | 5.7156 | 6.9803 |  | 176.595 | -1.7338
8 | NO | 259 | 3999.69 | 0.63828 | -1.1266 | 5.1699 | 6.2966 |  | 149.272 | -1.9894
8 | NO | 260 | 3501.07 | 0.30709 | -1.0315 | 4.2947 | 5.3262 |  | 124.915 | -2.2831
8 | NO | 261 | 2998.49 | 0.15812 | -0.89612 | 3.4125 | 4.3087 |  | 101.107 | -2.2789
8 | NO | 262 | 2499.86 | 0.21967 | -0.67139 | 2.5373 | 3.2087 |  | 72.8961 | -2.0531
8 | NO | 263 | 2001.45 | 0.63236 | -0.4547 | 1.6624 | 2.1171 |  | 50.1174 | -2.2848
8 | NO | 264 | 1500.75 | 0.29274 | -0.19967 | 0.78357 | 0.98324 |  | 25.4089 | -2.2878
8 | YES | 265 | 1248.49 | 0.28315 | -0.1488 | 0.34078 | 0.48958 |  | 10.7292 | -2.2749
8 | YES | 266 | 999.947 | 1.3105 | -0.12826 | -0.095486 | 0.03277 |  | 5.3663 | -2.2713
8 | YES | 267 | 828.006 | 2.0944 | -0.11748 | -0.39729 | -0.27981 |  | 1498.98 | -0.182
8 | YES | 268 | 598.926 | 0.68876 | -0.15055 | -0.79939 | -0.64885 |  | 1251.03 | -0.16666
8 | YES | 269 | 402.571 | 0.61032 | -0.35693 | -1.1441 | -0.78712 |  | 1001.73 | -0.18773
8 | YES | 270 | 300.608 | 0.75269 | -0.52769 | -1.323 | -0.79534 |  | 798.45 | -0.21784
8 | YES | 271 | 250.852 | 0.22462 | -0.72222 | -1.4104 | -0.68814 |  | 603.792 | -0.27803
8 | YES | 272 | 202.783 | 1.25 | -0.84868 | -1.4947 | -0.64606 |  | 400.409 | -0.4945
8 | YES | 273 | 175.01 | 1.3186 | -1.2137 | -1.5435 | -0.32975 |  | 299.06 | -0.75237
8 | YES | 274 | 149.167 | 0.46681 | -1.3835 | -1.5888 | -0.20537 |  | 250.306 | -1.0257
8 | YES | 275 | 123.53 | 0.7145 | -1.8549 | -1.6338 | 0.22107 |  | 202.134 | -1.5291
8 | YES | 276 | 98.4833 | 0.23182 | -2.3795 | -1.6778 | 0.70171 |  | 176.465 | -2.0452
8 | YES | 277 | 74.6983 | 0.76241 | -2.3988 | -1.7196 | 0.67928 |  | 151.844 | -2.2591
8 | YES | 278 | 50.6391 | 0.4675 | -2.319 | -1.7618 | 0.55724 |  | 123.934 | -2.2598
8 | YES | 279 | 25.6736 | 0.4654 | -2.3495 | -1.8056 | 0.54383 |  | 100.078 | -2.3028
8 | YES | 280 | 11.6197 | 0.3429 | -2.357 | -1.8303 | 0.52668 |  | 75.1951 | -2.2638
8 | NO | 281 | 4.6998 | 0.14191 | -2.3403 | -1.8424 | 0.49792 |  | 50.1043 | -2.2843
9 | NO | 296 | 2001.13 | 1.0251 | -0.48255 | 1.6619 | 2.1444 |  | 26.2796 | -2.2945
9 | NO | 297 | 1500.21 | 0.34241 | -0.18203 | 0.78262 | 0.96465 |  | 10.422 | -2.2897
9 | YES | 298 | 1248.27 | 0.76804 | -0.16171 | 0.34039 | 0.5021 |  | 5.39 | -2.2981
9 | YES | 299 | 998.363 | 0.99065 | -0.13557 | -0.098266 | 0.037304 |  | 1260.5 | -0.10423
9 | YES | 300 | 799.043 | 0.57849 | -0.22807 | -0.44813 | -0.22006 |  | 1002.15 | -0.096628
9 | YES | 301 | 599.199 | 0.57864 | -0.31764 | -0.79891 | -0.48128 |  | 798.81 | -0.19291
9 | YES | 302 | 399.813 | 0.79705 | -0.44714 | -1.1489 | -0.70175 |  | 593.42 | -0.39444
9 | YES | 303 | 302.534 | 1.5837 | -0.73312 | -1.3196 | -0.58653 |  | 399.468 | -0.78195
9 | YES | 304 | 250.467 | 0.38517 | -1.002 | -1.411 | -0.40902 |  | 297.774 | -1.3573
9 | YES | 305 | 201.606 | 0.55856 | -1.5052 | -1.4968 | 0.0084465 |  | 249.757 | -1.5525
9 | YES | 306 | 176.595 | 1.0024 | -1.7338 | -1.5407 | 0.19308 |  | 201.071 | -1.8786
9 | YES | 307 | 149.272 | 0.62809 | -1.9894 | -1.5887 | 0.40079 |  | 175.289 | -2.0525
9 | YES | 308 | 124.915 | 0.28729 | -2.2831 | -1.6314 | 0.6517 |  | 149.378 | -2.1975
9 | YES | 309 | 101.107 | 0.70513 | -2.2789 | -1.6732 | 0.6057 |  | 125.288 | -2.1823
9 | YES | 310 | 72.8961 | 1.519 | -2.0531 | -1.7227 | 0.33038 |  | 98.579 | -2.3017
9 | YES | 311 | 50.1174 | 0.30115 | -2.2848 | -1.7627 | 0.52211 |  | 75.6397 | -2.1963
9 | YES | 312 | 25.4089 | 1.2266 | -2.2878 | -1.8061 | 0.48169 |  | 49.8857 | -2.2154
9 | YES | 313 | 10.7292 | 0.4704 | -2.2749 | -1.8318 | 0.44309 |  | 24.7805 | -2.2261
9 | YES | 314 | 5.3663 | 0.18718 | -2.2713 | -1.8413 | 0.43008 |  | 10.087 | -2.197
10 | NO | 315 | 4880.81 | 0.20966 | -1.1033 | 6.7165 | 7.8199 |  | 1250.28 | -0.10074
10 | NO | 316 | 4001.56 | 0.40515 | -1.2132 | 5.1732 | 6.3864 |  | 1000.49 | -0.098033
10 | NO | 317 | 3499.19 | 0.50342 | -1.1142 | 4.2914 | 5.4056 |  | 801.742 | -0.17107
10 | NO | 318 | 2998.99 | 0.25781 | -1.0156 | 3.4134 | 4.429 |  | 600.251 | -0.34906
10 | NO | 319 | 2501.33 | 0.088744 | -0.71593 | 2.5399 | 3.2558 |  | 398.905 | -0.68064
10 | NO | 320 | 2000.25 | 0.13442 | -0.4561 | 1.6603 | 2.1164 |  | 301.608 | -1.0426
10 | YES | 321 | 1498.98 | 1.1157 | -0.182 | 0.78046 | 0.96246 |  | 251.725 | -1.4142
10 | YES | 322 | 1251.03 | 0.60304 | -0.16666 | 0.34524 | 0.5119 |  | 200.658 | -1.4913
10 | YES | 323 | 1001.73 | 0.70181 | -0.18773 | -0.092356 | 0.095371 |  | 175.629 | -1.7719
10 | YES | 324 | 798.45 | 0.87808 | -0.21784 | -0.44917 | -0.23134 |  | 148.758 | -1.8517
10 | YES | 325 | 603.792 | 0.26731 | -0.27803 | -0.79085 | -0.51283 |  | 125.149 | -2.063
10 | YES | 326 | 400.409 | 0.29807 | -0.4945 | -1.1478 | -0.65335 |  | 100.411 | -2.1979
10 | YES | 327 | 299.06 | 0.15771 | -0.75237 | -1.3257 | -0.57337 |  | 77.0474 | -2.2531
10 | YES | 328 | 250.306 | 0.54952 | -1.0257 | -1.4113 | -0.38558 |  | 50.7156 | -2.2416
10 | YES | 329 | 202.134 | 0.4043 | -1.5291 | -1.4959 | 0.033183 |  | 25.2391 | -2.284
10 | YES | 330 | 176.465 | 0.42054 | -2.0452 | -1.5409 | 0.50429 |  | 10.1999 | -2.261
10 | YES | 331 | 151.844 | 0.53805 | -2.2591 | -1.5842 | 0.67492 |  | 1499.71 | -0.15896
10 | YES | 332 | 123.934 | 0.52989 | -2.2598 | -1.6331 | 0.62665 |  | 1249.27 | -0.088537
10 | YES | 333 | 100.078 | 0.54369 | -2.3028 | -1.675 | 0.62784 |  | 1000.34 | -0.070752
10 | YES | 334 | 75.1951 | 0.32323 | -2.2638 | -1.7187 | 0.54513 |  | 801.259 | -0.25675
10 | YES | 335 | 50.1043 | 0.31573 | -2.2843 | -1.7627 | 0.52161 |  | 598.737 | -0.30009
10 | YES | 336 | 26.2796 | 0.65328 | -2.2945 | -1.8046 | 0.48997 |  | 399.773 | -0.94149
10 | YES | 337 | 10.422 | 0.4802 | -2.2897 | -1.8324 | 0.45735 |  | 301.302 | -1.2091
10 | YES | 338 | 5.39 | 0.29041 | -2.2981 | -1.8412 | 0.45688 |  | 250.568 | -1.4097
11 | NO | 340 | 4569.91 | 0.23267 | -1.1954 | 6.1708 | 7.3662 |  | 199.803 | -1.515
11 | NO | 0 | 0 | 0 | -3.175 | -1.8507 | 1.3243 |  | 175.51 | -1.7168
11 | NO | 342 | 3501.73 | 0.17905 | -1.0723 | 4.2959 | 5.3682 |  | 149.628 | -1.8006
11 | NO | 343 | 2990.45 | 0.1111 | -0.91895 | 3.3984 | 4.3174 |  | 125.132 | -2.0368
11 | NO | 344 | 2488.34 | 1.0528 | -0.67337 | 2.5171 | 3.1904 |  | 99.9763 | -2.2717
11 | NO | 345 | 2009.89 | 0.69746 | -0.45441 | 1.6773 | 2.1317 |  | 74.4035 | -2.2861
11 | NO | 346 | 1507.32 | 0.46572 | -0.15266 | 0.7951 | 0.94776 |  | 50.5914 | -2.2247
11 | YES | 347 | 1260.5 | 0.19854 | -0.10423 | 0.36186 | 0.46609 |  | 25.2013 | -2.2766
11 | YES | 348 | 1002.15 | 0.53413 | -0.096628 | -0.091619 | 0.0050093 |  | 9.7596 | -2.2626
11 | YES | 349 | 798.81 | 0.81482 | -0.19291 | -0.44854 | -0.25563 |  | 5.6067 | -2.2524
11 | YES | 350 | 593.42 | 0.70698 | -0.39444 | -0.80906 | -0.41462 |  | 1496.83 | -0.2328
11 | YES | 351 | 399.468 | 0.43363 | -0.78195 | -1.1495 | -0.36755 |  | 1247.79 | -0.10179
11 | YES | 352 | 297.774 | 0.29935 | -1.3573 | -1.328 | 0.029279 |  | 996.482 | -0.058372
11 | YES | 353 | 249.757 | 0.42768 | -1.5525 | -1.4123 | 0.14023 |  | 806.108 | -0.11096
11 | YES | 354 | 201.071 | 0.14409 | -1.8786 | -1.4977 | 0.38083 |  | 604.872 | -0.21788
11 | YES | 355 | 175.289 | 0.3426 | -2.0525 | -1.543 | 0.5095 |  | 407.077 | -0.64533
11 | YES | 356 | 149.378 | 0.2337 | -2.1975 | -1.5885 | 0.60904 |  | 303.923 | -1.0978
11 | YES | 357 | 125.288 | 0.36369 | -2.1823 | -1.6308 | 0.55155 |  | 253.186 | -1.2942
11 | YES | 358 | 98.579 | 0.59052 | -2.3017 | -1.6776 | 0.62407 |  | 201.27 | -1.5737
11 | YES | 359 | 75.6397 | 0.31522 | -2.1963 | -1.7179 | 0.4784 |  | 173.963 | -1.4983
11 | YES | 360 | 49.8857 | 0.14984 | -2.2154 | -1.7631 | 0.45229 |  | 151.536 | -1.5191
11 | YES | 361 | 24.7805 | 0.18714 | -2.2261 | -1.8072 | 0.41889 |  | 126.586 | -1.9213
11 | YES | 362 | 10.087 | 0.11165 | -2.197 | -1.833 | 0.36401 |  | 100.402 | -2.2354
11 | NO | 363 | 4.5509 | 0.27574 | -2.2028 | -1.8427 | 0.36014 |  | 74.0483 | -2.1936
12 | NO | 366 | 4296.11 | 0.85002 | -1.2765 | 5.6902 | 6.9667 |  | 49.93 | -2.2862
12 | NO | 367 | 3998.81 | 0.82125 | -1.148 | 5.1684 | 6.3164 |  | 24.5799 | -2.6577
12 | NO | 368 | 3500.42 | 0.28291 | -1.079 | 4.2936 | 5.3726 |  | 11.1119 | -2.1795
12 | NO | 369 | 3001.6 | 0.44799 | -0.92051 | 3.418 | 4.3385 |  | 1251.43 | -0.13093
12 | NO | 370 | 2501.5 | 0.84328 | -0.68736 | 2.5402 | 3.2275 |  | 1001.33 | -0.10345
12 | NO | 371 | 2001.51 | 0.47385 | -0.44562 | 1.6625 | 2.1082 |  | 803.855 | -0.12119
12 | NO | 372 | 1500.79 | 0.67265 | -0.16092 | 0.78364 | 0.94455 |  | 601.676 | -0.20373
12 | YES | 373 | 1250.28 | 0.12335 | -0.10074 | 0.34392 | 0.44466 |  | 401.528 | -0.66156
12 | YES | 374 | 1000.49 | 0.12977 | -0.098033 | -0.094533 | 0.0035005 |  | 299.696 | -0.87577
12 | YES | 375 | 801.742 | 0.52064 | -0.17107 | -0.44339 | -0.27233 |  | 251.149 | -1.1463
12 | YES | 376 | 600.251 | 0.23163 | -0.34906 | -0.79707 | -0.44801 |  | 200.618 | -1.4824
12 | YES | 377 | 398.905 | 0.58172 | -0.68064 | -1.1505 | -0.46985 |  | 175.661 | -1.6392
12 | YES | 378 | 301.608 | 0.37111 | -1.0426 | -1.3213 | -0.27864 |  | 150.244 | -1.7454
12 | YES | 379 | 251.725 | 0.50634 | -1.4142 | -1.4088 | 0.0053797 |  | 125.428 | -1.7734
12 | YES | 380 | 200.658 | 0.83374 | -1.4913 | -1.4985 | -0.0071675 |  | 100.394 | -2.2294
12 | YES | 381 | 175.629 | 0.2906 | -1.7719 | -1.5424 | 0.22953 |  | 74.9905 | -2.2349
12 | YES | 382 | 148.758 | 0.44328 | -1.8517 | -1.5896 | 0.26215 |  | 50.6495 | -2.269
12 | YES | 383 | 125.149 | 0.78861 | -2.063 | -1.631 | 0.43196 |  | 25.1855 | -2.2529
12 | YES | 384 | 100.411 | 0.39308 | -2.1979 | -1.6744 | 0.52349 |  | 9.6049 | -2.2362
12 | YES | 385 | 77.0474 | 0.48443 | -2.2531 | -1.7154 | 0.53768 |  | 5.2854 | -2.3651
12 | YES | 386 | 50.7156 | 0.28242 | -2.2416 | -1.7617 | 0.47992 |  | 1251.23 | -0.1746
12 | YES | 387 | 25.2391 | 0.17237 | -2.284 | -1.8064 | 0.47766 |  | 998.646 | -0.075008
12 | YES | 388 | 10.1999 | 0.17594 | -2.261 | -1.8328 | 0.42822 |  | 802.636 | -0.11404
12 | NO | 389 | 4.4886 | 0.24641 | -2.2587 | -1.8428 | 0.41594 |  | 599.335 | -0.18919
13 | NO | 392 | 4201.11 | 0.2505 | -1.1478 | 5.5235 | 6.6713 |  | 399.909 | -0.56999
13 | NO | 393 | 3998.58 | 0.26882 | -1.1126 | 5.168 | 6.2805 |  | 300.394 | -0.87488
13 | NO | 394 | 3498.99 | 0.63928 | -1.0651 | 4.2911 | 5.3562 |  | 249.659 | -1.1046
13 | NO | 395 | 3000.35 | 0.12517 | -0.90141 | 3.4158 | 4.3172 |  | 198.585 | -1.3585
13 | NO | 396 | 2500.04 | 0.38573 | -0.71727 | 2.5376 | 3.2549 |  | 174.145 | -1.456
13 | NO | 397 | 2000.94 | 0.31785 | -0.46879 | 1.6615 | 2.1303 |  | 151.187 | -1.6382
13 | YES | 398 | 1499.71 | 0.35183 | -0.15896 | 0.78174 | 0.9407 |  | 125.718 | -1.6541
13 | YES | 399 | 1249.27 | 0.21818 | -0.088537 | 0.34215 | 0.43068 |  | 98.6873 | -2.0075
13 | YES | 400 | 1000.34 | 0.21223 | -0.070752 | -0.094796 | -0.024044 |  | 73.0154 | -2.2525
13 | YES | 401 | 801.259 | 0.8392 | -0.25675 | -0.44424 | -0.18749 |  | 50.4418 | -2.2031
13 | YES | 402 | 598.737 | 0.83297 | -0.30009 | -0.79972 | -0.49963 |  | 24.57 | -2.2173
13 | YES | 403 | 399.773 | 0.37799 | -0.94149 | -1.149 | -0.20748 |  | 11.6758 | -2.2171
13 | YES | 404 | 301.302 | 0.28849 | -1.2091 | -1.3218 | -0.11269 |  | 6.0236 | -2.2572
13 | YES | 405 | 250.568 | 0.27467 | -1.4097 | -1.4109 | -0.0011611 |  | 1249.63 | -0.13361
13 | YES | 406 | 199.803 | 0.4091 | -1.515 | -1.5 | 0.015052 |  | 999.412 | -0.1299
13 | YES | 407 | 175.51 | 0.3087 | -1.7168 | -1.5426 | 0.17421 |  | 799.052 | -0.093562
13 | YES | 408 | 149.628 | 0.32269 | -1.8006 | -1.588 | 0.21252 |  | 598.112 | -0.13651
13 | YES | 409 | 125.132 | 0.26626 | -2.0368 | -1.631 | 0.40574 |  | 399.795 | -0.46688
13 | YES | 410 | 99.9763 | 0.35596 | -2.2717 | -1.6752 | 0.59653 |  | 297.419 | -0.74576
13 | YES | 411 | 74.4035 | 0.088755 | -2.2861 | -1.7201 | 0.56602 |  | 253.01 | -0.9932
13 | YES | 412 | 50.5914 | 0.38225 | -2.2247 | -1.7619 | 0.46277 |  | 204.763 | -1.2601
13 | YES | 413 | 25.2013 | 0.1047 | -2.2766 | -1.8064 | 0.47013 |  | 174.843 | -1.1856
13 | YES | 414 | 9.7596 | 0.2665 | -2.2626 | -1.8335 | 0.42905 |  | 151.252 | -1.1637
13 | YES | 415 | 5.6067 | 0.28208 | -2.2524 | -1.8408 | 0.4116 |  | 127.108 | -1.348
14 | NO | 418 | 3342.31 | 0.35376 | -1.0436 | 4.016 | 5.0596 |  | 99.6608 | -1.5634
14 | NO | 419 | 3001.04 | 0.49995 | -0.84295 | 3.417 | 4.26 |  | 76.8172 | -1.9608
14 | NO | 420 | 2497.99 | 0.84452 | -0.69182 | 2.534 | 3.2258 |  | 49.8571 | -2.1901
14 | NO | 421 | 1995.74 | 0.36664 | -0.47956 | 1.6524 | 2.132 |  | 24.4847 | -2.225
14 | YES | 422 | 1496.83 | 0.094953 | -0.2328 | 0.77669 | 1.0095 |  | 11.8118 | -2.1828
14 | YES | 423 | 1247.79 | 0.12816 | -0.10179 | 0.33955 | 0.44134 |  | 6.2732 | -2.1911
14 | YES | 424 | 996.482 | 0.15313 | -0.058372 | -0.10157 | -0.043196 |  | 1251 | -0.12394
14 | YES | 425 | 806.108 | 0.41673 | -0.11096 | -0.43573 | -0.32477 |  | 1001.76 | -0.058984
14 | YES | 426 | 604.872 | 0.31697 | -0.21788 | -0.78896 | -0.57107 |  | 800.357 | -0.11489
14 | YES | 427 | 407.077 | 0.34465 | -0.64533 | -1.1361 | -0.49081 |  | 599.476 | -0.30103
14 | YES | 428 | 303.923 | 0.13549 | -1.0978 | -1.3172 | -0.21939 |  | 403.584 | -0.61578
14 | YES | 429 | 253.186 | 0.20356 | -1.2942 | -1.4063 | -0.11205 |  | 300.032 | -0.65826
14 | YES | 430 | 201.27 | 0.51275 | -1.5737 | -1.4974 | 0.076337 |  | 248.451 | -0.68814
14 | YES | 431 | 173.963 | 0.49482 | -1.4983 | -1.5453 | -0.047065 |  | 199.203 | -0.80225
14 | YES | 432 | 151.536 | 0.14839 | -1.5191 | -1.5847 | -0.065571 |  | 175.318 | -0.80311
14 | YES | 433 | 126.586 | 0.29554 | -1.9213 | -1.6285 | 0.29283 |  | 151.213 | -0.8403
14 | YES | 434 | 100.402 | 0.29562 | -2.2354 | -1.6744 | 0.56095 |  | 124.946 | -0.90175
14 | YES | 435 | 74.0483 | 0.22137 | -2.1936 | -1.7207 | 0.47287 |  | 100.14 | -1.0214
14 | YES | 436 | 49.93 | 0.39272 | -2.2862 | -1.763 | 0.5232 |  | 75.0758 | -1.2107
14 | YES | 437 | 24.5799 | 0.27341 | -2.6577 | -1.8075 | 0.8502 |  | 49.3962 | -1.4535
14 | YES | 438 | 11.1119 | 0.17292 | -2.1795 | -1.8312 | 0.34835 |  | 25.0857 | -2.2462
14 | NO | 439 | 3.6794 | 0.1964 | -2.2125 | -1.8442 | 0.36831 |  | 9.8932 | -2.3363
15 | NO | 441 | 1999.55 | 0.22737 | -1.2821 | 1.6591 | 2.9412 |  | 5.4559 | -2.2833
15 | NO | 442 | 1500.86 | 0.45987 | -0.21554 | 0.78376 | 0.9993
15 | YES | 443 | 1251.43 | 0.82942 | -0.13093 | 0.34594 | 0.47687
15 | YES | 444 | 1001.33 | 0.15632 | -0.10345 | -0.093058 | 0.010388
15 | YES | 445 | 803.855 | 0.11578 | -0.12119 | -0.43968 | -0.31849
15 | YES | 446 | 601.676 | 0.59367 | -0.20373 | -0.79457 | -0.59084
15 | YES | 447 | 401.528 | 0.58964 | -0.66156 | -1.1459 | -0.48432
15 | YES | 448 | 299.696 | 0.31095 | -0.87577 | -1.3246 | -0.44886
15 | YES | 449 | 251.149 | 0.12937 | -1.1463 | -1.4098 | -0.26349
15 | YES | 450 | 200.618 | 1.0845 | -1.4824 | -1.4985 | -0.016188
15 | YES | 451 | 175.661 | 0.14709 | -1.6392 | -1.5423 | 0.096826
15 | YES | 452 | 150.244 | 0.082279 | -1.7454 | -1.587 | 0.15843
15 | YES | 453 | 125.428 | 0.21249 | -1.7734 | -1.6305 | 0.14288
15 | YES | 454 | 100.394 | 0.52829 | -2.2294 | -1.6745 | 0.5549
15 | YES | 455 | 74.9905 | 0.39506 | -2.2349 | -1.7191 | 0.51583
15 | YES | 456 | 50.6495 | 0.22069 | -2.269 | -1.7618 | 0.50718
15 | YES | 457 | 25.1855 | 0.23854 | -2.2529 | -1.8065 | 0.44641
15 | YES | 458 | 9.6049 | 0.19968 | -2.2362 | -1.8338 | 0.40237
15 | YES | 459 | 5.2854 | 0.077476 | -2.3651 | -1.8414 | 0.52366
16 | NO | 461 | 2572.39 | 0.48374 | -0.74349 | 2.6646 | 3.4081
16 | NO | 462 | 2000.27 | 0.69889 | -0.52011 | 1.6604 | 2.1805
16 | NO | 463 | 1501.18 | 0.3593 | -0.31237 | 0.78432 | 1.0967
16 | YES | 464 | 1251.23 | 0.5052 | -0.1746 | 0.34559 | 0.52019
16 | YES | 465 | 998.646 | 0.23251 | -0.075008 | -0.097769 | -0.022761
16 | NO | 0 | 0 | 0 | -1.383 | -1.8507 | -0.46768
16 | YES | 466 | 802.636 | 0.22548 | -0.11404 | -0.44182 | -0.32778
16 | YES | 467 | 599.335 | 0.87627 | -0.18919 | -0.79867 | -0.60948
16 | YES | 468 | 399.909 | 0.12728 | -0.56999 | -1.1487 | -0.57873
16 | YES | 469 | 300.394 | 0.44982 | -0.87488 | -1.3234 | -0.44853
16 | YES | 470 | 249.659 | 0.06262 | -1.1046 | -1.4125 | -0.30781
16 | YES | 471 | 198.585 | 0.54379 | -1.3585 | -1.5021 | -0.1436
16 | YES | 472 | 174.145 | 0.18582 | -1.456 | -1.545 | -0.089035
16 | YES | 473 | 151.187 | 0.86085 | -1.6382 | -1.5853 | 0.052887
16 | YES | 474 | 125.718 | 0.69051 | -1.6541 | -1.63 | 0.024081
16 | YES | 475 | 98.6873 | 0.55424 | -2.0075 | -1.6775 | 0.33005
16 | YES | 476 | 73.0154 | 0.16895 | -2.2525 | -1.7225 | 0.52998
16 | YES | 477 | 50.4418 | 0.18342 | -2.2031 | -1.7621 | 0.44095
16 | YES | 478 | 24.57 | 0.30696 | -2.2173 | -1.8076 | 0.40977
16 | YES | 479 | 11.6758 | 0.17642 | -2.2171 | -1.8302 | 0.38691
16 | YES | 480 | 6.0236 | 0.19825 | -2.2572 | -1.8401 | 0.41706
17 | NO | 483 | 2633.92 | 0.2719 | -0.7003 | 2.7726 | 3.4729
17 | NO | 484 | 2500.54 | 0.23963 | -0.67292 | 2.5385 | 3.2114
17 | NO | 485 | 2002.19 | 0.21459 | -0.5019 | 1.6637 | 2.1656
17 | NO | 486 | 1500.99 | 0.097577 | -0.24887 | 0.78399 | 1.0329
17 | YES | 487 | 1249.63 | 0.21728 | -0.13361 | 0.34278 | 0.47639
17 | YES | 488 | 999.412 | 0.2711 | -0.1299 | -0.096425 | 0.033476
17 | YES | 489 | 799.052 | 0.38068 | -0.093562 | -0.44811 | -0.35455
17 | YES | 490 | 598.112 | 0.26103 | -0.13651 | -0.80082 | -0.66431
17 | YES | 491 | 399.795 | 0.37405 | -0.46688 | -1.1489 | -0.68205
17 | YES | 492 | 297.419 | 0.17995 | -0.74576 | -1.3286 | -0.58286
17 | YES | 493 | 253.01 | 0.25307 | -0.9932 | -1.4066 | -0.41338
17 | YES | 494 | 204.763 | 0.060919 | -1.2601 | -1.4913 | -0.23117
17 | YES | 495 | 174.843 | 0.61067 | -1.1856 | -1.5438 | -0.35814
17 | YES | 496 | 151.252 | 0.54264 | -1.1637 | -1.5852 | -0.4215
17 | YES | 497 | 127.108 | 0.3692 | -1.348 | -1.6276 | -0.27955
17 | YES | 498 | 99.6608 | 0.29939 | -1.5634 | -1.6757 | -0.11237
17 | YES | 499 | 76.8172 | 0.26404 | -1.9608 | -1.7158 | 0.245
17 | YES | 500 | 49.8571 | 0.44803 | -2.1901 | -1.7632 | 0.42692
17 | YES | 501 | 24.4847 | 0.15707 | -2.225 | -1.8077 | 0.4173
17 | YES | 502 | 11.8118 | 0.35219 | -2.1828 | -1.8299 | 0.35286
17 | YES | 503 | 6.2732 | 0.15298 | -2.1911 | -1.8397 | 0.35138
18 | NO | 506 | 1756.28 | 0.58251 | -0.58406 | 1.2321 | 1.8162
18 | NO | 507 | 1501.19 | 0.21378 | -0.42645 | 0.78434 | 1.2108
18 | YES | 508 | 1251 | 0.5765 | -0.12394 | 0.34518 | 0.46913
18 | YES | 509 | 1001.76 | 0.37376 | -0.058984 | -0.092303 | -0.03332
18 | YES | 510 | 800.357 | 0.26156 | -0.11489 | -0.44582 | -0.33093
18 | YES | 511 | 599.476 | 0.47336 | -0.30103 | -0.79843 | -0.49739
18 | YES | 512 | 403.584 | 0.092177 | -0.61578 | -1.1423 | -0.5265
18 | YES | 513 | 300.032 | 0.15906 | -0.65826 | -1.324 | -0.66577
18 | YES | 514 | 248.451 | 0.52123 | -0.68814 | -1.4146 | -0.72644
18 | YES | 515 | 199.203 | 0.38477 | -0.80225 | -1.501 | -0.69877
18 | YES | 516 | 175.318 | 0.31151 | -0.80311 | -1.5429 | -0.73984
18 | YES | 517 | 151.213 | 0.067558 | -0.8403 | -1.5853 | -0.74496
18 | YES | 518 | 124.946 | 0.52959 | -0.90175 | -1.6314 | -0.72961
18 | YES | 519 | 100.14 | 0.53719 | -1.0214 | -1.6749 | -0.65355
18 | YES | 520 | 75.0758 | 0.22863 | -1.2107 | -1.7189 | -0.50821
18 | YES | 521 | 49.3962 | 0.14441 | -1.4535 | -1.764 | -0.31047
18 | YES | 522 | 25.0857 | 0.10492 | -2.2462 | -1.8066 | 0.43956
18 | YES | 523 | 9.8932 | 0.25334 | -2.3363 | -1.8333 | 0.503
18 | YES | 524 | 5.4559 | 0.1619 | -2.2833 | -1.8411 | 0.44222
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. Pressure
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0010115 |  | X | CALC Y
m= | -0.0053468 | 0.008844 |  | Std Dev of Resid |  | 0.79356 |  | 1 | -1.1176
b= | -1.1122 | 0.098308 |  | Resid Sum of Squares |  | 227.3359 |  | 18 | -1.2084
# of Values Extracted = |  |  | 367
# of Values Used in Fit = |  |  | 363
File | Use | Sample # | CTD DOX | CTD DOX | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 4 | 0.048 | 0 | 0.004 | -1.1176 | -1.1216 |  | 1 | -0.0012562
1 | NO | 5 | 0.04676 | 0.0014834 | -0.05324 | -1.1176 | -1.0643 |  | 1 | -1.0448
1 | YES | 6 | 0.50074 | 0.0087174 | -0.0012562 | -1.1176 | -1.1163 |  | 1 | -1.1088
1 | YES | 7 | 2.8522 | 0.0043797 | -1.0448 | -1.1176 | -0.072762 |  | 1 | -1.264
1 | YES | 8 | 3.2522 | 0.002912 | -1.1088 | -1.1176 | -0.0087815 |  | 1 | -1.411
1 | YES | 9 | 3.687 | 0.0021289 | -1.264 | -1.1176 | 0.14648 |  | 1 | -1.3767
1 | YES | 10 | 3.877 | 0.013065 | -1.411 | -1.1176 | 0.29347 |  | 2 | -1.3573
1 | YES | 11 | 4.2193 | 0.0032873 | -1.3767 | -1.1176 | 0.25912 |  | 2 | -1.4682
2 | YES | 13 | 3.4827 | 0.0042317 | -1.3573 | -1.1229 | 0.23439 |  | 2 | -1.4437
2 | YES | 14 | 4.2358 | 0.0015184 | -1.4682 | -1.1229 | 0.34527 |  | 2 | -1.4255
2 | YES | 15 | 4.3923 | 0.0015558 | -1.4437 | -1.1229 | 0.32078 |  | 2 | -1.4218
2 | YES | 16 | 4.3895 | 0.0016078 | -1.4255 | -1.1229 | 0.3026 |  | 2 | -1.4634
2 | YES | 17 | 4.3972 | 0.0035285 | -1.4218 | -1.1229 | 0.29886 |  | 2 | -1.4856
2 | YES | 18 | 4.5156 | 0.0024659 | -1.4634 | -1.1229 | 0.34053 |  | 2 | -1.5192
2 | YES | 19 | 4.6644 | 0.012831 | -1.4856 | -1.1229 | 0.36272 |  | 3 | -0.1977
2 | YES | 20 | 4.6928 | 0.0017865 | -1.5192 | -1.1229 | 0.39632 |  | 3 | -0.099
3 | YES | 22 | 0.3693 | 0.0017447 | -0.1977 | -1.1282 | -0.93054 |  | 3 | -0.12207
3 | YES | 23 | 0.21 | 0 | -0.099 | -1.1282 | -1.0292 |  | 3 | -0.17247
3 | YES | 24 | 0.20993 | 0.00057973 | -0.12207 | -1.1282 | -1.0062 |  | 3 | -0.2097
3 | YES | 25 | 0.16153 | 0.0013606 | -0.17247 | -1.1282 | -0.95577 |  | 3 | -0.2234
3 | YES | 26 | 0.1613 | 0.0010539 | -0.2097 | -1.1282 | -0.91854 |  | 3 | -0.40936
3 | YES | 27 | 0.4346 | 0.0012006 | -0.2234 | -1.1282 | -0.90485 |  | 3 | -0.70209
3 | YES | 28 | 0.92265 | 0.0023447 | -0.40936 | -1.1282 | -0.71889 |  | 3 | -0.7866
3 | YES | 29 | 1.3029 | 0.0011328 | -0.70209 | -1.1282 | -0.42615 |  | 3 | -0.7414
3 | YES | 31 | 1.9264 | 0.0010691 | -0.7866 | -1.1282 | -0.34165 |  | 3 | -0.80164
3 | YES | 32 | 1.8416 | 0.0012006 | -0.7414 | -1.1282 | -0.38685 |  | 3 | -0.94776
3 | YES | 33 | 2.0064 | 0.0013354 | -0.80164 | -1.1282 | -0.32661 |  | 3 | -1.3317
3 | YES | 34 | 2.3332 | 0.0016834 | -0.94776 | -1.1282 | -0.18049 |  | 3 | -1.5895
3 | YES | 35 | 3.2813 | 0.0027519 | -1.3317 | -1.1282 | 0.20342 |  | 3 | -1.7234
3 | YES | 36 | 4.0995 | 0.0030114 | -1.5895 | -1.1282 | 0.46128 |  | 3 | -1.7332
3 | YES | 37 | 4.4806 | 0.0012583 | -1.7234 | -1.1282 | 0.59511 |  | 3 | -1.7302
3 | YES | 38 | 4.4338 | 0.0029481 | -1.7332 | -1.1282 | 0.60497 |  | 3 | -1.7118
3 | YES | 39 | 4.5799 | 0.0018105 | -1.7302 | -1.1282 | 0.6019 |  | 4 | -0.64814
3 | YES | 40 | 4.5822 | 0.001127 | -1.7118 | -1.1282 | 0.58354 |  | 4 | -0.53817
4 | YES | 58 | 1.3029 | 0.0022558 | -0.64814 | -1.1336 | -0.48545 |  | 4 | -0.218
4 | YES | 59 | 0.95083 | 0.0012803 | -0.53817 | -1.1336 | -0.59543 |  | 4 | -0.1416
4 | YES | 60 | 0.478 | 0 | -0.218 | -1.1336 | -0.91559 |  | 4 | -0.11782
4 | YES | 61 | 0.2704 | 0.0012005 | -0.1416 | -1.1336 | -0.99199 |  | 4 | -0.10126
4 | YES | 62 | 0.18318 | 0.0023805 | -0.11782 | -1.1336 | -1.0158 |  | 4 | -0.24649
4 | YES | 63 | 0.17374 | 0.00099807 | -0.10126 | -1.1336 | -1.0323 |  | 4 | -0.26674
4 | YES | 64 | 0.33051 | 0.001674 | -0.24649 | -1.1336 | -0.8871 |  | 4 | -0.54804
4 | YES | 65 | 0.33226 | 0.0014866 | -0.26674 | -1.1336 | -0.86685 |  | 4 | -0.86334
4 | YES | 66 | 0.99696 | 0.0016852 | -0.54804 | -1.1336 | -0.58555 |  | 4 | -1.0023
4 | YES | 67 | 1.6957 | 0.002548 | -0.86334 | -1.1336 | -0.27025 |  | 4 | -1.2267
4 | YES | 68 | 2.0467 | 0.0023078 | -1.0023 | -1.1336 | -0.13125 |  | 4 | -1.3421
4 | YES | 69 | 2.5483 | 0.0018423 | -1.2267 | -1.1336 | 0.093108 |  | 4 | -1.4324
4 | YES | 70 | 2.8469 | 0.0011569 | -1.3421 | -1.1336 | 0.20855 |  | 4 | -3.2447
4 | YES | 71 | 3.0726 | 0.0018841 | -1.4324 | -1.1336 | 0.29877 |  | 4 | -1.4465
4 | YES | 72 | 3.1943 | 0.0016615 | -3.2447 | -1.1336 | 2.1111 |  | 4 | -2.0579
4 | YES | 73 | 3.1975 | 0.0013041 | -1.4465 | -1.1336 | 0.31287 |  | 4 | -2.6084
4 | YES | 74 | 4.4961 | 0.0029635 | -2.0579 | -1.1336 | 0.92428 |  | 4 | -2.0307
4 | YES | 75 | 4.5756 | 0.0016338 | -2.6084 | -1.1336 | 1.4748 |  | 4 | -2.0381
4 | YES | 76 | 4.5983 | 0.0018676 | -2.0307 | -1.1336 | 0.89716 |  | 4 | -1.9826
4 | YES | 77 | 4.5999 | 0.0019402 | -2.0381 | -1.1336 | 0.90446 |  | 5 | -0.95205
4 | YES | 78 | 4.6054 | 0.0016601 | -1.9826 | -1.1336 | 0.84902 |  | 5 | -0.8287
5 | YES | 82 | 1.5379 | 0.0019487 | -0.95205 | -1.1389 | -0.18689 |  | 5 | -0.7406
5 | YES | 83 | 1.5143 | 0.0013079 | -0.8287 | -1.1389 | -0.31024 |  | 5 | -0.50653
5 | YES | 84 | 1.2764 | 0.0014111 | -0.7406 | -1.1389 | -0.39834 |  | 5 | -0.31036
5 | YES | 85 | 0.89347 | 0.0028754 | -0.50653 | -1.1389 | -0.63241 |  | 5 | -0.18648
5 | YES | 86 | 0.43764 | 0.0024118 | -0.31036 | -1.1389 | -0.82857 |  | 5 | -0.1434
5 | YES | 87 | 0.25452 | 0.0027931 | -0.18648 | -1.1389 | -0.95246 |  | 5 | -0.13086
5 | YES | 88 | 0.1806 | 0.0012006 | -0.1434 | -1.1389 | -0.99554 |  | 5 | -0.44645
5 | YES | 89 | 0.18014 | 0.0016498 | -0.13086 | -1.1389 | -1.0081 |  | 5 | -0.19421
5 | YES | 90 | 0.18255 | 0.0014434 | -0.44645 | -1.1389 | -0.69248 |  | 5 | -0.41079
5 | YES | 91 | 0.28279 | 0.0016017 | -0.19421 | -1.1389 | -0.94473 |  | 5 | -0.58407
5 | YES | 92 | 0.74521 | 0.0019745 | -0.41079 | -1.1389 | -0.72815 |  | 5 | -0.78649
5 | YES | 93 | 1.0969 | 0.0017547 | -0.58407 | -1.1389 | -0.55487 |  | 5 | -0.95894
5 | YES | 94 | 1.3515 | 0.0014383 | -0.78649 | -1.1389 | -0.35245 |  | 5 | -0.87647
5 | YES | 95 | 1.8271 | 0.0020829 | -0.95894 | -1.1389 | -0.18 |  | 5 | -1.2899
5 | YES | 96 | 1.9475 | 0.0013421 | -0.87647 | -1.1389 | -0.26247 |  | 5 | -1.4495
5 | YES | 97 | 2.4851 | 0.0016243 | -1.2899 | -1.1389 | 0.151 |  | 5 | -1.946
5 | YES | 98 | 2.9115 | 0.014906 | -1.4495 | -1.1389 | 0.31057 |  | 5 | -2.1335
5 | YES | 99 | 3.944 | 0.0056295 | -1.946 | -1.1389 | 0.80704 |  | 5 | -2.1058
5 | YES | 100 | 4.4955 | 0.0013044 | -2.1335 | -1.1389 | 0.99458 |  | 5 | -2.0639
5 | YES | 101 | 4.5092 | 0.0018394 | -2.1058 | -1.1389 | 0.96688 |  | 5 | -2.0937
5 | YES | 102 | 4.5181 | 0.0014259 | -2.0639 | -1.1389 | 0.92497 |  | 5 | -2.1109
5 | YES | 103 | 4.5223 | 0.0015575 | -2.0937 | -1.1389 | 0.95472 |  | 6 | -1.0939
5 | YES | 104 | 4.524 | 0.0012706 | -2.1109 | -1.1389 | 0.97201 |  | 6 | -0.95145
6 | YES | 122 | 1.8111 | 0.00089619 | -1.0939 | -1.1443 | -0.050406 |  | 6 | -0.7953
6 | YES | 123 | 1.7655 | 0.0015809 | -0.95145 | -1.1443 | -0.19284 |  | 6 | -0.6011
6 | YES | 124 | 1.5257 | 0.0012493 | -0.7953 | -1.1443 | -0.34899 |  | 6 | -0.4314
6 | YES | 125 | 1.2279 | 0.00076817 | -0.6011 | -1.1443 | -0.54319 |  | 6 | -0.17702
6 | YES | 126 | 0.8526 | 0.0026692 | -0.4314 | -1.1443 | -0.71289 |  | 6 | -0.1298
6 | YES | 127 | 0.36298 | 0.0024562 | -0.17702 | -1.1443 | -0.96727 |  | 6 | -0.090719
6 | YES | 128 | 0.2222 | 0.0021433 | -0.1298 | -1.1443 | -1.0145 |  | 6 | -0.07
6 | YES | 129 | 0.22228 | 0.0021917 | -0.090719 | -1.1443 | -1.0536 |  | 6 | -0.080546
6 | YES | 130 | 0.186 | 0 | -0.07 | -1.1443 | -1.0743 |  | 6 | -0.23669
6 | YES | 131 | 0.22246 | 0.0019322 | -0.080546 | -1.1443 | -1.0637 |  | 6 | -0.57549
6 | YES | 132 | 0.47631 | 0.0014436 | -0.23669 | -1.1443 | -0.9076 |  | 6 | -0.84797
6 | YES | 133 | 1.1175 | 0.0017084 | -0.57549 | -1.1443 | -0.5688 |  | 6 | -0.97879
6 | YES | 134 | 1.569 | 0.0039453 | -0.84797 | -1.1443 | -0.29632 |  | 6 | -1.3153
6 | YES | 135 | 1.9172 | 0.0018221 | -0.97879 | -1.1443 | -0.1655 |  | 6 | -1.4349
6 | YES | 136 | 2.5827 | 0.0021456 | -1.3153 | -1.1443 | 0.17103 |  | 6 | -1.7284
6 | YES | 137 | 3.1111 | 0.0014921 | -1.4349 | -1.1443 | 0.29063 |  | 6 | -1.8489
6 | YES | 138 | 3.4436 | 0.0019536 | -1.7284 | -1.1443 | 0.58407 |  | 6 | -2.108
6 | YES | 139 | 3.6641 | 0.001536 | -1.8489 | -1.1443 | 0.70459 |  | 6 | -2.2185
6 | YES | 140 | 4.197 | 0.0031303 | -2.108 | -1.1443 | 0.96374 |  | 6 | -2.2005
6 | YES | 141 | 4.5605 | 0.0015763 | -2.2185 | -1.1443 | 1.0742 |  | 6 | -2.1808
6 | YES | 142 | 4.5715 | 0.0010493 | -2.2005 | -1.1443 | 1.0562 |  | 6 | -2.1774
6 | YES | 143 | 4.5812 | 0.0013225 | -2.1808 | -1.1443 | 1.0365 |  | 6 | -2.1664
6 | YES | 144 | 4.5865 | 0.0018528 | -2.1774 | -1.1443 | 1.0332 |  | 7 | -1.1079
6 | YES | 145 | 4.5886 | 0.0013785 | -2.1664 | -1.1443 | 1.0221 |  | 7 | -1.0223
7 | YES | 186 | 1.9311 | 0.0012286 | -1.1079 | -1.1496 | -0.041752 |  | 7 | -0.86338
7 | YES | 187 | 1.8347 | 0.002743 | -1.0223 | -1.1496 | -0.12734 |  | 7 | -0.85335
7 | YES | 188 | 1.5906 | 0.0026761 | -0.86338 | -1.1496 | -0.28625 |  | 7 | -0.68212
7 | YES | 189 | 1.5897 | 0.002941 | -0.85335 | -1.1496 | -0.29628 |  | 7 | -0.42386
7 | YES | 190 | 1.2279 | 0.00078071 | -0.68212 | -1.1496 | -0.46751 |  | 7 | -0.16389
7 | YES | 191 | 0.79214 | 0.0074926 | -0.42386 | -1.1496 | -0.72577 |  | 7 | -0.084479
7 | YES | 192 | 0.34511 | 0.0023587 | -0.16389 | -1.1496 | -0.98574 |  | 7 | -0.068397
7 | YES | 193 | 0.22352 | 0.0025103 | -0.084479 | -1.1496 | -1.0652 |  | 7 | -0.10731
7 | YES | 194 | 0.1986 | 0.0012006 | -0.068397 | -1.1496 | -1.0812 |  | 7 | -0.14869
7 | YES | 195 | 0.22269 | 0.0019834 | -0.10731 | -1.1496 | -1.0423 |  | 7 | -0.49131
7 | YES | 196 | 0.32331 | 0.0021089 | -0.14869 | -1.1496 | -1.0009 |  | 7 | -0.67841
7 | YES | 197 | 0.80869 | 0.013204 | -0.49131 | -1.1496 | -0.65833 |  | 7 | -0.91867
7 | YES | 198 | 1.2606 | 0.0018827 | -0.67841 | -1.1496 | -0.47122 |  | 7 | -1.2879
7 | YES | 199 | 1.6433 | 0.0019382 | -0.91867 | -1.1496 | -0.23096 |  | 7 | -1.211
7 | YES | 200 | 2.3511 | 0.0008104 | -1.2879 | -1.1496 | 0.13827 |  | 7 | -1.521
7 | YES | 201 | 2.604 | 0.0016973 | -1.211 | -1.1496 | 0.061418 |  | 7 | -1.7093
7 | YES | 202 | 2.871 | 0.0032544 | -1.521 | -1.1496 | 0.37135 |  | 7 | -2.2897
7 | YES | 203 | 3.3187 | 0.0023088 | -1.7093 | -1.1496 | 0.55968 |  | 7 | -2.2709
7 | YES | 204 | 4.4733 | 0.0024342 | -2.2897 | -1.1496 | 1.1401 |  | 7 | -2.2669
7 | YES | 205 | 4.5131 | 0.001646 | -2.2709 | -1.1496 | 1.1213 |  | 7 | -2.2247
7 | YES | 206 | 4.5501 | 0.0016361 | -2.2669 | -1.1496 | 1.1173 |  | 7 | -2.2487
7 | YES | 207 | 4.5613 | 0.001885 | -2.2247 | -1.1496 | 1.075 |  | 7 | -2.2656
7 | YES | 208 | 4.5683 | 0.0016469 | -2.2487 | -1.1496 | 1.099 |  | 8 | -1.2647
7 | YES | 209 | 4.5684 | 0.0017776 | -2.2656 | -1.1496 | 1.116 |  | 8 | -1.1266
8 | YES | 258 | 2.0893 | 0.0030343 | -1.2647 | -1.155 | 0.10973 |  | 8 | -1.0315
8 | YES | 259 | 2.0264 | 0.0015592 | -1.1266 | -1.155 | -0.028339 |  | 8 | -0.89612
8 | YES | 260 | 1.8585 | 0.0029751 | -1.0315 | -1.155 | -0.12351 |  | 8 | -0.67139
8 | YES | 261 | 1.6089 | 0.0018999 | -0.89612 | -1.155 | -0.25886 |  | 8 | -0.4547
8 | YES | 262 | 1.2276 | 0.0013806 | -0.67139 | -1.155 | -0.48359 |  | 8 | -0.19967
8 | YES | 263 | 0.8383 | 0.0023897 | -0.4547 | -1.155 | -0.70028 |  | 8 | -0.1488
8 | YES | 264 | 0.42333 | 0.001106 | -0.19967 | -1.155 | -0.95531 |  | 8 | -0.12826
8 | YES | 265 | 0.2852 | 0.00087197 | -0.1488 | -1.155 | -1.0062 |  | 8 | -0.11748
8 | YES | 266 | 0.24974 | 0.0018643 | -0.12826 | -1.155 | -1.0267 |  | 8 | -0.15055
8 | YES | 267 | 0.25752 | 0.0021451 | -0.11748 | -1.155 | -1.0375 |  | 8 | -0.35693
8 | YES | 268 | 0.30445 | 0.0013724 | -0.15055 | -1.155 | -1.0044 |  | 8 | -0.52769
8 | YES | 269 | 0.57207 | 0.0015692 | -0.35693 | -1.155 | -0.79805 |  | 8 | -0.72222
8 | YES | 270 | 0.87831 | 0.002141 | -0.52769 | -1.155 | -0.62729 |  | 8 | -0.84868
8 | YES | 271 | 1.2268 | 0.00079044 | -0.72222 | -1.155 | -0.43276 |  | 8 | -1.2137
8 | YES | 272 | 1.7213 | 0.005514 | -0.84868 | -1.155 | -0.3063 |  | 8 | -1.3835
8 | YES | 273 | 2.0763 | 0.0055284 | -1.2137 | -1.155 | 0.058761 |  | 8 | -1.8549
8 | YES | 274 | 2.5045 | 0.0017083 | -1.3835 | -1.155 | 0.2285 |  | 8 | -2.3795
8 | YES | 275 | 3.4371 | 0.0025153 | -1.8549 | -1.155 | 0.69994 |  | 8 | -2.3988
8 | YES | 276 | 4.5215 | 0.0025791 | -2.3795 | -1.155 | 1.2245 |  | 8 | -2.319
8 | YES | 277 | 4.5762 | 0.0024665 | -2.3988 | -1.155 | 1.2439 |  | 8 | -2.3495
8 | YES | 278 | 4.593 | 0.0017391 | -2.319 | -1.155 | 1.1641 |  | 8 | -2.357
8 | YES | 279 | 4.6045 | 0.0015811 | -2.3495 | -1.155 | 1.1945 |  | 8 | -2.3403
8 | YES | 280 | 4.609 | 0.0016146 | -2.357 | -1.155 | 1.202 |  | 9 | -0.48255
8 | YES | 281 | 4.6136 | 0.0015638 | -2.3403 | -1.155 | 1.1854 |  | 9 | -0.18203
9 | YES | 296 | 0.78446 | 0.0014434 | -0.48255 | -1.1603 | -0.67778 |  | 9 | -0.16171
9 | YES | 297 | 0.34697 | 0.0020328 | -0.18203 | -1.1603 | -0.97829 |  | 9 | -0.13557
9 | YES | 298 | 0.26829 | 0.0018772 | -0.16171 | -1.1603 | -0.99862 |  | 9 | -0.22807
9 | YES | 299 | 0.25443 | 0.0023303 | -0.13557 | -1.1603 | -1.0248 |  | 9 | -0.31764
9 | YES | 300 | 0.32993 | 0.00072727 | -0.22807 | -1.1603 | -0.93226 |  | 9 | -0.44714
9 | YES | 301 | 0.45436 | 0.0007746 | -0.31764 | -1.1603 | -0.84269 |  | 9 | -0.73312
9 | YES | 302 | 0.74686 | 0.0017287 | -0.44714 | -1.1603 | -0.71319 |  | 9 | -1.002
9 | YES | 303 | 1.2119 | 0.001771 | -0.73312 | -1.1603 | -0.42721 |  | 9 | -1.5052
9 | YES | 304 | 1.807 | 0.003248 | -1.002 | -1.1603 | -0.15831 |  | 9 | -1.7338
9 | YES | 305 | 2.7778 | 0.0023032 | -1.5052 | -1.1603 | 0.34492 |  | 9 | -1.9894
9 | YES | 306 | 3.1532 | 0.0023219 | -1.7338 | -1.1603 | 0.57345 |  | 9 | -2.2831
9 | YES | 307 | 3.6766 | 0.0046223 | -1.9894 | -1.1603 | 0.82912 |  | 9 | -2.2789
9 | YES | 308 | 4.3459 | 0.0022517 | -2.2831 | -1.1603 | 1.1228 |  | 9 | -2.0531
9 | YES | 309 | 4.4501 | 0.013204 | -2.2789 | -1.1603 | 1.1186 |  | 9 | -2.2848
9 | YES | 310 | 4.5979 | 0.0017311 | -2.0531 | -1.1603 | 0.89278 |  | 9 | -2.2878
9 | YES | 311 | 4.6012 | 0.0015276 | -2.2848 | -1.1603 | 1.1245 |  | 9 | -2.2749
9 | YES | 312 | 4.6052 | 0.0019398 | -2.2878 | -1.1603 | 1.1274 |  | 9 | -2.2713
9 | YES | 313 | 4.6101 | 0.001881 | -2.2749 | -1.1603 | 1.1146 |  | 10 | -1.1033
9 | YES | 314 | 4.6127 | 0.001021 | -2.2713 | -1.1603 | 1.111 |  | 10 | -1.2132
10 | YES | 315 | 2.0817 | 0.0024677 | -1.1033 | -1.1657 | -0.062343 |  | 10 | -1.1142
10 | YES | 316 | 2.0068 | 0.0014182 | -1.2132 | -1.1657 | 0.047497 |  | 10 | -1.0156
10 | YES | 317 | 1.8758 | 0.0020761 | -1.1142 | -1.1657 | -0.051473 |  | 10 | -0.71593
10 | YES | 318 | 1.5924 | 0.0025964 | -1.0156 | -1.1657 | -0.15007 |  | 10 | -0.4561
10 | YES | 319 | 1.2051 | 0.002893 | -0.71593 | -1.1657 | -0.44975 |  | 10 | -0.182
10 | YES | 320 | 0.7649 | 0.0026344 | -0.4561 | -1.1657 | -0.70957 |  | 10 | -0.16666
10 | YES | 321 | 0.329 | 0 | -0.182 | -1.1657 | -0.98367 |  | 10 | -0.18773
10 | YES | 322 | 0.23434 | 0.0014919 | -0.16666 | -1.1657 | -0.99901 |  | 10 | -0.21784
10 | YES | 323 | 0.22027 | 0.0015492 | -0.18773 | -1.1657 | -0.97795 |  | 10 | -0.27803
10 | YES | 324 | 0.29217 | 0.00079945 | -0.21784 | -1.1657 | -0.94784 |  | 10 | -0.4945
10 | YES | 325 | 0.45698 | 0.0017534 | -0.27803 | -1.1657 | -0.88765 |  | 10 | -0.75237
10 | YES | 326 | 0.8295 | 0.00059335 | -0.4945 | -1.1657 | -0.67118 |  | 10 | -1.0257
10 | YES | 327 | 1.2576 | 0.0043092 | -0.75237 | -1.1657 | -0.4133 |  | 10 | -1.5291
10 | YES | 328 | 1.7763 | 0.0030435 | -1.0257 | -1.1657 | -0.13993 |  | 10 | -2.0452
10 | YES | 329 | 3.0289 | 0.0020587 | -1.5291 | -1.1657 | 0.36339 |  | 10 | -2.2591
10 | YES | 330 | 3.6728 | 0.0047478 | -2.0452 | -1.1657 | 0.87955 |  | 10 | -2.2598
10 | YES | 331 | 4.1619 | 0.0033134 | -2.2591 | -1.1657 | 1.0934 |  | 10 | -2.3028
10 | YES | 332 | 4.3362 | 0.0065157 | -2.2598 | -1.1657 | 1.0941 |  | 10 | -2.2638
10 | YES | 333 | 4.4581 | 0.00087248 | -2.3028 | -1.1657 | 1.1372 |  | 10 | -2.2843
10 | YES | 334 | 4.5042 | 0.0016126 | -2.2638 | -1.1657 | 1.0981 |  | 10 | -2.2945
10 | YES | 335 | 4.5157 | 0.0013758 | -2.2843 | -1.1657 | 1.1187 |  | 10 | -2.2897
10 | YES | 336 | 4.5275 | 0.0016689 | -2.2945 | -1.1657 | 1.1288 |  | 10 | -2.2981
10 | YES | 337 | 4.5323 | 0.0015943 | -2.2897 | -1.1657 | 1.1241 |  | 11 | -1.1954
10 | YES | 338 | 4.5289 | 0.0016854 | -2.2981 | -1.1657 | 1.1324 |  | 11 | -1.0723
11 | YES | 340 | 2.0236 | 0.0043433 | -1.1954 | -1.171 | 0.02434 |  | 11 | -0.91895
11 | NO | 0 | 0 | 0 | -3.175 | -1.171 | 2.004 |  | 11 | -0.67337
11 | YES | 342 | 1.8446 | 0.001982 | -1.0723 | -1.171 | -0.09867 |  | 11 | -0.45441
11 | YES | 343 | 1.5831 | 0.002152 | -0.91895 | -1.171 | -0.25207 |  | 11 | -0.15266
11 | YES | 344 | 1.1926 | 0.0022141 | -0.67337 | -1.171 | -0.49765 |  | 11 | -0.10423
11 | YES | 345 | 0.78359 | 0.0022718 | -0.45441 | -1.171 | -0.71661 |  | 11 | -0.096628
11 | YES | 346 | 0.30934 | 0.0017008 | -0.15266 | -1.171 | -1.0184 |  | 11 | -0.19291
11 | YES | 347 | 0.19877 | 0.00093773 | -0.10423 | -1.171 | -1.0668 |  | 11 | -0.39444
11 | YES | 348 | 0.21537 | 0.0014613 | -0.096628 | -1.171 | -1.0744 |  | 11 | -0.78195
11 | YES | 349 | 0.34909 | 0.00249 | -0.19291 | -1.171 | -0.97811 |  | 11 | -1.3573
11 | YES | 350 | 0.62656 | 0.0032375 | -0.39444 | -1.171 | -0.77658 |  | 11 | -1.5525
11 | YES | 351 | 1.3231 | 0.0017168 | -0.78195 | -1.171 | -0.38907 |  | 11 | -1.8786
11 | YES | 352 | 2.4287 | 0.0014037 | -1.3573 | -1.171 | 0.18626 |  | 11 | -2.0525
11 | YES | 353 | 3.0405 | 0.003679 | -1.5525 | -1.171 | 0.38149 |  | 11 | -2.1975
11 | YES | 354 | 3.4124 | 0.0028862 | -1.8786 | -1.171 | 0.70755 |  | 11 | -2.1823
11 | YES | 355 | 3.7705 | 0.0031573 | -2.0525 | -1.171 | 0.88148 |  | 11 | -2.3017
11 | YES | 356 | 4.0585 | 0.0032993 | -2.1975 | -1.171 | 1.0265 |  | 11 | -2.1963
11 | YES | 357 | 4.1637 | 0.0017177 | -2.1823 | -1.171 | 1.0113 |  | 11 | -2.2154
11 | YES | 358 | 4.4143 | 0.0018895 | -2.3017 | -1.171 | 1.1307 |  | 11 | -2.2261
11 | YES | 359 | 4.3857 | 0.0012572 | -2.1963 | -1.171 | 1.0253 |  | 11 | -2.197
11 | YES | 360 | 4.3946 | 0.0015036 | -2.2154 | -1.171 | 1.0444 |  | 11 | -2.2028
11 | YES | 361 | 4.4029 | 0.0022066 | -2.2261 | -1.171 | 1.0551 |  | 12 | -1.2765
11 | YES | 362 | 4.403 | 0.0018797 | -2.197 | -1.171 | 1.026 |  | 12 | -1.148
11 | YES | 363 | 4.4072 | 0.001711 | -2.2028 | -1.171 | 1.0318 |  | 12 | -1.079
12 | YES | 366 | 1.9875 | 0.0056304 | -1.2765 | -1.1764 | 0.10013 |  | 12 | -0.92051
12 | YES | 367 | 1.977 | 0.0031318 | -1.148 | -1.1764 | -0.028357 |  | 12 | -0.68736
12 | YES | 368 | 1.844 | 0 | -1.079 | -1.1764 | -0.097367 |  | 12 | -0.44562
12 | YES | 369 | 1.5765 | 0.0026016 | -0.92051 | -1.1764 | -0.25585 |  | 12 | -0.16092
12 | YES | 370 | 1.2116 | 0.0013283 | -0.68736 | -1.1764 | -0.48901 |  | 12 | -0.10074
12 | YES | 371 | 0.79838 | 0.0016546 | -0.44562 | -1.1764 | -0.73075 |  | 12 | -0.098033
12 | YES | 372 | 0.32808 | 0.00064014 | -0.16092 | -1.1764 | -1.0154 |  | 12 | -0.17107
12 | YES | 373 | 0.21626 | 0.0017867 | -0.10074 | -1.1764 | -1.0756 |  | 12 | -0.34906
12 | YES | 374 | 0.21597 | 0.00036364 | -0.098033 | -1.1764 | -1.0783 |  | 12 | -0.68064
12 | YES | 375 | 0.32793 | 0.00051211 | -0.17107 | -1.1764 | -1.0053 |  | 12 | -1.0426
12 | YES | 376 | 0.61294 | 0.0017715 | -0.34906 | -1.1764 | -0.82731 |  | 12 | -1.4142
12 | YES | 377 | 1.2254 | 0.0030304 | -0.68064 | -1.1764 | -0.49573 |  | 12 | -1.4913
12 | YES | 378 | 1.8644 | 0.0019371 | -1.0426 | -1.1764 | -0.13374 |  | 12 | -1.7719
12 | YES | 379 | 2.4838 | 0.0092709 | -1.4142 | -1.1764 | 0.23784 |  | 12 | -1.8517
12 | YES | 380 | 3.0897 | 0.0023932 | -1.4913 | -1.1764 | 0.31493 |  | 12 | -2.063
12 | YES | 381 | 3.2801 | 0.0019439 | -1.7719 | -1.1764 | 0.59556 |  | 12 | -2.1979
12 | YES | 382 | 3.4803 | 0.0025891 | -1.8517 | -1.1764 | 0.67535 |  | 12 | -2.2531
12 | YES | 383 | 3.857 | 0.009024 | -2.063 | -1.1764 | 0.8866 |  | 12 | -2.2416
12 | YES | 384 | 4.1701 | 0.0022345 | -2.1979 | -1.1764 | 1.0216 |  | 12 | -2.284
12 | YES | 385 | 4.3789 | 0.002038 | -2.2531 | -1.1764 | 1.0768 |  | 12 | -2.261
12 | YES | 386 | 4.4604 | 0.0012024 | -2.2416 | -1.1764 | 1.0652 |  | 12 | -2.2587
12 | YES | 387 | 4.462 | 0.0016602 | -2.284 | -1.1764 | 1.1077 |  | 13 | -1.1478
12 | YES | 388 | 4.464 | 0.0003873 | -2.261 | -1.1764 | 1.0846 |  | 13 | -1.1126
12 | YES | 389 | 4.4673 | 0.0012418 | -2.2587 | -1.1764 | 1.0824 |  | 13 | -1.0651
13 | YES | 392 | 1.9142 | 0.0022484 | -1.1478 | -1.1817 | -0.033903 |  | 13 | -0.90141
13 | YES | 393 | 1.8954 | 0.0052708 | -1.1126 | -1.1817 | -0.069153 |  | 13 | -0.71727
13 | YES | 394 | 1.7689 | 0.0029451 | -1.0651 | -1.1817 | -0.11661 |  | 13 | -0.46879
13 | YES | 395 | 1.5056 | 0.0026573 | -0.90141 | -1.1817 | -0.2803 |  | 13 | -0.15896
13 | YES | 396 | 1.1867 | 0.0025 | -0.71727 | -1.1817 | -0.46444 |  | 13 | -0.088537
13 | YES | 397 | 0.82421 | 0.00099931 | -0.46879 | -1.1817 | -0.71292 |  | 13 | -0.070752
13 | YES | 398 | 0.34704 | 0.00045454 | -0.15896 | -1.1817 | -1.0228 |  | 13 | -0.25675
13 | YES | 399 | 0.21146 | 0.0015548 | -0.088537 | -1.1817 | -1.0932 |  | 13 | -0.30009
13 | YES | 400 | 0.17125 | 0.0019928 | -0.070752 | -1.1817 | -1.111 |  | 13 | -0.94149
13 | YES | 401 | 0.24225 | 0.0024334 | -0.25675 | -1.1817 | -0.92496 |  | 13 | -1.2091
13 | YES | 402 | 0.55191 | 0.0017888 | -0.30009 | -1.1817 | -0.88162 |  | 13 | -1.4097
13 | YES | 403 | 1.3935 | 0.0015659 | -0.94149 | -1.1817 | -0.24023 |  | 13 | -1.515
13 | YES | 404 | 2.1879 | 0.0017008 | -1.2091 | -1.1817 | 0.027407 |  | 13 | -1.7168
13 | YES | 405 | 2.6173 | 0.0019649 | -1.4097 | -1.1817 | 0.22799 |  | 13 | -1.8006
13 | YES | 406 | 3.102 | 0.001981 | -1.515 | -1.1817 | 0.33331 |  | 13 | -2.0368
13 | YES | 407 | 3.1672 | 0.0012974 | -1.7168 | -1.1817 | 0.53511 |  | 13 | -2.2717
13 | YES | 408 | 3.4144 | 0.0037168 | -1.8006 | -1.1817 | 0.61885 |  | 13 | -2.2861
13 | YES | 409 | 3.7742 | 0.005372 | -2.0368 | -1.1817 | 0.85507 |  | 13 | -2.2247
13 | YES | 410 | 4.3573 | 0.0032074 | -2.2717 | -1.1817 | 1.09 |  | 13 | -2.2766
13 | YES | 411 | 4.4359 | 0.00062452 | -2.2861 | -1.1817 | 1.1044 |  | 13 | -2.2626
13 | YES | 412 | 4.4173 | 0.0018245 | -2.2247 | -1.1817 | 1.0429 |  | 13 | -2.2524
13 | YES | 413 | 4.4224 | 0.0011819 | -2.2766 | -1.1817 | 1.0949 |  | 14 | -1.0436
13 | YES | 414 | 4.4264 | 0.0011142 | -2.2626 | -1.1817 | 1.0809 |  | 14 | -0.84295
13 | YES | 415 | 4.4256 | 0.001617 | -2.2524 | -1.1817 | 1.0707 |  | 14 | -0.69182
14 | YES | 418 | 1.6874 | 0.0018858 | -1.0436 | -1.1871 | -0.14345 |  | 14 | -0.47956
14 | YES | 419 | 1.4561 | 0.00054546 | -0.84295 | -1.1871 | -0.34411 |  | 14 | -0.2328
14 | YES | 420 | 1.2162 | 0.0012845 | -0.69182 | -1.1871 | -0.49524 |  | 14 | -0.10179
14 | YES | 421 | 0.87144 | 0.0031699 | -0.47956 | -1.1871 | -0.7075 |  | 14 | -0.058372
14 | YES | 422 | 0.4072 | 0.00087197 | -0.2328 | -1.1871 | -0.95426 |  | 14 | -0.11096
14 | YES | 423 | 0.22221 | 0.0026106 | -0.10179 | -1.1871 | -1.0853 |  | 14 | -0.21788
14 | YES | 424 | 0.15563 | 0.0017895 | -0.058372 | -1.1871 | -1.1287 |  | 14 | -0.64533
14 | YES | 425 | 0.15604 | 0.0017146 | -0.11096 | -1.1871 | -1.0761 |  | 14 | -1.0978
14 | YES | 426 | 0.31112 | 0.0017088 | -0.21788 | -1.1871 | -0.96918 |  | 14 | -1.2942
14 | YES | 427 | 1.0677 | 0.0033176 | -0.64533 | -1.1871 | -0.54173 |  | 14 | -1.5737
14 | YES | 428 | 1.9182 | 0.0032042 | -1.0978 | -1.1871 | -0.08924 |  | 14 | -1.4983
14 | YES | 429 | 2.2928 | 0.00092633 | -1.2942 | -1.1871 | 0.10716 |  | 14 | -1.5191
14 | YES | 430 | 2.8803 | 0.002273 | -1.5737 | -1.1871 | 0.38667 |  | 14 | -1.9213
14 | YES | 431 | 3.1247 | 0.0020064 | -1.4983 | -1.1871 | 0.3112 |  | 14 | -2.2354
14 | YES | 432 | 2.9449 | 0.0018055 | -1.5191 | -1.1871 | 0.33206 |  | 14 | -2.1936
14 | YES | 433 | 3.5627 | 0.002504 | -1.9213 | -1.1871 | 0.73426 |  | 14 | -2.2862
14 | YES | 434 | 4.2226 | 0.0023223 | -2.2354 | -1.1871 | 1.0483 |  | 14 | -2.6577
14 | YES | 435 | 4.3244 | 0.0017313 | -2.1936 | -1.1871 | 1.0065 |  | 14 | -2.1795
14 | YES | 436 | 4.3228 | 0.0020977 | -2.2862 | -1.1871 | 1.0992 |  | 14 | -2.2125
14 | YES | 437 | 4.3443 | 0.002411 | -2.6577 | -1.1871 | 1.4707 |  | 15 | -1.2821
14 | YES | 438 | 4.3595 | 0.0029918 | -2.1795 | -1.1871 | 0.99247 |  | 15 | -0.21554
14 | YES | 439 | 4.3625 | 0.0029554 | -2.2125 | -1.1871 | 1.0255 |  | 15 | -0.13093
15 | YES | 441 | 0.84788 | 0.0020904 | -1.2821 | -1.1924 | 0.089713 |  | 15 | -0.10345
15 | YES | 442 | 0.38846 | 0.0015548 | -0.21554 | -1.1924 | -0.97687 |  | 15 | -0.12119
15 | YES | 443 | 0.22507 | 0.00072727 | -0.13093 | -1.1924 | -1.0615 |  | 15 | -0.20373
15 | YES | 444 | 0.16955 | 0.0019577 | -0.10345 | -1.1924 | -1.089 |  | 15 | -0.66156
15 | YES | 445 | 0.19381 | 0.0018363 | -0.12119 | -1.1924 | -1.0712 |  | 15 | -0.87577
15 | YES | 446 | 0.34427 | 0.0015492 | -0.20373 | -1.1924 | -0.98868 |  | 15 | -1.1463
15 | YES | 447 | 1.1684 | 0.0015963 | -0.66156 | -1.1924 | -0.53084 |  | 15 | -1.4824
15 | YES | 448 | 1.5772 | 0.0039386 | -0.87577 | -1.1924 | -0.31664 |  | 15 | -1.6392
15 | YES | 449 | 2.0046 | 0.002128 | -1.1463 | -1.1924 | -0.046057 |  | 15 | -1.7454
15 | YES | 450 | 2.6787 | 0.0081011 | -1.4824 | -1.1924 | 0.28994 |  | 15 | -1.7734
15 | YES | 451 | 3.0428 | 0.0013396 | -1.6392 | -1.1924 | 0.44676 |  | 15 | -2.2294
15 | YES | 452 | 3.2226 | 0.0019595 | -1.7454 | -1.1924 | 0.55298 |  | 15 | -2.2349
15 | YES | 453 | 3.4016 | 0.0017677 | -1.7734 | -1.1924 | 0.58099 |  | 15 | -2.269
15 | YES | 454 | 4.3486 | 0.0013157 | -2.2294 | -1.1924 | 1.037 |  | 15 | -2.2529
15 | YES | 455 | 4.4121 | 0.00081207 | -2.2349 | -1.1924 | 1.0425 |  | 15 | -2.2362
15 | YES | 456 | 4.428 | 0.0012676 | -2.269 | -1.1924 | 1.0766 |  | 15 | -2.3651
15 | YES | 457 | 4.4291 | 0.0013817 | -2.2529 | -1.1924 | 1.0605 |  | 16 | -0.74349
15 | YES | 458 | 4.4308 | 0.0022261 | -2.2362 | -1.1924 | 1.0438 |  | 16 | -0.52011
15 | YES | 459 | 4.4309 | 0.00048819 | -2.3651 | -1.1924 | 1.1727 |  | 16 | -0.31237
16 | YES | 461 | 1.1895 | 0.0015007 | -0.74349 | -1.1978 | -0.45427 |  | 16 | -0.1746
16 | YES | 462 | 0.90289 | 0.0024793 | -0.52011 | -1.1978 | -0.67765 |  | 16 | -0.075008
16 | YES | 463 | 0.43963 | 0.0013174 | -0.31237 | -1.1978 | -0.88538 |  | 16 | -0.11404
16 | YES | 464 | 0.2794 | 0.0012005 | -0.1746 | -1.1978 | -1.0232 |  | 16 | -0.18919
16 | YES | 465 | 0.16799 | 0.00058446 | -0.075008 | -1.1978 | -1.1227 |  | 16 | -0.56999
16 | NO | 0 | 0 | 0 | -1.383 | -1.1978 | 0.18525 |  | 16 | -0.87488
16 | YES | 466 | 0.16296 | 0.0017146 | -0.11404 | -1.1978 | -1.0837 |  | 16 | -1.1046
16 | YES | 467 | 0.35181 | 0.0019075 | -0.18919 | -1.1978 | -1.0086 |  | 16 | -1.3585
16 | YES | 468 | 0.98201 | 0.0017344 | -0.56999 | -1.1978 | -0.62776 |  | 16 | -1.456
16 | YES | 469 | 1.5411 | 0.0016861 | -0.87488 | -1.1978 | -0.32288 |  | 16 | -1.6382
16 | YES | 470 | 1.8523 | 0.0019439 | -1.1046 | -1.1978 | -0.093104 |  | 16 | -1.6541
16 | YES | 471 | 2.4955 | 0.0034907 | -1.3585 | -1.1978 | 0.16076 |  | 16 | -2.0075
16 | YES | 472 | 2.749 | 0.001668 | -1.456 | -1.1978 | 0.25822 |  | 16 | -2.2525
16 | YES | 473 | 2.8798 | 0.0018226 | -1.6382 | -1.1978 | 0.44044 |  | 16 | -2.2031
16 | YES | 474 | 3.1219 | 0.0013291 | -1.6541 | -1.1978 | 0.45634 |  | 16 | -2.2173
16 | YES | 475 | 3.5725 | 0.0039436 | -2.0075 | -1.1978 | 0.80976 |  | 16 | -2.2171
16 | YES | 476 | 4.3425 | 0.0017893 | -2.2525 | -1.1978 | 1.0547 |  | 16 | -2.2572
16 | YES | 477 | 4.3829 | 0.0013355 | -2.2031 | -1.1978 | 1.0053 |  | 17 | -0.7003
16 | YES | 478 | 4.4097 | 0.0021378 | -2.2173 | -1.1978 | 1.0196 |  | 17 | -0.67292
16 | YES | 479 | 4.4159 | 0.0015077 | -2.2171 | -1.1978 | 1.0193 |  | 17 | -0.5019
16 | YES | 480 | 4.4178 | 0.0017766 | -2.2572 | -1.1978 | 1.0594 |  | 17 | -0.24887
17 | YES | 483 | 1.1427 | 0.0025483 | -0.7003 | -1.2031 | -0.5028 |  | 17 | -0.13361
17 | YES | 484 | 1.1621 | 0.00090911 | -0.67292 | -1.2031 | -0.53018 |  | 17 | -0.1299
17 | YES | 485 | 0.8981 | 0.0024372 | -0.5019 | -1.2031 | -0.7012 |  | 17 | -0.093562
17 | YES | 486 | 0.45113 | 0.0021831 | -0.24887 | -1.2031 | -0.95423 |  | 17 | -0.13651
17 | YES | 487 | 0.26439 | 0.0015512 | -0.13361 | -1.2031 | -1.0695 |  | 17 | -0.46688
17 | YES | 488 | 0.1841 | 0.00062456 | -0.1299 | -1.2031 | -1.0732 |  | 17 | -0.74576
17 | YES | 489 | 0.14844 | 0.0013958 | -0.093562 | -1.2031 | -1.1095 |  | 17 | -0.9932
17 | YES | 490 | 0.24249 | 0.0019413 | -0.13651 | -1.2031 | -1.0666 |  | 17 | -1.2601
17 | YES | 491 | 0.84912 | 0.0018054 | -0.46688 | -1.2031 | -0.73622 |  | 17 | -1.1856
17 | YES | 492 | 1.3692 | 0.0009489 | -0.74576 | -1.2031 | -0.45734 |  | 17 | -1.1637
17 | YES | 493 | 1.8068 | 0.0018557 | -0.9932 | -1.2031 | -0.2099 |  | 17 | -1.348
17 | YES | 494 | 2.2439 | 0.0033862 | -1.2601 | -1.2031 | 0.056989 |  | 17 | -1.5634
17 | YES | 495 | 2.2014 | 0.0032068 | -1.1856 | -1.2031 | -0.017461 |  | 17 | -1.9608
17 | YES | 496 | 2.2263 | 0.0069041 | -1.1637 | -1.2031 | -0.039411 |  | 17 | -2.1901
17 | YES | 497 | 2.521 | 0.0025478 | -1.348 | -1.2031 | 0.14492 |  | 17 | -2.225
17 | YES | 498 | 3.0116 | 0.0016894 | -1.5634 | -1.2031 | 0.36028 |  | 17 | -2.1828
17 | YES | 499 | 3.5302 | 0.0068678 | -1.9608 | -1.2031 | 0.75774 |  | 17 | -2.1911
17 | YES | 500 | 4.3769 | 0.0019192 | -2.1901 | -1.2031 | 0.98699 |  | 18 | -0.58406
17 | YES | 501 | 4.414 | 0.001449 | -2.225 | -1.2031 | 1.0219 |  | 18 | -0.42645
17 | YES | 502 | 4.4292 | 0.0009602 | -2.1828 | -1.2031 | 0.97971 |  | 18 | -0.12394
17 | YES | 503 | 4.4269 | 0.0010634 | -2.1911 | -1.2031 | 0.98795 |  | 18 | -0.058984
18 | YES | 506 | 0.69494 | 0.00077779 | -0.58406 | -1.2084 | -0.62439 |  | 18 | -0.11489
18 | YES | 507 | 0.44155 | 0.0014434 | -0.42645 | -1.2084 | -0.78199 |  | 18 | -0.30103
18 | YES | 508 | 0.26006 | 0.0017715 | -0.12394 | -1.2084 | -1.0845 |  | 18 | -0.61578
18 | YES | 509 | 0.15902 | 0.0018529 | -0.058984 | -1.2084 | -1.1495 |  | 18 | -0.65826
18 | YES | 510 | 0.18711 | 0.0016723 | -0.11489 | -1.2084 | -1.0936 |  | 18 | -0.68814
18 | YES | 511 | 0.50397 | 0.00036364 | -0.30103 | -1.2084 | -0.90741 |  | 18 | -0.80225
18 | YES | 512 | 1.0882 | 0.0018506 | -0.61578 | -1.2084 | -0.59267 |  | 18 | -0.80311
18 | YES | 513 | 1.2397 | 0.00099809 | -0.65826 | -1.2084 | -0.55018 |  | 18 | -0.8403
18 | YES | 514 | 1.3259 | 0.00064942 | -0.68814 | -1.2084 | -0.52031 |  | 18 | -0.90175
18 | YES | 515 | 1.4877 | 0.0026054 | -0.80225 | -1.2084 | -0.4062 |  | 18 | -1.0214
18 | YES | 516 | 1.5429 | 0.0016723 | -0.80311 | -1.2084 | -0.40534 |  | 18 | -1.2107
18 | YES | 517 | 1.6147 | 0.0017061 | -0.8403 | -1.2084 | -0.36815 |  | 18 | -1.4535
18 | YES | 518 | 1.7423 | 0.0015614 | -0.90175 | -1.2084 | -0.3067 |  | 18 | -2.2462
18 | YES | 519 | 1.9836 | 0.0018574 | -1.0214 | -1.2084 | -0.18709 |  | 18 | -2.3363
18 | YES | 520 | 2.3543 | 0.0017361 | -1.2107 | -1.2084 | 0.0022425 |  | 18 | -2.2833
18 | YES | 521 | 2.8475 | 0.0021951 | -1.4535 | -1.2084 | 0.24506
18 | YES | 522 | 4.4148 | 0.0016083 | -2.2462 | -1.2084 | 1.0378
18 | YES | 523 | 4.4787 | 0.0017787 | -2.3363 | -1.2084 | 1.1279
18 | YES | 524 | 4.5417 | 0.0016444 | -2.2833 | -1.2084 | 1.0749
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. CTD DOX
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=CTD DOX
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.99502 |  | X | CALC Y
m= | -0.50072 | 0.0023159 |  | Std Dev of Resid |  | 0.060276 |  | 0.14844 | -0.11924
b= | -0.044914 | 0.0068476 |  | Resid Sum of Squares |  | 0.85016 |  | 4.6136 | -2.3551
# of Values Extracted = |  |  | 367
# of Values Used in Fit = |  |  | 236
File | Use | Sample # | CTD DOX | CTD DOX | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | CTD DOX
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 4 | 0.048 | 0 | 0.004 | -0.068949 | -0.072949 |  | 0.2704 | -0.1416
1 | NO | 5 | 0.04676 | 0.0014834 | -0.05324 | -0.068328 | -0.015089 |  | 0.18318 | -0.11782
1 | NO | 6 | 0.50074 | 0.0087174 | -0.0012562 | -0.29565 | -0.29439 |  | 0.17374 | -0.10126
1 | NO | 7 | 2.8522 | 0.0043797 | -1.0448 | -1.4731 | -0.42828 |  | 0.33051 | -0.24649
1 | NO | 8 | 3.2522 | 0.002912 | -1.1088 | -1.6734 | -0.5646 |  | 0.33226 | -0.26674
1 | NO | 9 | 3.687 | 0.0021289 | -1.264 | -1.8911 | -0.62702 |  | 0.99696 | -0.54804
1 | NO | 10 | 3.877 | 0.013065 | -1.411 | -1.9862 | -0.57517 |  | 1.6957 | -0.86334
1 | NO | 11 | 4.2193 | 0.0032873 | -1.3767 | -2.1576 | -0.78095 |  | 2.0467 | -1.0023
2 | NO | 13 | 3.4827 | 0.0042317 | -1.3573 | -1.7888 | -0.43149 |  | 2.5483 | -1.2267
2 | NO | 14 | 4.2358 | 0.0015184 | -1.4682 | -2.1659 | -0.69771 |  | 2.8469 | -1.3421
2 | NO | 15 | 4.3923 | 0.0015558 | -1.4437 | -2.2442 | -0.80055 |  | 3.0726 | -1.4324
2 | NO | 16 | 4.3895 | 0.0016078 | -1.4255 | -2.2428 | -0.81732 |  | 0.25452 | -0.18648
2 | NO | 17 | 4.3972 | 0.0035285 | -1.4218 | -2.2467 | -0.82494 |  | 0.1806 | -0.1434
2 | NO | 18 | 4.5156 | 0.0024659 | -1.4634 | -2.3059 | -0.84252 |  | 0.18014 | -0.13086
2 | NO | 19 | 4.6644 | 0.012831 | -1.4856 | -2.3805 | -0.89484 |  | 0.28279 | -0.19421
2 | NO | 20 | 4.6928 | 0.0017865 | -1.5192 | -2.3947 | -0.87546 |  | 0.74521 | -0.41079
3 | NO | 22 | 0.3693 | 0.0017447 | -0.1977 | -0.22983 | -0.032127 |  | 1.0969 | -0.58407
3 | NO | 23 | 0.21 | 0 | -0.099 | -0.15007 | -0.051066 |  | 1.3515 | -0.78649
3 | NO | 24 | 0.20993 | 0.00057973 | -0.12207 | -0.15003 | -0.027954 |  | 1.8271 | -0.95894
3 | NO | 25 | 0.16153 | 0.0013606 | -0.17247 | -0.1258 | 0.046676 |  | 2.4851 | -1.2899
3 | NO | 26 | 0.1613 | 0.0010539 | -0.2097 | -0.12568 | 0.084023 |  | 2.9115 | -1.4495
3 | NO | 27 | 0.4346 | 0.0012006 | -0.2234 | -0.26253 | -0.039132 |  | 3.944 | -1.946
3 | NO | 28 | 0.92265 | 0.0023447 | -0.40936 | -0.5069 | -0.097546 |  | 4.4955 | -2.1335
3 | NO | 29 | 1.3029 | 0.0011328 | -0.70209 | -0.69731 | 0.0047842 |  | 0.2222 | -0.1298
3 | NO | 31 | 1.9264 | 0.0010691 | -0.7866 | -1.0095 | -0.22291 |  | 0.22228 | -0.090719
3 | NO | 32 | 1.8416 | 0.0012006 | -0.7414 | -0.96704 | -0.22564 |  | 0.186 | -0.07
3 | NO | 33 | 2.0064 | 0.0013354 | -0.80164 | -1.0495 | -0.2479 |  | 0.22246 | -0.080546
3 | NO | 34 | 2.3332 | 0.0016834 | -0.94776 | -1.2132 | -0.26545 |  | 0.47631 | -0.23669
3 | NO | 35 | 3.2813 | 0.0027519 | -1.3317 | -1.6879 | -0.35627 |  | 1.1175 | -0.57549
3 | NO | 36 | 4.0995 | 0.0030114 | -1.5895 | -2.0976 | -0.50807 |  | 1.569 | -0.84797
3 | NO | 37 | 4.4806 | 0.0012583 | -1.7234 | -2.2885 | -0.5651 |  | 1.9172 | -0.97879
3 | NO | 38 | 4.4338 | 0.0029481 | -1.7332 | -2.265 | -0.53177 |  | 2.5827 | -1.3153
3 | NO | 39 | 4.5799 | 0.0018105 | -1.7302 | -2.3381 | -0.60798 |  | 3.4436 | -1.7284
3 | NO | 40 | 4.5822 | 0.001127 | -1.7118 | -2.3393 | -0.62753 |  | 3.6641 | -1.8489
4 | NO | 58 | 1.3029 | 0.0022558 | -0.64814 | -0.69728 | -0.049142 |  | 4.197 | -2.108
4 | NO | 59 | 0.95083 | 0.0012803 | -0.53817 | -0.52102 | 0.017149 |  | 4.5605 | -2.2185
4 | NO | 60 | 0.478 | 0 | -0.218 | -0.28426 | -0.066258 |  | 4.5715 | -2.2005
4 | YES | 61 | 0.2704 | 0.0012005 | -0.1416 | -0.18031 | -0.038704 |  | 4.5812 | -2.1808
4 | YES | 62 | 0.18318 | 0.0023805 | -0.11782 | -0.13664 | -0.018819 |  | 4.5865 | -2.1774
4 | YES | 63 | 0.17374 | 0.00099807 | -0.10126 | -0.13191 | -0.030643 |  | 0.22352 | -0.084479
4 | YES | 64 | 0.33051 | 0.001674 | -0.24649 | -0.21041 | 0.036086 |  | 0.1986 | -0.068397
4 | YES | 65 | 0.33226 | 0.0014866 | -0.26674 | -0.21128 | 0.05546 |  | 0.22269 | -0.10731
4 | YES | 66 | 0.99696 | 0.0016852 | -0.54804 | -0.54411 | 0.0039298 |  | 0.32331 | -0.14869
4 | YES | 67 | 1.6957 | 0.002548 | -0.86334 | -0.89396 | -0.030625 |  | 0.80869 | -0.49131
4 | YES | 68 | 2.0467 | 0.0023078 | -1.0023 | -1.0697 | -0.067377 |  | 1.2606 | -0.67841
4 | YES | 69 | 2.5483 | 0.0018423 | -1.2267 | -1.3209 | -0.094198 |  | 1.6433 | -0.91867
4 | YES | 70 | 2.8469 | 0.0011569 | -1.3421 | -1.4704 | -0.12825 |  | 2.3511 | -1.2879
4 | YES | 71 | 3.0726 | 0.0018841 | -1.4324 | -1.5834 | -0.15109 |  | 2.604 | -1.211
4 | NO | 72 | 3.1943 | 0.0016615 | -3.2447 | -1.6444 | 1.6003 |  | 2.871 | -1.521
4 | NO | 73 | 3.1975 | 0.0013041 | -1.4465 | -1.646 | -0.19952 |  | 3.3187 | -1.7093
4 | NO | 74 | 4.4961 | 0.0029635 | -2.0579 | -2.2962 | -0.23834 |  | 4.4733 | -2.2897
4 | NO | 75 | 4.5756 | 0.0016338 | -2.6084 | -2.336 | 0.27235 |  | 4.5131 | -2.2709
4 | NO | 76 | 4.5983 | 0.0018676 | -2.0307 | -2.3473 | -0.3166 |  | 4.5501 | -2.2669
4 | NO | 77 | 4.5999 | 0.0019402 | -2.0381 | -2.3482 | -0.31015 |  | 4.5613 | -2.2247
4 | NO | 78 | 4.6054 | 0.0016601 | -1.9826 | -2.3509 | -0.36831 |  | 4.5683 | -2.2487
5 | NO | 82 | 1.5379 | 0.0019487 | -0.95205 | -0.815 | 0.13705 |  | 4.5684 | -2.2656
5 | NO | 83 | 1.5143 | 0.0013079 | -0.8287 | -0.80315 | 0.025548 |  | 0.2852 | -0.1488
5 | NO | 84 | 1.2764 | 0.0014111 | -0.7406 | -0.68403 | 0.05657 |  | 0.24974 | -0.12826
5 | NO | 85 | 0.89347 | 0.0028754 | -0.50653 | -0.49229 | 0.014236 |  | 0.25752 | -0.11748
5 | NO | 86 | 0.43764 | 0.0024118 | -0.31036 | -0.26405 | 0.046317 |  | 0.30445 | -0.15055
5 | YES | 87 | 0.25452 | 0.0027931 | -0.18648 | -0.17236 | 0.014121 |  | 0.57207 | -0.35693
5 | YES | 88 | 0.1806 | 0.0012006 | -0.1434 | -0.13535 | 0.0080511 |  | 0.87831 | -0.52769
5 | YES | 89 | 0.18014 | 0.0016498 | -0.13086 | -0.13511 | -0.004255 |  | 1.2268 | -0.72222
5 | NO | 90 | 0.18255 | 0.0014434 | -0.44645 | -0.13632 | 0.31014 |  | 1.7213 | -0.84868
5 | YES | 91 | 0.28279 | 0.0016017 | -0.19421 | -0.18651 | 0.0076926 |  | 2.0763 | -1.2137
5 | YES | 92 | 0.74521 | 0.0019745 | -0.41079 | -0.41805 | -0.0072611 |  | 2.5045 | -1.3835
5 | YES | 93 | 1.0969 | 0.0017547 | -0.58407 | -0.59417 | -0.010103 |  | 3.4371 | -1.8549
5 | YES | 94 | 1.3515 | 0.0014383 | -0.78649 | -0.72164 | 0.064846 |  | 4.5215 | -2.3795
5 | YES | 95 | 1.8271 | 0.0020829 | -0.95894 | -0.95976 | -0.0008169 |  | 4.5762 | -2.3988
5 | NO | 96 | 1.9475 | 0.0013421 | -0.87647 | -1.0201 | -0.14361 |  | 4.593 | -2.319
5 | YES | 97 | 2.4851 | 0.0016243 | -1.2899 | -1.2892 | 0.0007077 |  | 4.6045 | -2.3495
5 | YES | 98 | 2.9115 | 0.014906 | -1.4495 | -1.5028 | -0.053244 |  | 4.609 | -2.357
5 | YES | 99 | 3.944 | 0.0056295 | -1.946 | -2.0198 | -0.073782 |  | 4.6136 | -2.3403
5 | YES | 100 | 4.4955 | 0.0013044 | -2.1335 | -2.2959 | -0.16237 |  | 0.26829 | -0.16171
5 | NO | 101 | 4.5092 | 0.0018394 | -2.1058 | -2.3027 | -0.19693 |  | 0.25443 | -0.13557
5 | NO | 102 | 4.5181 | 0.0014259 | -2.0639 | -2.3072 | -0.2433 |  | 0.32993 | -0.22807
5 | NO | 103 | 4.5223 | 0.0015575 | -2.0937 | -2.3093 | -0.21568 |  | 0.45436 | -0.31764
5 | NO | 104 | 4.524 | 0.0012706 | -2.1109 | -2.3102 | -0.19924 |  | 0.74686 | -0.44714
6 | NO | 122 | 1.8111 | 0.00089619 | -1.0939 | -0.95178 | 0.1421 |  | 1.2119 | -0.73312
6 | NO | 123 | 1.7655 | 0.0015809 | -0.95145 | -0.92896 | 0.022486 |  | 1.807 | -1.002
6 | NO | 124 | 1.5257 | 0.0012493 | -0.7953 | -0.80886 | -0.013564 |  | 2.7778 | -1.5052
6 | NO | 125 | 1.2279 | 0.00076817 | -0.6011 | -0.65975 | -0.058649 |  | 3.1532 | -1.7338
6 | NO | 126 | 0.8526 | 0.0026692 | -0.4314 | -0.47183 | -0.040432 |  | 3.6766 | -1.9894
6 | NO | 127 | 0.36298 | 0.0024562 | -0.17702 | -0.22667 | -0.04965 |  | 4.3459 | -2.2831
6 | YES | 128 | 0.2222 | 0.0021433 | -0.1298 | -0.15617 | -0.026371 |  | 4.4501 | -2.2789
6 | YES | 129 | 0.22228 | 0.0021917 | -0.090719 | -0.15621 | -0.065496 |  | 4.6012 | -2.2848
6 | YES | 130 | 0.186 | 0 | -0.07 | -0.13805 | -0.068048 |  | 4.6052 | -2.2878
6 | YES | 131 | 0.22246 | 0.0019322 | -0.080546 | -0.1563 | -0.075756 |  | 4.6101 | -2.2749
6 | YES | 132 | 0.47631 | 0.0014436 | -0.23669 | -0.28341 | -0.046728 |  | 4.6127 | -2.2713
6 | YES | 133 | 1.1175 | 0.0017084 | -0.57549 | -0.60447 | -0.028986 |  | 0.23434 | -0.16666
6 | YES | 134 | 1.569 | 0.0039453 | -0.84797 | -0.83056 | 0.017409 |  | 0.22027 | -0.18773
6 | YES | 135 | 1.9172 | 0.0018221 | -0.97879 | -1.0049 | -0.026114 |  | 0.29217 | -0.21784
6 | YES | 136 | 2.5827 | 0.0021456 | -1.3153 | -1.3381 | -0.022793 |  | 0.45698 | -0.27803
6 | NO | 137 | 3.1111 | 0.0014921 | -1.4349 | -1.6027 | -0.16777 |  | 0.8295 | -0.4945
6 | YES | 138 | 3.4436 | 0.0019536 | -1.7284 | -1.7692 | -0.040852 |  | 1.2576 | -0.75237
6 | YES | 139 | 3.6641 | 0.001536 | -1.8489 | -1.8796 | -0.030731 |  | 1.7763 | -1.0257
6 | YES | 140 | 4.197 | 0.0031303 | -2.108 | -2.1464 | -0.038389 |  | 3.0289 | -1.5291
6 | YES | 141 | 4.5605 | 0.0015763 | -2.2185 | -2.3284 | -0.10992 |  | 4.1619 | -2.2591
6 | YES | 142 | 4.5715 | 0.0010493 | -2.2005 | -2.3339 | -0.13341 |  | 4.3362 | -2.2598
6 | YES | 143 | 4.5812 | 0.0013225 | -2.1808 | -2.3388 | -0.15803 |  | 4.4581 | -2.3028
6 | YES | 144 | 4.5865 | 0.0018528 | -2.1774 | -2.3415 | -0.16404 |  | 4.5042 | -2.2638
6 | NO | 145 | 4.5886 | 0.0013785 | -2.1664 | -2.3425 | -0.17618 |  | 4.5157 | -2.2843
7 | NO | 186 | 1.9311 | 0.0012286 | -1.1079 | -1.0119 | 0.096016 |  | 4.5275 | -2.2945
7 | NO | 187 | 1.8347 | 0.002743 | -1.0223 | -0.96359 | 0.058701 |  | 4.5323 | -2.2897
7 | NO | 188 | 1.5906 | 0.0026761 | -0.86338 | -0.84137 | 0.022011 |  | 4.5289 | -2.2981
7 | NO | 189 | 1.5897 | 0.002941 | -0.85335 | -0.84088 | 0.012467 |  | 0.19877 | -0.10423
7 | NO | 190 | 1.2279 | 0.00078071 | -0.68212 | -0.65974 | 0.022386 |  | 0.21537 | -0.096628
7 | NO | 191 | 0.79214 | 0.0074926 | -0.42386 | -0.44155 | -0.017695 |  | 0.34909 | -0.19291
7 | NO | 192 | 0.34511 | 0.0023587 | -0.16389 | -0.21772 | -0.053823 |  | 0.62656 | -0.39444
7 | YES | 193 | 0.22352 | 0.0025103 | -0.084479 | -0.15684 | -0.072356 |  | 1.3231 | -0.78195
7 | YES | 194 | 0.1986 | 0.0012006 | -0.068397 | -0.14436 | -0.075962 |  | 2.4287 | -1.3573
7 | YES | 195 | 0.22269 | 0.0019834 | -0.10731 | -0.15642 | -0.049104 |  | 3.0405 | -1.5525
7 | YES | 196 | 0.32331 | 0.0021089 | -0.14869 | -0.2068 | -0.058106 |  | 3.4124 | -1.8786
7 | YES | 197 | 0.80869 | 0.013204 | -0.49131 | -0.44984 | 0.041463 |  | 3.7705 | -2.0525
7 | YES | 198 | 1.2606 | 0.0018827 | -0.67841 | -0.67612 | 0.0022966 |  | 4.0585 | -2.1975
7 | YES | 199 | 1.6433 | 0.0019382 | -0.91867 | -0.86776 | 0.050907 |  | 4.1637 | -2.1823
7 | YES | 200 | 2.3511 | 0.0008104 | -1.2879 | -1.2222 | 0.065744 |  | 4.4143 | -2.3017
7 | YES | 201 | 2.604 | 0.0016973 | -1.211 | -1.3488 | -0.13771 |  | 4.3857 | -2.1963
7 | YES | 202 | 2.871 | 0.0032544 | -1.521 | -1.4825 | 0.03849 |  | 4.3946 | -2.2154
7 | YES | 203 | 3.3187 | 0.0023088 | -1.7093 | -1.7066 | 0.0026629 |  | 4.4029 | -2.2261
7 | YES | 204 | 4.4733 | 0.0024342 | -2.2897 | -2.2848 | 0.0049773 |  | 4.403 | -2.197
7 | YES | 205 | 4.5131 | 0.001646 | -2.2709 | -2.3047 | -0.033781 |  | 4.4072 | -2.2028
7 | YES | 206 | 4.5501 | 0.0016361 | -2.2669 | -2.3232 | -0.056308 |  | 0.21626 | -0.10074
7 | YES | 207 | 4.5613 | 0.001885 | -2.2247 | -2.3289 | -0.10418 |  | 0.21597 | -0.098033
7 | YES | 208 | 4.5683 | 0.0016469 | -2.2487 | -2.3324 | -0.083726 |  | 0.32793 | -0.17107
7 | YES | 209 | 4.5684 | 0.0017776 | -2.2656 | -2.3324 | -0.066801 |  | 0.61294 | -0.34906
8 | NO | 258 | 2.0893 | 0.0030343 | -1.2647 | -1.0911 | 0.17365 |  | 1.2254 | -0.68064
8 | NO | 259 | 2.0264 | 0.0015592 | -1.1266 | -1.0596 | 0.067088 |  | 1.8644 | -1.0426
8 | NO | 260 | 1.8585 | 0.0029751 | -1.0315 | -0.97552 | 0.055953 |  | 2.4838 | -1.4142
8 | NO | 261 | 1.6089 | 0.0018999 | -0.89612 | -0.85051 | 0.045612 |  | 3.0897 | -1.4913
8 | NO | 262 | 1.2276 | 0.0013806 | -0.67139 | -0.6596 | 0.011785 |  | 3.2801 | -1.7719
8 | NO | 263 | 0.8383 | 0.0023897 | -0.4547 | -0.46467 | -0.009964 |  | 3.4803 | -1.8517
8 | NO | 264 | 0.42333 | 0.001106 | -0.19967 | -0.25688 | -0.057215 |  | 3.857 | -2.063
8 | YES | 265 | 0.2852 | 0.00087197 | -0.1488 | -0.18772 | -0.038917 |  | 4.1701 | -2.1979
8 | YES | 266 | 0.24974 | 0.0018643 | -0.12826 | -0.16997 | -0.04171 |  | 4.3789 | -2.2531
8 | YES | 267 | 0.25752 | 0.0021451 | -0.11748 | -0.17386 | -0.056381 |  | 4.4604 | -2.2416
8 | YES | 268 | 0.30445 | 0.0013724 | -0.15055 | -0.19736 | -0.046816 |  | 4.462 | -2.284
8 | YES | 269 | 0.57207 | 0.0015692 | -0.35693 | -0.33136 | 0.025575 |  | 4.464 | -2.261
8 | YES | 270 | 0.87831 | 0.002141 | -0.52769 | -0.4847 | 0.042983 |  | 4.4673 | -2.2587
8 | YES | 271 | 1.2268 | 0.00079044 | -0.72222 | -0.65919 | 0.063036 |  | 0.21146 | -0.088537
8 | YES | 272 | 1.7213 | 0.005514 | -0.84868 | -0.90681 | -0.058135 |  | 0.17125 | -0.070752
8 | YES | 273 | 2.0763 | 0.0055284 | -1.2137 | -1.0845 | 0.1292 |  | 0.24225 | -0.25675
8 | YES | 274 | 2.5045 | 0.0017083 | -1.3835 | -1.299 | 0.084504 |  | 0.55191 | -0.30009
8 | YES | 275 | 3.4371 | 0.0025153 | -1.8549 | -1.7659 | 0.088993 |  | 2.1879 | -1.2091
8 | YES | 276 | 4.5215 | 0.0025791 | -2.3795 | -2.3089 | 0.070613 |  | 2.6173 | -1.4097
8 | YES | 277 | 4.5762 | 0.0024665 | -2.3988 | -2.3363 | 0.062553 |  | 3.102 | -1.515
8 | YES | 278 | 4.593 | 0.0017391 | -2.319 | -2.3447 | -0.025679 |  | 3.1672 | -1.7168
8 | YES | 279 | 4.6045 | 0.0015811 | -2.3495 | -2.3505 | -0.0010522 |  | 3.4144 | -1.8006
8 | YES | 280 | 4.609 | 0.0016146 | -2.357 | -2.3528 | 0.0042096 |  | 3.7742 | -2.0368
8 | YES | 281 | 4.6136 | 0.0015638 | -2.3403 | -2.3551 | -0.014709 |  | 4.3573 | -2.2717
9 | NO | 296 | 0.78446 | 0.0014434 | -0.48255 | -0.43771 | 0.044839 |  | 4.4359 | -2.2861
9 | NO | 297 | 0.34697 | 0.0020328 | -0.18203 | -0.21865 | -0.036615 |  | 4.4173 | -2.2247
9 | YES | 298 | 0.26829 | 0.0018772 | -0.16171 | -0.17925 | -0.017541 |  | 4.4224 | -2.2766
9 | YES | 299 | 0.25443 | 0.0023303 | -0.13557 | -0.17231 | -0.036742 |  | 4.4264 | -2.2626
9 | YES | 300 | 0.32993 | 0.00072727 | -0.22807 | -0.21012 | 0.017947 |  | 4.4256 | -2.2524
9 | YES | 301 | 0.45436 | 0.0007746 | -0.31764 | -0.27242 | 0.045213 |  | 0.22221 | -0.10179
9 | YES | 302 | 0.74686 | 0.0017287 | -0.44714 | -0.41888 | 0.02826 |  | 0.15563 | -0.058372
9 | YES | 303 | 1.2119 | 0.001771 | -0.73312 | -0.65173 | 0.08139 |  | 0.15604 | -0.11096
9 | YES | 304 | 1.807 | 0.003248 | -1.002 | -0.9497 | 0.052316 |  | 0.31112 | -0.21788
9 | YES | 305 | 2.7778 | 0.0023032 | -1.5052 | -1.4358 | 0.069462 |  | 1.0677 | -0.64533
9 | YES | 306 | 3.1532 | 0.0023219 | -1.7338 | -1.6238 | 0.10999 |  | 1.9182 | -1.0978
9 | YES | 307 | 3.6766 | 0.0046223 | -1.9894 | -1.8858 | 0.10362 |  | 2.2928 | -1.2942
9 | YES | 308 | 4.3459 | 0.0022517 | -2.2831 | -2.221 | 0.062139 |  | 2.8803 | -1.5737
9 | YES | 309 | 4.4501 | 0.013204 | -2.2789 | -2.2732 | 0.005749 |  | 3.1247 | -1.4983
9 | NO | 310 | 4.5979 | 0.0017311 | -2.0531 | -2.3472 | -0.29406 |  | 2.9449 | -1.5191
9 | YES | 311 | 4.6012 | 0.0015276 | -2.2848 | -2.3488 | -0.063995 |  | 3.5627 | -1.9213
9 | YES | 312 | 4.6052 | 0.0019398 | -2.2878 | -2.3508 | -0.063073 |  | 4.2226 | -2.2354
9 | YES | 313 | 4.6101 | 0.001881 | -2.2749 | -2.3533 | -0.078321 |  | 4.3244 | -2.1936
9 | YES | 314 | 4.6127 | 0.001021 | -2.2713 | -2.3546 | -0.083223 |  | 4.3228 | -2.2862
10 | NO | 315 | 2.0817 | 0.0024677 | -1.1033 | -1.0872 | 0.016083 |  | 4.3595 | -2.1795
10 | NO | 316 | 2.0068 | 0.0014182 | -1.2132 | -1.0498 | 0.1634 |  | 4.3625 | -2.2125
10 | NO | 317 | 1.8758 | 0.0020761 | -1.1142 | -0.98416 | 0.13004 |  | 0.22507 | -0.13093
10 | NO | 318 | 1.5924 | 0.0025964 | -1.0156 | -0.84226 | 0.17334 |  | 0.16955 | -0.10345
10 | NO | 319 | 1.2051 | 0.002893 | -0.71593 | -0.64832 | 0.06761 |  | 0.19381 | -0.12119
10 | NO | 320 | 0.7649 | 0.0026344 | -0.4561 | -0.42792 | 0.028184 |  | 0.34427 | -0.20373
10 | NO | 321 | 0.329 | 0 | -0.182 | -0.20965 | -0.027651 |  | 1.1684 | -0.66156
10 | YES | 322 | 0.23434 | 0.0014919 | -0.16666 | -0.16225 | 0.0044085 |  | 1.5772 | -0.87577
10 | YES | 323 | 0.22027 | 0.0015492 | -0.18773 | -0.15521 | 0.032518 |  | 2.0046 | -1.1463
10 | YES | 324 | 0.29217 | 0.00079945 | -0.21784 | -0.19121 | 0.026628 |  | 2.6787 | -1.4824
10 | YES | 325 | 0.45698 | 0.0017534 | -0.27803 | -0.27373 | 0.0042943 |  | 3.0428 | -1.6392
10 | YES | 326 | 0.8295 | 0.00059335 | -0.4945 | -0.46026 | 0.034233 |  | 3.2226 | -1.7454
10 | YES | 327 | 1.2576 | 0.0043092 | -0.75237 | -0.67463 | 0.077738 |  | 3.4016 | -1.7734
10 | YES | 328 | 1.7763 | 0.0030435 | -1.0257 | -0.93432 | 0.091418 |  | 4.3486 | -2.2294
10 | YES | 329 | 3.0289 | 0.0020587 | -1.5291 | -1.5616 | -0.032504 |  | 4.4121 | -2.2349
10 | NO | 330 | 3.6728 | 0.0047478 | -2.0452 | -1.8839 | 0.16127 |  | 4.428 | -2.269
10 | YES | 331 | 4.1619 | 0.0033134 | -2.2591 | -2.1289 | 0.1302 |  | 4.4291 | -2.2529
10 | YES | 332 | 4.3362 | 0.0065157 | -2.2598 | -2.2161 | 0.043646 |  | 4.4308 | -2.2362
10 | YES | 333 | 4.4581 | 0.00087248 | -2.3028 | -2.2772 | 0.025653 |  | 4.4309 | -2.3651
10 | YES | 334 | 4.5042 | 0.0016126 | -2.2638 | -2.3002 | -0.036425 |  | 0.2794 | -0.1746
10 | YES | 335 | 4.5157 | 0.0013758 | -2.2843 | -2.306 | -0.021653 |  | 0.16799 | -0.075008
10 | YES | 336 | 4.5275 | 0.0016689 | -2.2945 | -2.3119 | -0.017392 |  | 0.16296 | -0.11404
10 | YES | 337 | 4.5323 | 0.0015943 | -2.2897 | -2.3143 | -0.024565 |  | 0.35181 | -0.18919
10 | YES | 338 | 4.5289 | 0.0016854 | -2.2981 | -2.3126 | -0.014523 |  | 0.98201 | -0.56999
11 | NO | 340 | 2.0236 | 0.0043433 | -1.1954 | -1.0582 | 0.13717 |  | 1.5411 | -0.87488
11 | NO | 0 | 0 | 0 | -3.175 | -0.044914 | 3.1301 |  | 1.8523 | -1.1046
11 | NO | 342 | 1.8446 | 0.001982 | -1.0723 | -0.96857 | 0.10378 |  | 2.4955 | -1.3585
11 | NO | 343 | 1.5831 | 0.002152 | -0.91895 | -0.83758 | 0.081373 |  | 2.749 | -1.456
11 | NO | 344 | 1.1926 | 0.0022141 | -0.67337 | -0.64209 | 0.031285 |  | 3.1219 | -1.6541
11 | NO | 345 | 0.78359 | 0.0022718 | -0.45441 | -0.43727 | 0.017141 |  | 4.3425 | -2.2525
11 | NO | 346 | 0.30934 | 0.0017008 | -0.15266 | -0.19981 | -0.047145 |  | 4.3829 | -2.2031
11 | YES | 347 | 0.19877 | 0.00093773 | -0.10423 | -0.14444 | -0.040211 |  | 4.4097 | -2.2173
11 | YES | 348 | 0.21537 | 0.0014613 | -0.096628 | -0.15276 | -0.056127 |  | 4.4159 | -2.2171
11 | YES | 349 | 0.34909 | 0.00249 | -0.19291 | -0.21971 | -0.026802 |  | 4.4178 | -2.2572
11 | YES | 350 | 0.62656 | 0.0032375 | -0.39444 | -0.35865 | 0.035792 |  | 0.26439 | -0.13361
11 | YES | 351 | 1.3231 | 0.0017168 | -0.78195 | -0.70739 | 0.074559 |  | 0.1841 | -0.1299
11 | YES | 352 | 2.4287 | 0.0014037 | -1.3573 | -1.261 | 0.096258 |  | 0.14844 | -0.093562
11 | YES | 353 | 3.0405 | 0.003679 | -1.5525 | -1.5673 | -0.014837 |  | 0.24249 | -0.13651
11 | YES | 354 | 3.4124 | 0.0028862 | -1.8786 | -1.7536 | 0.12499 |  | 0.84912 | -0.46688
11 | YES | 355 | 3.7705 | 0.0031573 | -2.0525 | -1.9329 | 0.11962 |  | 1.3692 | -0.74576
11 | YES | 356 | 4.0585 | 0.0032993 | -2.1975 | -2.0771 | 0.12045 |  | 1.8068 | -0.9932
11 | YES | 357 | 4.1637 | 0.0017177 | -2.1823 | -2.1298 | 0.052555 |  | 2.2439 | -1.2601
11 | YES | 358 | 4.4143 | 0.0018895 | -2.3017 | -2.2552 | 0.04649 |  | 2.2014 | -1.1856
11 | YES | 359 | 4.3857 | 0.0012572 | -2.1963 | -2.2409 | -0.044605 |  | 2.2263 | -1.1637
11 | YES | 360 | 4.3946 | 0.0015036 | -2.2154 | -2.2454 | -0.029961 |  | 2.521 | -1.348
11 | YES | 361 | 4.4029 | 0.0022066 | -2.2261 | -2.2495 | -0.023477 |  | 3.0116 | -1.5634
11 | YES | 362 | 4.403 | 0.0018797 | -2.197 | -2.2496 | -0.052612 |  | 4.3769 | -2.1901
11 | YES | 363 | 4.4072 | 0.001711 | -2.2028 | -2.2517 | -0.04884 |  | 4.414 | -2.225
12 | NO | 366 | 1.9875 | 0.0056304 | -1.2765 | -1.0401 | 0.23641 |  | 4.4292 | -2.1828
12 | NO | 367 | 1.977 | 0.0031318 | -1.148 | -1.0348 | 0.11318 |  | 4.4269 | -2.1911
12 | NO | 368 | 1.844 | 0 | -1.079 | -0.96824 | 0.11076 |  | 0.26006 | -0.12394
12 | NO | 369 | 1.5765 | 0.0026016 | -0.92051 | -0.83429 | 0.086218 |  | 0.15902 | -0.058984
12 | NO | 370 | 1.2116 | 0.0013283 | -0.68736 | -0.65161 | 0.035749 |  | 0.18711 | -0.11489
12 | NO | 371 | 0.79838 | 0.0016546 | -0.44562 | -0.44468 | 0.00094118 |  | 0.50397 | -0.30103
12 | NO | 372 | 0.32808 | 0.00064014 | -0.16092 | -0.20919 | -0.048275 |  | 1.0882 | -0.61578
12 | YES | 373 | 0.21626 | 0.0017867 | -0.10074 | -0.1532 | -0.052454 |  | 1.2397 | -0.65826
12 | YES | 374 | 0.21597 | 0.00036364 | -0.098033 | -0.15305 | -0.05502 |  | 1.3259 | -0.68814
12 | YES | 375 | 0.32793 | 0.00051211 | -0.17107 | -0.20912 | -0.038051 |  | 1.4877 | -0.80225
12 | YES | 376 | 0.61294 | 0.0017715 | -0.34906 | -0.35183 | -0.0027684 |  | 1.5429 | -0.80311
12 | YES | 377 | 1.2254 | 0.0030304 | -0.68064 | -0.65848 | 0.02216 |  | 1.6147 | -0.8403
12 | YES | 378 | 1.8644 | 0.0019371 | -1.0426 | -0.97844 | 0.064189 |  | 1.7423 | -0.90175
12 | YES | 379 | 2.4838 | 0.0092709 | -1.4142 | -1.2886 | 0.12561 |  | 1.9836 | -1.0214
12 | YES | 380 | 3.0897 | 0.0023932 | -1.4913 | -1.592 | -0.10069 |  | 2.3543 | -1.2107
12 | YES | 381 | 3.2801 | 0.0019439 | -1.7719 | -1.6873 | 0.084621 |  | 2.8475 | -1.4535
12 | YES | 382 | 3.4803 | 0.0025891 | -1.8517 | -1.7876 | 0.064162 |  | 4.4148 | -2.2462
12 | YES | 383 | 3.857 | 0.009024 | -2.063 | -1.9762 | 0.086766 |  | 4.4787 | -2.3363
12 | YES | 384 | 4.1701 | 0.0022345 | -2.1979 | -2.133 | 0.064965 |  | 4.5417 | -2.2833
12 | YES | 385 | 4.3789 | 0.002038 | -2.2531 | -2.2375 | 0.015615
12 | YES | 386 | 4.4604 | 0.0012024 | -2.2416 | -2.2783 | -0.036753
12 | YES | 387 | 4.462 | 0.0016602 | -2.284 | -2.2791 | 0.0049354
12 | YES | 388 | 4.464 | 0.0003873 | -2.261 | -2.2801 | -0.019126
12 | YES | 389 | 4.4673 | 0.0012418 | -2.2587 | -2.2818 | -0.023018
13 | NO | 392 | 1.9142 | 0.0022484 | -1.1478 | -1.0034 | 0.14442
13 | NO | 393 | 1.8954 | 0.0052708 | -1.1126 | -0.994 | 0.11856
13 | NO | 394 | 1.7689 | 0.0029451 | -1.0651 | -0.93064 | 0.13446
13 | NO | 395 | 1.5056 | 0.0026573 | -0.90141 | -0.79879 | 0.10262
13 | NO | 396 | 1.1867 | 0.0025 | -0.71727 | -0.63913 | 0.07814
13 | NO | 397 | 0.82421 | 0.00099931 | -0.46879 | -0.45761 | 0.011182
13 | NO | 398 | 0.34704 | 0.00045454 | -0.15896 | -0.21868 | -0.059726
13 | YES | 399 | 0.21146 | 0.0015548 | -0.088537 | -0.1508 | -0.062261
13 | YES | 400 | 0.17125 | 0.0019928 | -0.070752 | -0.13066 | -0.05991
13 | YES | 401 | 0.24225 | 0.0024334 | -0.25675 | -0.16621 | 0.090539
13 | YES | 402 | 0.55191 | 0.0017888 | -0.30009 | -0.32127 | -0.021175
13 | NO | 403 | 1.3935 | 0.0015659 | -0.94149 | -0.74267 | 0.19882
13 | YES | 404 | 2.1879 | 0.0017008 | -1.2091 | -1.1404 | 0.068692
13 | YES | 405 | 2.6173 | 0.0019649 | -1.4097 | -1.3554 | 0.054253
13 | YES | 406 | 3.102 | 0.001981 | -1.515 | -1.5981 | -0.083116
13 | YES | 407 | 3.1672 | 0.0012974 | -1.7168 | -1.6308 | 0.086037
13 | YES | 408 | 3.4144 | 0.0037168 | -1.8006 | -1.7546 | 0.045969
13 | YES | 409 | 3.7742 | 0.005372 | -2.0368 | -1.9347 | 0.10204
13 | YES | 410 | 4.3573 | 0.0032074 | -2.2717 | -2.2267 | 0.045031
13 | YES | 411 | 4.4359 | 0.00062452 | -2.2861 | -2.2661 | 0.020044
13 | YES | 412 | 4.4173 | 0.0018245 | -2.2247 | -2.2568 | -0.032117
13 | YES | 413 | 4.4224 | 0.0011819 | -2.2766 | -2.2593 | 0.017259
13 | YES | 414 | 4.4264 | 0.0011142 | -2.2626 | -2.2613 | 0.001301
13 | YES | 415 | 4.4256 | 0.001617 | -2.2524 | -2.2609 | -0.0084384
14 | NO | 418 | 1.6874 | 0.0018858 | -1.0436 | -0.88982 | 0.15379
14 | NO | 419 | 1.4561 | 0.00054546 | -0.84295 | -0.77399 | 0.068963
14 | NO | 420 | 1.2162 | 0.0012845 | -0.69182 | -0.65388 | 0.037939
14 | NO | 421 | 0.87144 | 0.0031699 | -0.47956 | -0.48126 | -0.0016984
14 | NO | 422 | 0.4072 | 0.00087197 | -0.2328 | -0.24881 | -0.016004
14 | YES | 423 | 0.22221 | 0.0026106 | -0.10179 | -0.15618 | -0.054385
14 | YES | 424 | 0.15563 | 0.0017895 | -0.058372 | -0.12284 | -0.064468
14 | YES | 425 | 0.15604 | 0.0017146 | -0.11096 | -0.12305 | -0.012088
14 | YES | 426 | 0.31112 | 0.0017088 | -0.21788 | -0.2007 | 0.017188
14 | YES | 427 | 1.0677 | 0.0033176 | -0.64533 | -0.57952 | 0.065813
14 | YES | 428 | 1.9182 | 0.0032042 | -1.0978 | -1.0054 | 0.092436
14 | YES | 429 | 2.2928 | 0.00092633 | -1.2942 | -1.193 | 0.10127
14 | YES | 430 | 2.8803 | 0.002273 | -1.5737 | -1.4871 | 0.086608
14 | YES | 431 | 3.1247 | 0.0020064 | -1.4983 | -1.6095 | -0.11127
14 | YES | 432 | 2.9449 | 0.0018055 | -1.5191 | -1.5195 | -0.00035312
14 | YES | 433 | 3.5627 | 0.002504 | -1.9213 | -1.8288 | 0.092502
14 | YES | 434 | 4.2226 | 0.0023223 | -2.2354 | -2.1593 | 0.076148
14 | YES | 435 | 4.3244 | 0.0017313 | -2.1936 | -2.2102 | -0.016671
14 | YES | 436 | 4.3228 | 0.0020977 | -2.2862 | -2.2094 | 0.076836
14 | NO | 437 | 4.3443 | 0.002411 | -2.6577 | -2.2202 | 0.43757
14 | YES | 438 | 4.3595 | 0.0029918 | -2.1795 | -2.2278 | -0.048256
14 | YES | 439 | 4.3625 | 0.0029554 | -2.2125 | -2.2293 | -0.016758
15 | NO | 441 | 0.84788 | 0.0020904 | -1.2821 | -0.46947 | 0.81265
15 | NO | 442 | 0.38846 | 0.0015548 | -0.21554 | -0.23943 | -0.023888
15 | YES | 443 | 0.22507 | 0.00072727 | -0.13093 | -0.15761 | -0.026675
15 | YES | 444 | 0.16955 | 0.0019577 | -0.10345 | -0.12981 | -0.026367
15 | YES | 445 | 0.19381 | 0.0018363 | -0.12119 | -0.14196 | -0.020769
15 | YES | 446 | 0.34427 | 0.0015492 | -0.20373 | -0.2173 | -0.013572
15 | YES | 447 | 1.1684 | 0.0015963 | -0.66156 | -0.62998 | 0.031587
15 | YES | 448 | 1.5772 | 0.0039386 | -0.87577 | -0.83466 | 0.041105
15 | YES | 449 | 2.0046 | 0.002128 | -1.1463 | -1.0487 | 0.097668
15 | YES | 450 | 2.6787 | 0.0081011 | -1.4824 | -1.3862 | 0.096183
15 | YES | 451 | 3.0428 | 0.0013396 | -1.6392 | -1.5685 | 0.070651
15 | YES | 452 | 3.2226 | 0.0019595 | -1.7454 | -1.6585 | 0.086852
15 | YES | 453 | 3.4016 | 0.0017677 | -1.7734 | -1.7482 | 0.025238
15 | YES | 454 | 4.3486 | 0.0013157 | -2.2294 | -2.2224 | 0.0069969
15 | YES | 455 | 4.4121 | 0.00081207 | -2.2349 | -2.2541 | -0.019269
15 | YES | 456 | 4.428 | 0.0012676 | -2.269 | -2.2621 | 0.0068398
15 | YES | 457 | 4.4291 | 0.0013817 | -2.2529 | -2.2627 | -0.009781
15 | YES | 458 | 4.4308 | 0.0022261 | -2.2362 | -2.2635 | -0.027317
15 | YES | 459 | 4.4309 | 0.00048819 | -2.3651 | -2.2636 | 0.10149
16 | NO | 461 | 1.1895 | 0.0015007 | -0.74349 | -0.64053 | 0.10296
16 | NO | 462 | 0.90289 | 0.0024793 | -0.52011 | -0.49701 | 0.023096
16 | NO | 463 | 0.43963 | 0.0013174 | -0.31237 | -0.26504 | 0.047327
16 | YES | 464 | 0.2794 | 0.0012005 | -0.1746 | -0.18481 | -0.010211
16 | YES | 465 | 0.16799 | 0.00058446 | -0.075008 | -0.12903 | -0.054023
16 | NO | 0 | 0 | 0 | -1.383 | -0.044914 | 1.3381
16 | YES | 466 | 0.16296 | 0.0017146 | -0.11404 | -0.12651 | -0.01247
16 | YES | 467 | 0.35181 | 0.0019075 | -0.18919 | -0.22107 | -0.031883
16 | YES | 468 | 0.98201 | 0.0017344 | -0.56999 | -0.53662 | 0.033367
16 | YES | 469 | 1.5411 | 0.0016861 | -0.87488 | -0.81658 | 0.058293
16 | YES | 470 | 1.8523 | 0.0019439 | -1.1046 | -0.97242 | 0.13223
16 | YES | 471 | 2.4955 | 0.0034907 | -1.3585 | -1.2945 | 0.064055
16 | YES | 472 | 2.749 | 0.001668 | -1.456 | -1.4214 | 0.034563
16 | NO | 473 | 2.8798 | 0.0018226 | -1.6382 | -1.4869 | 0.1513
16 | YES | 474 | 3.1219 | 0.0013291 | -1.6541 | -1.6081 | 0.045975
16 | NO | 475 | 3.5725 | 0.0039436 | -2.0075 | -1.8337 | 0.17378
16 | YES | 476 | 4.3425 | 0.0017893 | -2.2525 | -2.2193 | 0.033211
16 | YES | 477 | 4.3829 | 0.0013355 | -2.2031 | -2.2395 | -0.036433
16 | YES | 478 | 4.4097 | 0.0021378 | -2.2173 | -2.2529 | -0.035607
16 | YES | 479 | 4.4159 | 0.0015077 | -2.2171 | -2.256 | -0.038941
16 | YES | 480 | 4.4178 | 0.0017766 | -2.2572 | -2.257 | 0.00016216
17 | NO | 483 | 1.1427 | 0.0025483 | -0.7003 | -0.61709 | 0.083211
17 | NO | 484 | 1.1621 | 0.00090911 | -0.67292 | -0.62679 | 0.046127
17 | NO | 485 | 0.8981 | 0.0024372 | -0.5019 | -0.49461 | 0.0072909
17 | NO | 486 | 0.45113 | 0.0021831 | -0.24887 | -0.2708 | -0.021937
17 | YES | 487 | 0.26439 | 0.0015512 | -0.13361 | -0.1773 | -0.043687
17 | YES | 488 | 0.1841 | 0.00062456 | -0.1299 | -0.1371 | -0.0071954
17 | YES | 489 | 0.14844 | 0.0013958 | -0.093562 | -0.11924 | -0.025678
17 | YES | 490 | 0.24249 | 0.0019413 | -0.13651 | -0.16633 | -0.029821
17 | YES | 491 | 0.84912 | 0.0018054 | -0.46688 | -0.47009 | -0.0032113
17 | YES | 492 | 1.3692 | 0.0009489 | -0.74576 | -0.73052 | 0.01524
17 | YES | 493 | 1.8068 | 0.0018557 | -0.9932 | -0.94961 | 0.043584
17 | YES | 494 | 2.2439 | 0.0033862 | -1.2601 | -1.1685 | 0.091606
17 | YES | 495 | 2.2014 | 0.0032068 | -1.1856 | -1.1472 | 0.038462
17 | YES | 496 | 2.2263 | 0.0069041 | -1.1637 | -1.1597 | 0.0040189
17 | YES | 497 | 2.521 | 0.0025478 | -1.348 | -1.3072 | 0.040802
17 | YES | 498 | 3.0116 | 0.0016894 | -1.5634 | -1.5529 | 0.010489
17 | NO | 499 | 3.5302 | 0.0068678 | -1.9608 | -1.8125 | 0.14831
17 | YES | 500 | 4.3769 | 0.0019192 | -2.1901 | -2.2365 | -0.046428
17 | YES | 501 | 4.414 | 0.001449 | -2.225 | -2.2551 | -0.03009
17 | YES | 502 | 4.4292 | 0.0009602 | -2.1828 | -2.2627 | -0.079886
17 | YES | 503 | 4.4269 | 0.0010634 | -2.1911 | -2.2616 | -0.070524
18 | NO | 506 | 0.69494 | 0.00077779 | -0.58406 | -0.39289 | 0.19117
18 | NO | 507 | 0.44155 | 0.0014434 | -0.42645 | -0.266 | 0.16045
18 | YES | 508 | 0.26006 | 0.0017715 | -0.12394 | -0.17513 | -0.051189
18 | YES | 509 | 0.15902 | 0.0018529 | -0.058984 | -0.12454 | -0.065554
18 | YES | 510 | 0.18711 | 0.0016723 | -0.11489 | -0.1386 | -0.02371
18 | YES | 511 | 0.50397 | 0.00036364 | -0.30103 | -0.29726 | 0.0037725
18 | YES | 512 | 1.0882 | 0.0018506 | -0.61578 | -0.58981 | 0.02597
18 | YES | 513 | 1.2397 | 0.00099809 | -0.65826 | -0.66568 | -0.0074114
18 | YES | 514 | 1.3259 | 0.00064942 | -0.68814 | -0.7088 | -0.020657
18 | YES | 515 | 1.4877 | 0.0026054 | -0.80225 | -0.78986 | 0.012388
18 | YES | 516 | 1.5429 | 0.0016723 | -0.80311 | -0.81747 | -0.014362
18 | YES | 517 | 1.6147 | 0.0017061 | -0.8403 | -0.85343 | -0.013129
18 | YES | 518 | 1.7423 | 0.0015614 | -0.90175 | -0.91729 | -0.015541
18 | YES | 519 | 1.9836 | 0.0018574 | -1.0214 | -1.0382 | -0.016802
18 | YES | 520 | 2.3543 | 0.0017361 | -1.2107 | -1.2238 | -0.013073
18 | YES | 521 | 2.8475 | 0.0021951 | -1.4535 | -1.4707 | -0.017198
18 | YES | 522 | 4.4148 | 0.0016083 | -2.2462 | -2.2555 | -0.0092757
18 | YES | 523 | 4.4787 | 0.0017787 | -2.3363 | -2.2875 | 0.048828
18 | YES | 524 | 4.5417 | 0.0016444 | -2.2833 | -2.319 | -0.035702
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. CTD DOX
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_5
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=CTD DOX
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.9855 |  | X | CALC Y
m= | -0.35587 | 0.017621 |  | Std Dev of Resid |  | 0.083938 |  | 0.04676 | 0.023588
b= | 0.040229 | 0.05038 |  | Resid Sum of Squares |  | 0.042274 |  | 4.2193 | -1.4613
# of Values Extracted = |  |  | 8
# of Values Used in Fit = |  |  | 8
File | Use | Sample # | CTD DOX | CTD DOX | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | CTD DOX
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 4 | 0.048 | 0 | 0.004 | 0.023147 | 0.019147 |  | 0.048 | 0.004
1 | YES | 5 | 0.04676 | 0.0014834 | -0.05324 | 0.023588 | 0.076828 |  | 0.04676 | -0.05324
1 | YES | 6 | 0.50074 | 0.0087174 | -0.0012562 | -0.13797 | -0.13671 |  | 0.50074 | -0.0012562
1 | YES | 7 | 2.8522 | 0.0043797 | -1.0448 | -0.97478 | 0.070006 |  | 2.8522 | -1.0448
1 | YES | 8 | 3.2522 | 0.002912 | -1.1088 | -1.1171 | -0.008369 |  | 3.2522 | -1.1088
1 | YES | 9 | 3.687 | 0.0021289 | -1.264 | -1.2718 | -0.0078194 |  | 3.687 | -1.264
1 | YES | 10 | 3.877 | 0.013065 | -1.411 | -1.3395 | 0.071552 |  | 3.877 | -1.411
1 | YES | 11 | 4.2193 | 0.0032873 | -1.3767 | -1.4613 | -0.08463 |  | 4.2193 | -1.3767
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. CTD DOX
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_6
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=CTD DOX
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.7575 |  | X | CALC Y
m= | -0.11125 | 0.025697 |  | Std Dev of Resid |  | 0.025882 |  | 3.4827 | -1.352
b= | -0.96457 | 0.11206 |  | Resid Sum of Squares |  | 0.0040193 |  | 4.6928 | -1.4866
# of Values Extracted = |  |  | 8
# of Values Used in Fit = |  |  | 8
File | Use | Sample # | CTD DOX | CTD DOX | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | CTD DOX
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
2 | YES | 13 | 3.4827 | 0.0042317 | -1.3573 | -1.352 | 0.0052778 |  | 3.4827 | -1.3573
2 | YES | 14 | 4.2358 | 0.0015184 | -1.4682 | -1.4358 | 0.032375 |  | 4.2358 | -1.4682
2 | YES | 15 | 4.3923 | 0.0015558 | -1.4437 | -1.4532 | -0.0095243 |  | 4.3923 | -1.4437
2 | YES | 16 | 4.3895 | 0.0016078 | -1.4255 | -1.4529 | -0.027391 |  | 4.3895 | -1.4255
2 | YES | 17 | 4.3972 | 0.0035285 | -1.4218 | -1.4538 | -0.031992 |  | 4.3972 | -1.4218
2 | YES | 18 | 4.5156 | 0.0024659 | -1.4634 | -1.4669 | -0.0034856 |  | 4.5156 | -1.4634
2 | YES | 19 | 4.6644 | 0.012831 | -1.4856 | -1.4835 | 0.0021496 |  | 4.6644 | -1.4856
2 | YES | 20 | 4.6928 | 0.0017865 | -1.5192 | -1.4866 | 0.03259 |  | 4.6928 | -1.5192
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. CTD DOX
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_7
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=CTD DOX
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.99444 |  | X | CALC Y
m= | -0.36434 | 0.0068112 |  | Std Dev of Resid |  | 0.048841 |  | 0.1613 | -0.14929
b= | -0.090523 | 0.018225 |  | Resid Sum of Squares |  | 0.038166 |  | 4.5822 | -1.76
# of Values Extracted = |  |  | 18
# of Values Used in Fit = |  |  | 18
File | Use | Sample # | CTD DOX | CTD DOX | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | CTD DOX
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
3 | YES | 22 | 0.3693 | 0.0017447 | -0.1977 | -0.22507 | -0.027372 |  | 0.3693 | -0.1977
3 | YES | 23 | 0.21 | 0 | -0.099 | -0.16704 | -0.068035 |  | 0.21 | -0.099
3 | YES | 24 | 0.20993 | 0.00057973 | -0.12207 | -0.16701 | -0.044934 |  | 0.20993 | -0.12207
3 | YES | 25 | 0.16153 | 0.0013606 | -0.17247 | -0.14938 | 0.023096 |  | 0.16153 | -0.17247
3 | YES | 26 | 0.1613 | 0.0010539 | -0.2097 | -0.14929 | 0.060411 |  | 0.1613 | -0.2097
3 | YES | 27 | 0.4346 | 0.0012006 | -0.2234 | -0.24887 | -0.02547 |  | 0.4346 | -0.2234
3 | YES | 28 | 0.92265 | 0.0023447 | -0.40936 | -0.42668 | -0.017327 |  | 0.92265 | -0.40936
3 | YES | 29 | 1.3029 | 0.0011328 | -0.70209 | -0.56523 | 0.13686 |  | 1.3029 | -0.70209
3 | YES | 31 | 1.9264 | 0.0010691 | -0.7866 | -0.79239 | -0.0057973 |  | 1.9264 | -0.7866
3 | YES | 32 | 1.8416 | 0.0012006 | -0.7414 | -0.7615 | -0.020099 |  | 1.8416 | -0.7414
3 | YES | 33 | 2.0064 | 0.0013354 | -0.80164 | -0.82153 | -0.019889 |  | 2.0064 | -0.80164
3 | YES | 34 | 2.3332 | 0.0016834 | -0.94776 | -0.94062 | 0.0071387 |  | 2.3332 | -0.94776
3 | YES | 35 | 3.2813 | 0.0027519 | -1.3317 | -1.2861 | 0.045619 |  | 3.2813 | -1.3317
3 | YES | 36 | 4.0995 | 0.0030114 | -1.5895 | -1.5841 | 0.0053963 |  | 4.0995 | -1.5895
3 | YES | 37 | 4.4806 | 0.0012583 | -1.7234 | -1.723 | 0.00034996 |  | 4.4806 | -1.7234
3 | YES | 38 | 4.4338 | 0.0029481 | -1.7332 | -1.7059 | 0.027283 |  | 4.4338 | -1.7332
3 | YES | 39 | 4.5799 | 0.0018105 | -1.7302 | -1.7592 | -0.029006 |  | 4.5799 | -1.7302
3 | YES | 40 | 4.5822 | 0.001127 | -1.7118 | -1.76 | -0.048226 |  | 4.5822 | -1.7118
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. CTD DOX
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_8
Calculation of Linear Least Squares Fit For
(CTD DOX - Bottle DOX)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (CTD DOX - Bottle DOX), Prop=CTD DOX
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.97491 |  | X | CALC Y
m= | -0.45401 | 0.017166 |  | Std Dev of Resid |  | 0.12894 |  | 0.17374 | -0.13876
b= | -0.059882 | 0.047038 |  | Resid Sum of Squares |  | 0.29926 |  | 4.6054 | -2.1508
# of Values Extracted = |  |  | 21
# of Values Used in Fit = |  |  | 20
File | Use | Sample # | CTD DOX | CTD DOX | Diff (CTD DOX - Bottle DOX) |  |  |  | Fit vs. | CTD DOX
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
4 | YES | 58 | 1.3029 | 0.0022558 | -0.64814 | -0.6514 | -0.003259 |  | 1.3029 | -0.64814
4 | YES | 59 | 0.95083 | 0.0012803 | -0.53817 | -0.49157 | 0.04659 |  | 0.95083 | -0.53817
4 | YES | 60 | 0.478 | 0 | -0.218 | -0.2769 | -0.058901 |  | 0.478 | -0.218
4 | YES | 61 | 0.2704 | 0.0012005 | -0.1416 | -0.18265 | -0.041043 |  | 0.2704 | -0.1416
4 | YES | 62 | 0.18318 | 0.0023805 | -0.11782 | -0.14305 | -0.025231 |  | 0.18318 | -0.11782
4 | YES | 63 | 0.17374 | 0.00099807 | -0.10126 | -0.13876 | -0.037496 |  | 0.17374 | -0.10126
4 | YES | 64 | 0.33051 | 0.001674 | -0.24649 | -0.20994 | 0.036555 |  | 0.33051 | -0.24649
4 | YES | 65 | 0.33226 | 0.0014866 | -0.26674 | -0.21073 | 0.05601 |  | 0.33226 | -0.26674
4 | YES | 66 | 0.99696 | 0.0016852 | -0.54804 | -0.51252 | 0.035525 |  | 0.99696 | -0.54804
4 | YES | 67 | 1.6957 | 0.002548 | -0.86334 | -0.82974 | 0.033603 |  | 1.6957 | -0.86334
4 | YES | 68 | 2.0467 | 0.0023078 | -1.0023 | -0.9891 | 0.013245 |  | 2.0467 | -1.0023
4 | YES | 69 | 2.5483 | 0.0018423 | -1.2267 | -1.2168 | 0.0098532 |  | 2.5483 | -1.2267
4 | YES | 70 | 2.8469 | 0.0011569 | -1.3421 | -1.3524 | -0.010257 |  | 2.8469 | -1.3421
4 | YES | 71 | 3.0726 | 0.0018841 | -1.4324 | -1.4549 | -0.022545 |  | 3.0726 | -1.4324
4 | NO | 72 | 3.1943 | 0.0016615 | -3.2447 | -1.5101 | 1.7346 |  | 3.1975 | -1.4465
4 | YES | 73 | 3.1975 | 0.0013041 | -1.4465 | -1.5116 | -0.065151 |  | 4.4961 | -2.0579
4 | YES | 74 | 4.4961 | 0.0029635 | -2.0579 | -2.1012 | -0.04332 |  | 4.5756 | -2.6084
4 | YES | 75 | 4.5756 | 0.0016338 | -2.6084 | -2.1373 | 0.47109 |  | 4.5983 | -2.0307
4 | YES | 76 | 4.5983 | 0.0018676 | -2.0307 | -2.1476 | -0.1168 |  | 4.5999 | -2.0381
4 | YES | 77 | 4.5999 | 0.0019402 | -2.0381 | -2.1483 | -0.11028 |  | 4.6054 | -1.9826
4 | YES | 78 | 4.6054 | 0.0016601 | -1.9826 | -2.1508 | -0.16819
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
FIT vs. Sample #
FIT vs. CTD DOX
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE