## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 22 March 2021 | ### Added HPLC data. S.H. |
| 15 November 2005 | TSG Fluorescence recalibrated using manufacturers 2001 calibration |
| 5-Jul-2005 | Added nutrient data to the rosette files. J.L. |

## PROCESSING NOTES

Cruise: 2005-02

Agency: OSAP

Location: WCVI

Project: Optical Properties Off Vancouver Island

Party Chief: Peña A.

Platform: John P. Tully

Date: February 28, 2005 – March 5, 2005

Processed by: Germaine Gatien

Date of Processing: March 15, 2005 – June 7, 2005

Number of original CTD casts: 25

Number of CTD casts processed: 23 (2 files contained only surface data, not mentioned in log)

Number of rosette casts: 23

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with Chelsea/Seatech transmissometer (#333DR), Altimeter OA-916D (#1024), a Seapoint Fluorometer (#2229) with a 10X cable, Biospherical Instruments PAR sensor (#4656) and Biospherical Reference PAR (#16504). The deck unit was a model 911 (#0508) and the logging computer was FS03 (Shuttle Case). The salinometer was a Portasal model 8410 (#58879). A thermosalinograph (SeaBird 21, S/N 2487) was mounted with fluorometer WS3S-713P.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in excellent order.

There was no record in the log of when two of the loop samples were taken.

There are uncertainties about the salinity calibration. The analysis of bottles from both 2005-02 and 2005-01 were interrupted by a malfunction in the Portasal. Comparison between CTD and bottle samples run after repairs to the Portasal differ markedly from those done before. Recalibration was done based on the results from after repair, but there remains some doubt about the reliability of those results.

The salinity calibration sampling for casts #29 to 31, after the sensors were changed, is too scattered and too shallow to allow a useful comparison. No salinity recalibration was applied to those casts.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 200m

- ±0.1ml/l from 200db to 500db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Note was made in the log of problems with large differences between the conductivity sensors and problems with DO titrations.

The temperature and conductivity sensors were changed before cast #29.

File 2005-02chlarc.xls was obtained with extracted chlorophyll data including flags and comments.

Bottle salinity data was unavailable at the beginning of the processing job – only loop data was ready.

The dissolved oxygen data were obtained; flags and comments had been added. Two of the files had been recalibrated by the analyst. There are a few problems with the files that will need attention.

Chlorophyll data was available with flags and comments, but with headers with non-standard names and extra lines that needed to be removed.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no problems found.

There were two configurations used because the sensors were changed after cast #28.

Configuration file 2005-02-ctd1.con was prepared for casts #3 to 28 and 2005.02-ctd2 for casts 29-31.

##### Conversion of Raw Data

All data were converted.

A few casts were examined and all expected channels are present. The up and downcasts were mostly similar. The pairs of conductivity sensors are farther apart than usual before the sensors were changed. The altimetry is noisy but looks reasonable near the bottom. There are many spikes in T and S, mostly in the upcasts. PAR, SPAR, DO and transmissivity look ok.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format.

All files were put through CLEAN to add event numbers.

All BOT files were plotted. The only problem noted was noise in the primary temperature for bottle #15, cast #7. CTDEDIT was used to clean that data and the output was copied to BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5	Points per block = 50

##### ALIGNCTD

This step was skipped as the deck unit was one that advances the secondary conductivity.

##### CELLTM

Tests were run on the 3 casts (2 with the first configuration and 1 with the second) with (alpha, 1/beta) set to (0.01, 9), (0.02, 7), (0.03, 7), (0.02, 9), (0.0245, 9.5) and (0.03, 9). The best results were found with (0.02, 7) for both primary and secondary for the first two casts and for the primary of the third. For the secondary for cast #31 (0.01, 9) worked best.

CELLTM was run using (0.02, 7) for both channels for casts #1-28 and for the primary for casts #29-31. (0.01, 9) was used for the secondary for casts #29-31.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 250 | -0.0005 | -0.0011 | -0.011 |  |
| 3 | 450 | ~0 XN | -0.0011 | -0.0105 | High, moderate noise |
| 4 | 450 | -0.001 XN | -0.001 | -0.0105 | High |
| 4 | 600 | ~0 XN | -0.001 | -0.0105 | High, very noisy |
| 31 | 230 | -0.0015 XN | ~0 N | +0.0015 | High, moderate noise |

XN means the differences were extremely noisy.

The values for the first configuration are similar to those seen late in 2005-01 with the same equipment, but these casts are much shallower. During 2005-01 there was clearly pressure dependence in the conductivity and salinity at great depth. There were recent recalibrations for all the T and C sensors, so the large differences are unexpected.

For the second configuration we have no history of the same configuration. The differences in salinity look reasonably small.

One offshore cast was examined in detail to see if it was clear which temperature sensor was giving better data, but little was learned. While stopped, the primary temperature was consistently higher than the secondary. While moving there was no regular pattern. When the descent rate was very high they seemed closest. There are frequent excursions in one of the sensors while the other holds steady, but it happens to both. Shed wakes sometimes affected one more than the other, but again it could be either sensor. This was a rough cast so perhaps the CTD was swinging in a way that affected the pumps alternately.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run. One station name was added and 2 were corrected.

The cruise track was plotted and no further problems were noted.

The average surface pressure is 1.5db, which is a little lower than usual, but there is a wide range of values from 0.2 to 4.5db. The cast with the lowest surface reading was examined and values in the top 1db look like near-surface, but in-water values. The pressure sensor was recalibrated recently and appears to be close to correct so will not be adjusted further.

A few casts were checked to ensure the header got the right reading from the altimeter. There was a lot of noise, but the algorithm worked well.

The transmissivity and fluorometer traces look ok and upcasts and downcasts are reasonably similar.

##### BOTTLE FILE PREPARATION

The bottle files from casts #10 and 11 were joined and named 2005-02-0010.bot.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

Sample numbers were then added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

Dissolved oxygen files (*.add) created by the analyst had a flag channel. Problems were encountered in the analysis of two casts (9 and 13) and a recalibration was done (output: *.DOXX) so that there are two DO channels in the file. REMOVE was used to remove the first channel and REORDER was used to get the same order as for the rest of the files.

File 2005-02chlarc.xls was edited to change the channel names to standard form and event numbers were added; it was saved as 2005-02chlarc.csv.

Nutrient data were not yet available.

The CHL and ADD files were merged with SAMAVG in two steps. A format error in one of the DO flags had to be fixed to enable merging. (Output: MRG1, MRG2)

##### 12. COMPARE (except salinity)

Salinity – partial results from 2005-01

In the absence of salinity data a few casts from Line P were compared with the few bottles that were analyzed before the salinometer broke. The differences are unusually high for both channels and there is similar pressure dependence for both, but it differs from one cast to another. From the differences between the two channels described in section 8, we expect pressure-dependence in at least one channel and we expect time-dependence. But we also expect a large difference between the channels varying from +0.001 for cast #12 at 2000db to -0.0115 for cast #97 at the same pressure. COMPARE shows that for cast #6 the primary was lower than the secondary by about 0.0007psu and for cast #15 the primary is higher by about 0.0013spu. These results are consistent with expectations but insufficient to resolve which is better and what is creating such a large difference.

Dissolved Oxygen

COMPARE was run using sample number as the reference variable. The best fit was with the differences between bottles and CTD versus CTD DOX. When outliers were excluded by using standard deviation in the CTD data and residuals in the differences as indicators, the following fit was found:

CTD-BOT = 1.463 DOX-CTD + 0.111

During the two previous uses of this sensor the results were:

CTD-BOT = 1.493 DOX-CTD + 0.056 (2005-01)

CTD-BOT = 1.359 DOX-CTD + 0.069 (2004-37)

2004-37 was in the Juan de Fuca/Strait of Georgia area where the range of DO would be quite different. The results are close to those of 2005-01. The differences were fairly flat with time. The only severe outliers had already been flagged “d”. (See 2005-02-dox-comp1.xls.)

Plots of CTD DO and Titrated DO versus CTD Salinity from the bottle files were examined. For cast #15 there are a few records that look a little odd, but both the CTD and bottles show the same feature, and the range of values is quite small.

CHL vs fluorescence

COMPARE was used to plot the differences between FLUOR and CHL versus pressure and fluorescence. There was slight pressure dependence. The fit of CHL versus FL is as follows:

CHL = 1.145 * FL – 0.113

Removing more points based on standard deviations in the CTD data reduces the slope, but also removes all values with high FL values, so this is probably not a useful result. Removing points based on the fit residuals increases the slope; as the limit is gradually reduced the fit approaches something like

CHL = 1.22 * FL – 0.36.

The only severe outliers were for high CHL values near the surface; it has been noted frequently that the CTD fluorometer does not do a good job with such values, so there is no reason to think the CHL data should be flagged.

##### Other Comparisons

Previous experience with these sensors – The sensors used for casts #3 to 38 were used for 2004-37 and 2005-01. During 2004-37 the primary sensors were found to be high by 0.0035 and the secondary high by 0.0019psu. Doubts have arisen since the comparison was done, as it has been found that mud may have corrupted the samples for some of the bottom bottles. The differences between the two salinity channels varied from 0.001psu to 0.0025psu, so much smaller than during this cruise. During 2005-01 the primary salinity was found to be very time-dependent and the secondary was low by 0.0044psu.

The primary sensors used for casts #29-31 were freshly calibrated. The secondary sensors were used for 2004-26 when the salinity was very close to the bottles, but the salinity was very noisy.

Historic ranges –Cast #16 had low salinity near 80db. This cast was very close to shore, so may not be well represented in the historic range data.

##### 14. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this data set the fluorescence was low and the descent rate very noisy, so this calculation was not easy. Values were found from 0.7s to 1.5s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Dissolved Oxygen

A rough estimate was made using the technique described above for fluorescence and this suggests that an advancement of about 7s or 168 records is required to make the difference between the down and upcast DO look like that of the temperature. Tests were run using +120, +140, +160, +180, +200 records on a few casts. As well as making the trace offset similar to those of temperature we also judge by how the DO looks during a bottle stop. We expect the DO to settle to the correct value by the end of the bottle stop. That value should be roughly half way between the up and down values in motion. The results vary with DO gradient, so there is no clear RIGHT answer. However, the best choice overall appears to be to advance the DO channel by +180 records. This shift was applied to all casts.

Conductivity

Since there were two configurations of T and C sensors tests had to be done in two batches.

For the first configuration casts #7 and 13 were used since they had sufficient fine structure to test the alignment well. During 2004-37 when these sensors were used, the primary sensors were shifted by -0.8 records and the secondary by -0.5s, but the added pumped sensors may have been mounted in reversed positions. The best settings proved to be -0.7 for the primary and -0.4 for the secondary.

For the second configuration casts #29 and 31 were studied. During 2004-26 when these sensors were used the primary was shifted by -0.2s and the secondary by -0.4s. For this cruise the best choice was -0.2s for both the primary and secondary.

All casts were put through SHIFT using -0.7s/-0.4s for casts #3-28 and -0.2s/-0.2s for casts #29-31. (Output *.SHFC1 and SHFC2).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

##### DETAILED EDITING

Without bottle data it is not obvious which sensor pair to pick for the archive. The secondary was picked because those sensors were used during 2004-37 when the fluorometer was mounted on the primary pump. That is the case for this cruise too and there is a strong suspicion that fluorometers have a bad effect on the pump.

Page plots were produced using (T1,S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

Casts requiring heavy editing were: 3 & 4.

All other casts required light editing.

Note was made of the editing details in the headers of the relevant files.

##### Initial Calibration study

Shortly after the editing was completed post-cruise calibration information became available for the primary conductivity sensor used for casts 1-28 and the secondary temperature sensor used for casts 29-31.

The drift in the primary conductivity calibration suggests that the salinity was high by from 0.01psu to 0.015psu at the time of this cruise. This is close to the differences between the two sensors for casts 1-28, which implies that the error in the secondary salinity is reasonably small. The choice of the secondary salinity looks best.

The drift in the secondary temperature for casts 29-31 would lead to salinity values that are low by about 0.0015psu, but we have no idea what the error is due to conductivity drift. The differences showed that the secondary was higher than the primary by about 0.0015psu implying that the primary must be low by about 0.003psu. Once again, the secondary is probably the better choice.

No recalibration should be applied until bottle salinity analysis can be done. But there is reason to expect that the errors will be within ±0.005psu.

##### Initial Recalibration

Because the salinity data was unavailable and the data was needed by the PI, the DO channel was recalibrated. The salinity recalibration will be done later when the bottle data is available.

The DO channel was recalibrated in the bottle files using the relationship found in section 12:

CTD DOX (corrected) = 1.463 CTD-DOX + 0.111

COMPARE was rerun and when only points with -0.15<residuals <0.15 are included, the average difference is -0.004 ml/l. The edited downcast files, EDT, were also recalibrated.

##### Special Fluorometer Processing

The COR files were clipped to 100db and put through REMOVE and HEADEDIT to produce files PFCTD. The PFCTD files were written on a CD-ROM for the use of Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison

The averaged downcast files were thinned and the CTD DO values were compared with the upcast bottle DO values. The results show good correspondence near the surface. While there is a lot of scatter in the plot of differences versus DO, the average is near zero in the top 150m. Below that, there is a tendency for the CTD to read high by up to 0.15ml/l. This is probably due to deep water being underrepresented in the fit since there was little data from below 200db.  (See 2005-02-dox-comp3.xls)

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 200m

- ±0.15ml/l from 200db to 500db

##### FINAL CTD files steps (REMOVE and HEADEDIT) except for salinity recalibration

For all casts the following channels were removed: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Altimeter and Flag channels. (Output: *.REM)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and

unedited except that some records were removed in editing T and S

The dissolved oxygen data in the CTD files should be considered:

±0.5ml/l in the top 200m

±0.15ml/l from 200db to 500db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named PCTD and still need to have the salinity recalibrated.

##### Final Bottle Files except for salinity recalibration

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag for all casts.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named PCHE.

##### SALINITY COMPARISON

Salinity analysis data was received in May. The analyst noted that one sample was entered wrong in the computer as #501; this was renamed in the spreadsheet as #1. There was one duplicate sample, #25. An average was entered and the individual values recorded as a comment. There was a sample #933 which is not recorded in the Rosette Log nor the Daily Log; the value makes no sense as a duplicate of #93 and there is no sample missing. This value was removed from the file, which was saved as 2005-01-sal.csv.

That file was converted to individual SAL files.

The PCHE files were put through SORT to order on sample numbers and saved as PCHE2. These were merged with the SAL files and saved as PCHE3.

COMPARE was run using PCHE3 and SAMCOR1. The sampling was mostly very shallow and the scatter was very large, but similar to the results of 2005-01 for the same depths. When only the two samples below 275m are included, the primary sensors are high by about 0.006psu and the secondary are low by 0.0047psu. During 2005-01 (which immediately preceded this cruise) the secondary was found to be low by about 0.0044psu and steady with time. The primary was found to vary greatly with time being low by about 0.002psu early in the cruise and high by about 0.008psu by the end. There is no evidence of temporal drift during 2005-02, but the only deep cast was at the beginning of the cruise, so little can be deduced from that. The most we can say is that the results from this cruise are not out of line with those of 2005-01. So using the results of 2005-01 is the best plan.

Casts #29 to 31 had different sensors; there was salinity sampling, but none below 250m. The secondary salinity appears to be closer to the bottles than the primary, but there is so much scatter that it is impossible to make an estimate of offset.

##### SALINITY RECALIBRATION for casts #3-27

The salinity for casts #29, 30 and 31 will not be recalibrated; the PCTD, PBOF, PFCTD files were renamed as CTD, BOF and FCTD for those casts only. The PCHE3 files were put through sort again and renamed CHE.

The PCHE3 and SAMCOR1 files were recalibrated using file 2005-02-recal2.ccf to add 0.0044psu to the secondary salinity. The output files, PCHE4 and SAMCOR2, were used as input to a second run of COMPARE to ensure the calibration had been done correctly, and it had. Then PCHE4 was put through SORT to order by increasing pressure and the output was named CHE.

The PCTD, PBOF and PFCTD files were also recalibrated using file 2005-02-recal2.ccf. The output files were named CTD, BOF and FCTD, the latter two being for the use of Angelica Pena, not the archive.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T and S. Two sets of profile plots were produced with combinations of PAR, Surface PAR, DO, fluorescence and transmissivity profiles.

Cross-reference lists were produced for the final CTD and CHE files.

##### 25. Thermosalinograph Data

a.) Checking calibrations

There was 1 file containing TSG data. A report was printed for the con file and the calibrations were checked for temperature and conductivity. No errors were found. The primary  sensors were recalibrated in Dec. 2004, the secondary temperature in August 2004. When last calibrated the drift in salinity was found to be +0.0008psu per month and the temperature drift about -0.0012Cº. If drift has continued at the same rate, we would expect the temperature to be low by about -0.00025Cº (which would lead to a salinity value that is too high by only 0.0002psu) and a salinity error due to conductivity drift on the order of +0.002psu. However, the conductivity drift between the 2003 and 2004 calibrations was of the opposite sign, so we should not assume it will be the same for this deployment.

The TSG has been used for only 1 other cruise since recalibration.

b.) Converting to IOS Headers and adding position headers and time channels.

The data was converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

Note that UPloy0 was the actual channel name and that it contains the flow rate.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

Time-series plots were produced and no severe problems noted.

There is only one spike in each of the primary temperature and salinity channels. The secondary temperature shows no such spike so pad values were entered into the primary temperature and salinity channels for record #3877 of file 2004-25-0001.atc.

c.)  Checking Time Channel

The CTD data after DELETE and metre-averaging was thinned to reduce the files to a single point at or near 4db and exported to a spreadsheet. The CTD data was unedited. (2004-25CTD.csv)

The ATC file was opened in EXCEL and reduced to a single header line and the channels of interest (Date, Time, T0, T1, Sal, Flow, Fluor, Lat, Long). All records were removed EXCEPT those that correspond to the time of a CTD cast.

Two different files were derived, one studying the time, the other comparison T, S and FL. See section (e) for the latter comparison.

When latitude and longitude from the CTD and TSG were plotted against each other the fit was excellent. (See 2005-02-time-study.xls.)

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are no obvious problems this time, so this step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 For a first comparison the CNV files were opened in EXCEL. The differences between the two temperature channels are very noisy with a few quiet patches that probably correspond to stops. If all data is included the temperature from the lab is higher than the remote temperature by an average of 0.237Cº. When a fairly quiet 8-hour section was chosen the average was 0.182Cº. To get a good comparison cast #3 was chosen because there was a fairly well-mixed surface layer and little ship drift. During that stop the difference was 0.2Cº. During 2004-29 the differences were found to be 0.13Cº. The primary temperature was recalibrated recently, but not the intake sensor. It seems likely that the T1 vs T2 difference could vary by as much as 0.07Cº due to sensor drift and variable intake temperatures. During 2004-29 the surface water was around 13 to 17ºC whereas during 2005-02 it was around 9ºC.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. (2005-02-TSG-CTD-comp.xls)

The remote TSG temperature was higher by an average of 0.058Cº with the secondary CTD temperature being slightly closer to the TSG intake temperature. When a few outliers were excluded the salinity of the TSG was lower than the primary CTD salinity by about 0.03psu and lower than the secondary by 0.02psu. (This was done before recalibration of CTD salinity; after recalibration the TSG would be lower than the secondary CTD salinity by about 0.024psu.)

The same data was used to compare the fluorescence from the two instruments. The fluorescence from the TSG was very low, with CTD FL = 12.4 * TSG Fl - 1.6 with the CTD Fluorescence  being an average of 5.4 times the TSG Fluorescence.

- Loop Bottle Comparisons Next thing to check is the comparison of 18 loop salinity bottle samples and the TSG data. Unfortunately there is no record of when 2 of them were collected.

For 16 of the samples the loop bottle salinity values were compared with the TSG salinity. (See 2005-02-TSG-avg.xls.) The differences showed a lot of scatter. As a measure of TSG variability the difference between TSG salinity and the 2 minute-average TSG salinity was plotted against TSG variability. This showed that the variability in the TSG record was responsible for most of the scatter. When 4 records in high variability areas were excluded, the TSG was found to be lower than the bottles by 0.035psu. When the variability was very low the TSG was low by from 0.017psu and 0.027psu, with an average of 0.021psu. For the offshore part of the cruise, the low variability data corresponds to times when there were frequent stops for casts. The highest variability is in the Strait of Georgia.

Fluorescence values were compared with the loop chlorophyll values. (2005-02-TSG-Fl-chl-comp.xls) A fit of

CHL = 13.9*FL – 1.9ug/l

was found when the noisy TSG data is excluded. We expect a value of about 10. The results for other TULLY cruises using the same equipment were:

CHL = 9.7*FL – 2.2ug/l     (2004-25)

CHL = 9.3*FL – 0.15ug/l   (2004-29)

f.) Calibration History

The TSG was recalibrated in December 2004. It was used for 2005-01, during which there was a lot of noise in the salinity comparison, and concerns about the Portasal performance. The surface gradients were very low so vertical matching of TSG and CTD data was not critical. The temperature was found to be high by about 0.2Cº and the salinity low by about 0.03psu ±0.01psu.

Conclusions

Since the primary temperature sensor was recently calibrated, we expect good data. The temperature in the TSG is 0.21Cº higher than the intake temperature in areas of low variability. The intake temperature is higher than the CTD temperature by about 0.06 and the primary temperature is higher than the CTD by about 0.27Cº. The differences between the two TSG temperatures are close to those observed during 2005-01, but the differences from the CTD are larger. 2005-01 sampled waters with extremely low T and S gradients near the surface, so that comparison between TSG and CTD is presumed to be more useful.

The salinity is lower than the loop bottles by about 0.021psu when only data from low-variability regions are included. But based on results from 2005-01, it is believed that the Portasal analyses that were run in March 2005 were low by more than 0.01psu. This included the loop samples for this cruise. So the TSG salinity is probably low by at least 0.03psu. It is lower than the calibrated CTD secondary salinity by about 0.026.

The TSG Fluorescence is very noisy, but appears to be low by a factor of about 14.

f.) Editing

One record was replaced with pad values for temperature and salinity earlier in the processing; no further editing was done.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.20 Cº and +0.03psu to Temperature:Primary and Salinity:T0:C0. After calibration the two temperature channels were compared and are closer.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary, Flag, UPloy0 (flow rate)

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, add the depth of sampling to the header and change the fluorescence channel name to FLUORESCENCE and the corresponding units to VOLTS.

The sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

Particulars

1-12 Ship/cruise ID wrong in headers.

7. Salinity difference ~0.1psu due to conductivity sensor difference.

13. Problems in DO analysis.

29. Changed secondary sensors

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-02 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 28 February 2005 End: 5 March 2005 |  |  |  |  |  |  |
| Location: WCVI and Strait of Georgia |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Peña A. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-02

| Calibration Information – configuration for casts 1-28 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 18/11/04 | Factory “ |  |  |
| Conductivity | 2399 | 19/11/04 | “ |  |  |
| Secondary Temp. | 2106 | 18/11/04 | “ |  |  |
| Secondary Cond. | 1763 | 19/11/04 | “ |  |  |
| Transmissometer | 333DR | 29/11/04 | Factory |  |  |
| Biospherical PAR | 4656 | 11/02/03 |  |  |  |
| SBE 43 DO sensor | 0766 | 16/11/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| SPAR | 16504 | 02/01/04 |  |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |
| Calibration Information – configuration for casts 29-31 |  |  |  |  |  |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 16/10/04 | Factory “ |  |  |
| Conductivity | 2102 | 15/10/04 | “ |  |  |
| Secondary Temp. | 2023 | 14/10/04 | “ |  |  |
| Secondary Cond. | 1766 | 15/10/04 | “ |  |  |
| Transmissometer | 333DR | 29/11/04 | Factory |  |  |
| Biospherical PAR | 4656 | 11/02/03 |  |  |  |
| SBE 43 DO sensor | 0766 | 16/11/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| SPAR | 16504 | 02/01/04 |  |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2005-02

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 4/12/04 | Factory |  |  |
| Conductivity | 2487 | 4/12/04 | “ |  |  |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab WetstarFluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? |  |  |  |