## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2005-03

Agency: OSAP

Location: N. E. Pacific

Project: High Seas Salmon

Party Chief: Trudel M.

Platform: W. E. RICKER

Date: March 2, 2005 – March 28, 2005

Processed by: Germaine Gatien

Date of Processing: 21 April 2005 –8 August 2005

Number of original CTD casts: 127 (includes one test cast)

Number of casts processed: 126

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit was a SeaBird model 11 (S/N 0471.) The salinity was analyzed using an Autosal on loan from Guildline.

The thermosalinograph was not working.

# SUMMARY OF QUALITY AND CONCERNS

The CTD Daily Log Book was generally in good order.

The secondary conductivity was bad, probably due to a clogged bleed valve.

There was no Niskin sampling from cast #36 to #94 due to damage to the bottles.

The salinity comparison showed a lot of scatter; salinity is considered ±0.005psu.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and was in good order.

A few problems are noted in the log. There was no Niskin sampling from cast #36 to #94 due to damage to the bottles; a new one was brought to the ship and sampling resumed at cast #97.

Bottle salinity, nutrients and extracted chlorophyll data were obtained.

The configuration files were obtained and the calibration constants were checked. The pressure calibration date and one co-efficient were wrong in the pressure calibration and the date of the transmissivity calibration was wrong. A new configuration file was prepared with those errors corrected; it was named 2005-03-ctd.con.

The sensor history was found. During 2004-32 the primary sensor was found to be close to the bottles but the secondary was about 0.03psu lower than the primary and the bottles. It is possible that the problem was due to bleed valve clogging rather than calibration.

##### Conversion of Raw Data

The raw data were converted using configuration file 2005-03-CTD.con.

Examination of a few casts shows that all expected channels were present. There are large differences between the temperature and conductivity channels in both the downcast and upcast; the largest differences are in the conductivity in the top 50m of the downcast. The upcasts are quite different from the downcasts with a lot of noise in both channels. Examination on a T-S surface shows that the secondary salinity is bad in the top 50m. Based on previous experience with this equipment, it looks as though the bleed valve had clogged and cleared itself around 50db. This explains some of the differences, but even after that problem clears, there are still significant differences between these channels. Before cast #66 the bleed valve was cleared and thereafter the problems with the secondary data in the top 50db seem to disappear, but the large differences elsewhere persist.

The fluorescence has significant differences between downcast and upcast traces. The fluorescence could be related to pump problems as it was probably mounted on the secondary pump. (The log does not indicate which pump it was on, but the secondary is usually chosen.) The transmissivity seems ok.

Later in the processing it was found that cast #115 was corrupted. The file was repaired and reconverted with the name 2005-03-9115.cnv. It was then put through the steps described in §4 to §10. Once in IOS HEADERS it was renamed as 2005-03-0115.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2	Pass 2    Std Dev = 5     Points per block = 50

##### ALIGNCTD

ALIGNCTD was not run even though the deck unit was one of the older model since recent cruises suggest that even the secondary conductivity is being advanced. If this is not so, SHIFT can be used later to do this correction.

##### CELLTM

Tests were run on three casts running CELLTM with choices of (0.01,7), (0.01,9), (0.02,7), (0.02,9), (0.03,7),(0.03,7) and (0.0245,9.5) for (alpha, 1/beta). The selection of casts was based on being at least 200db deep and having a fairly quiet descent rate. Casts #54 and #279 were selected. The up and downcasts were quite different so it is difficult to determine which setting is best.

The best results for this cruise used (0.02, 9) for the primary and (0.03, 9) for the secondary, but the differences were slight. This is the same as when these sensors were used during 2004-32.

CELLTM was run using (0.02, 9) for the primary and (0.03, 9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Five casts that sampled to 500db were examined to look at differences between channels.

| Cast | T1-T0 | C1-C0 | S1-S0 | Descent rate |
| --- | --- | --- | --- | --- |
| 147 | -0.0014 | -0.0012 | -0.012 | Extremely noisy |
| 154 | -0.001 (X Noisy) | -0.0012 (X Noisy) | -0.013 | Extremely noisy |
| 157 | -0.001 (X Noisy) | -0.001 (X Noisy) | -0.01 (X noisy) | Extremely noisy |
| 160 | -0.0009 | -0.0012 | -0.012 | Noisy |
| 193 | -0.0012 | -0.0013 | -0.0135 | Extremely noisy |

These differences are very large.

The transmissivity is frequently extremely noisy below 250db. Editing will be attempted since the bad data all looks close to zero. Similar observations were made during the last Ricker cruise.

As noted earlier, the fluorescence is very odd near the surface for downcasts #1 to #63, going from near-zero values in the top 10m and suddenly moving to much higher values. This is similar to what was noted during 2004-32. After the bleed valve was cleaned before cast #66 the problem disappears, so this confirms that the source of the problem is in the pump system. The fluorescence channel should be removed from casts #1-63.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

The station names early in the cruise were wrong; these were changed to match the log.

##### Checking Headers

An initial run of header check showed there were errors in cast #115. The SeaBird headers were corrupted. The original DAT file was put through a routine to repair it (and the output was named 9115 to distinguish it from the original file). That improved the pressure data, but there remained problems in the headers - no position information and the time was in PST rather than UTC. After these were fixed using information from the log, the file was reconverted to IOS header format.

A header summary and a header check were produced. The only errors found were a few more errors in station names. Those were corrected.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 5.3db, with a minimum of 1.8db; these values are higher than usual given that Hugh Maclean was operating the CTD. Looking at cast #36 during the upcast the salinity is suspiciously close to zero when the pressure is about 1.5db and the pumps are on. However, casts #1 and 142 both show fairly high values of salinity at that depth, so the low salinity for #36 may be real; that cast was taken close to shore. In previous use of this equipment the pressure offset was zero, but we do know that as these instruments age the pressure calibration drifts. However, all other drifts observed in SBE 911+ CTDs have been in the opposite direction. No recalibration will be done at this time since the evidence is mixed, but an eye should be kept on this instrument to see if there is some drift occurring.

The mixed-layer depth calculation shows considerable variation at the level of the surface bottles (~10db), so these are likely of little use for salinity calibration purposes.

##### SHIFT

Fluorescence

The downcast fluorescence is suspect early in the cruise. It differs from the upcast in a way that does not suggest a misalignment. It is believed that the bleed valve was clogged. Tests run on the casts after the bleed valve was cleaned gave highly variable results, so it was decided to use a setting of +24 as has been applied to most cruises in the past.

Conductivity

For two recent cruises using this equipment a setting of -1.2 records for the primary conductivity was found effective in minimizing salinity spiking without oversmoothing. The secondary salinity has not been archived in recent cruises and for this cruise the secondary data looks very bad for some casts, so tests will not be done for secondary conductivity. Tests were run on casts #147 and 264 using advancements of from -0.8 to -1.4 records. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The setting of -1.2 looked best overall.

All data were put through SHIFT using -1.2 for the primary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Minimum Salinity: 5

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maximum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The only warnings referred to upcast data.

##### DETAILED EDITING

The primary sensors were selected for editing.

Page plots were produced using T0, S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case.

The descent rate was generally kept high, but it was extremely noisy at times, with frequent reversals in direction.

The following casts required fairly heavy editing: 10, 39, 42, 66, 70, 133, 142147, 150.

The following casts required no editing: 27, 48, 118, 139, 228, 285.

All other casts required only light editing. Note was made in the headers of any editing.

A few casts have regions with spikes to zero values in transmissivity. A second pass through CTDEDIT was made to remove bad data from the following casts: 142, 147, 150, 154, 157, 160, 193. Note was made in the header of editing done.

The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that do not require editing.

##### Niskin bottle data and comparison to CTD

Bottle data were received from the analysts in spreadsheet format:  2005-03 SALS.xls and 2005-03chl.xls and QFnuts2005-03.xls.

In the salinity spreadsheet the sample number is the same as the event number. So for 7 casts that had 2 bottles, numbering had to be adjusted. The analyst put a leading 0 in the shallow sample #s, but this does not seem clear and a leading 0 may be dropped in some uses. So the leading 0s were dropped and a 9 was put in front of the deep sample numbers. For example #142 is the shallow sample at cast #142 and 9142 is the deep sample from that cast. There were no flags assigned, but the analyst felt none were needed. A flag channel column was added.

Flags had been assigned to the chlorophyll and nutrient data.

A few casts were checked to see where the CTD stopped near the top of the upcast. In most cases the stop was at about 16 or 17db; if the bottle was mounted 5 above that as is usual, then the samples are from 11 to 12db. According to the log the sampling was at 10db. The deep samples are from about 500m, so the bottle would be at about 495db. The EDT files were thinned to 11db and 495db and the results were exported to a spreadsheet. This will enable comparison of bottles and CTD, but the CTD data is from the downcast whereas the CHL is from the upcast, so considerable scatter is expected. That, plus the variable depth of sampling, limits the usefulness of the results.

The salinity bottle data was transferred to that spreadsheet and the CTD primary salinity and bottle salinity were compared. The deep bottles were found to be high by an average of 0.0026psu, with a range from -0.004 to +0.009psu. The shallow bottles had a lot of scatter. After excluding differences >0.1psu, the average gave the CTD higher than the bottles by 0.0073psu. (See 2005-03-sal-comp.xls)

The same process was used to compare the CHL and CTD Fluorescence. The scatter was large; as has been noted before the CTD fluorescence does not do well for values of CHL > 4. When outliers were excluded it was found that the CHL was about 1.4 times the fluorescence. When casts 1-63 were excluded (casts with clogged pump) plus one outlier, CHL was about 1.7 times the fluorescence. (See 2005-03-chl-comp.xls).

##### Other comparisons

Previous experience with these sensors – This equipment has been used only once since the last recalibration, during 2004-32. There were only 2 deep bottles, one suggesting the CTD salinity was high, the other that it was low.  Surface bottles showed a lot of scatter with the primary salinity low by about 0.004. No recalibration was applied.

Historic ranges – There were excursions from the historic ranges, but most of these casts were close to shore and/or in areas of tidal mixing, so the ranges may not be appropriate. The excursions were in both directions (too high and too low) and in both temperature and salinity; this is not considered evidence of calibration problems.

##### Recalibration

Based on the bottle data the primary salinity is believed to be slightly higher than the bottles, probably within 0.002psu. There was too much scatter to justify recalibration.

##### Special Fluorometer Processing

For casts #66 to the end, the EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. Those files were put through REMOVE  and HEADEDIT to remove unwanted channels and fix formats and channel names. The output files were named FCTD. Files were also prepared containing upcast data from the top 20db. (The SHFFL files were put through REVERSE, then DELETE, CLIP and HEADEDIT.)

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S and Transmissivity profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For casts #1 to 63, the Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Fluorescence:Seapoint, Status:Pump, Descent_Rate and Flag channels were removed from all casts. HEADER EDIT was used to fix formats and header entries and to add the following comments:

The transmissivity data are nominal and unedited except

that some records were removed in editing T and S.

The comparison with bottle data showed a lot of scatter;

salinity is considered +/-0.005psu.

The fluorescence channel was removed because the data was bad due to a

clogged bleed value in the secondary pump.

For all other casts the Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and header entries and to add the following comments:

The transmissivity data are nominal and unedited except

that some records were removed in editing T and S.

The comparison with bottle data showed a lot of scatter;

salinity is considered +/-0.005psu.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

1. Large difference in salinity channels noted with primary always higher than secondary.

66. Salinity channels closer after pump tube cleaned

142. transmissometry bad from 320 to 400db - jellyfish.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2005-03

Dates: Start: Mrach 3, 2005 End: March 28, 2005

Location: N. W. Pacific

Vessel: W.E. Ricker

Party Chief: Trudel M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2005-03

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2663 | 16/10/04 | Factory |  |  |
| Conductivity | 2280 | 15/10/04 | “ |  |  |
| Secondary Temp. | 2095 | 16/10/04 | “ |  |  |
| Secondary Cond. | 1764 | 15/10/04 | “ |  |  |
| Transmissometer | 197 | 16/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |