## SBE C-sensors
SBE04-
Conductivity | date: | 2000-12-22 | 2002-02-15 | 2002-03-07 | 2002-12-31 | 2003-01-24
# 04-2232 | g= | -10.9583 | -10.9697 | -10.4419 | -10.4431 | -10.417
| h= | 1.5495 | 1.5528 | 1.4776 | 1.4777 | 1.4733
| i= | -0.0031019 | -0.003917 | -0.0034982 | -0.003469 | -0.0031073
| j= | 0.00027151 | 0.00034012 | 0.00029422 | 0.00029279 | 0.00026962
| Cpcor= | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  | 0.0006
Conductivity | date: | 2001-03-20 |  |  |  |  | 2004-12-16
# 04-2582 | g= | -10.4002 |  |  |  |  | -10.4027
| h= | 1.5713 |  |  |  |  | 1.5718
| i= | 0.0002858 |  |  |  |  | 0.00024741
| j= | 7.3728e-05 |  |  |  |  | 7.7726e-05
| CPcor= | -9.57e-08 |  |  |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  |  |  | 3.25e-06
drift PSU/month |  | new |  |  |  |  | -0.0001
Conductivity | date: | 2001-04-11 |  |  | 2002-12-31 | 2004-05-14 | 2005-01-20 | 2005-01-20
# 04-2607 | g= | -10.7508 |  |  | -10.7616 | -10.7617 | -10.7616 | -9.1256
| h= | 1.5007 |  |  | 1.5037 | 1.5034 | 1.5035 | 1.2734
| i= | -0.0023592 |  |  | -0.0031131 | -0.0029381 | -0.0029996 | -0.001792
| j= | 0.00025292 |  |  | 0.00030162 | 0.00028378 | 0.00029037 | 0.00018425
| Cpcor= | -9.57e-08 |  |  | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  | 0 | -0.0001 | 0 | -0.0001
Conductivity | date: | 2002-03-14 |  |  | 2002-12-31
# 04-2676 | g= | -10.6862 |  |  | -10.677
| h= | 1.4924 |  |  | 1.4889
| i= | -0.0070427 |  |  | -0.0059867
| j= | 0.00063624 |  |  | 0.00054955
| Cpcor= | -9.57e-08 |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06
drift since last cal: |  | 0 |  |  | +0.00010
post-cruise cal:
Conductivity | date: | 2003-02-21 |  |  | 2004-05-25 | 2004-10-05
# 04-2809 | g= | -9.9553 |  |  | -9.961 | -9.9605
| h= | 1.4118 |  |  | 1.4134 | 1.4134
| i= | -0.0002185 |  |  | -0.00058677 | -0.00060551
| j= | 9.6919e-05 |  |  | 0.00012396 | 0.00012563
| Cpcor= | -9.57e-08 |  |  | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | new |  |  | -0.0004 | 0.0002
Conductivity | date: | 2003-02-21 |  |  | 2004-05-25 | 2004-10-05
# 04-2810 | g= | -9.8663 |  |  | -9.8736 | -9.8721
| h= | 1.4576 |  |  | 1.4595 | 1.4592
| i= | -0.00065958 |  |  | -0.001096 | -0.0010272
| j= | 0.00013121 |  |  | 0.00016409 | 0.00016068
| Cpcor= | -9.57e-08 |  |  | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  | -0.0004 | 0
Conductivity | date: | 2001-06-04
# 04-2913 | g= | -9.8663
| h= | 1.4576
| i= | -0.00065958
| j= | 0.00013121
| Cpcor= | -9.57e-08
| CTcor= | 3.25e-06
drift since last cal: | PSU/mon | 0
Conductivity | date: |  |  | 2002-05-06
# 04-0771 | g= |  |  | -4.2696
(JAMSTEC?) | h= |  |  | 0.50542
| i= |  |  | 0.0002447
| j= |  |  | 1.0228e-05
| Cpcor= |  |  | -9.57e-08
| CTcor= |  |  | 3.25e-06
drift since last cal: | PSU/mon | 0 |  | -0.00105
Conductivity | date: | 11-Mar-04 NEW |  |  |  | 2005-01-20
# 04-2992 | g= | -9.9454 |  |  |  | -9.9409
| h= | 1.372 |  |  |  | 1.37
| i= | 0.00033338 |  |  |  | 0.00093495
| j= | 0.00047816 |  |  |  | 1.3188e-05
| Cpcor= | -9.57e-08 |  |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  |  | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  |  | -0.0014
Conductivity | date: | 11-Mar-04 NEW |  |  |  | 2005-01-26
# 04-2984 | g= | -10.1869 |  |  |  | -10.1896
| h= | 1.3835 |  |  |  | 1.3841
| i= | 0.00021555 |  |  |  | 6.7445e-05
| j= | 5.7268e-05 |  |  |  | 6.9622e-05
| Cpcor= | -9.57e-08 |  |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  |  | 3.25e-06
drift since last cal: | PSU/mon | 0 (new) |  |  |  | -0.0004
Note on C41: BvH: on new SBE9/11 Primary Conductivity NEW FROM FACTORY (IS THIS SAME ACCURACY AS SENSORS FOR SBE-25?)
Note on C50: BvH: on new SBE9/11 Secondary Conductivity NEW FROM FACTORY (IS THIS SAME ACCURACY AS SENSORS FOR SBE-25?)
Note on D5: Bon: these coeffs are for the post-cruise conductivity sensor calibration before electrode repair!!
Note on E5: Bon: new C-cell, was replaced under warranty (defective electrodes)
Note on F5: Bon: these coeffs are for post-cruise conductivity sensor calibration before electrode repair!!
Note on G5: Bon: C-cell had cracks, was replaced and then re-calibrated at our expense
Note on H23: BvH: POST CRUISE calibration
Note on I23: BvH: Calibration after modifications

## SBE T-sensors
SBE03-
Temperature | date: | 2001-06-04 |  |  |  | 2004-12-10
# 03-2913 | g= | 0.0043476 |  |  |  | 0.004348
(coef for ITS-90) | h= | 0.00064083 |  |  |  | 0.00064166
| i= | 2.2647e-05 |  |  |  | 2.3223e-05
| j= | 2.104e-06 |  |  |  | 2.2355e-06
| fo= | 1000 |  |  |  | 1000
drift degrC/yr: |  | new |  |  |  | 0.00111
Temperature | date: | 2001-01-01 | 2002-02-15 | 2002-12-31
# 03-4044 | g= | 0.0043426 | 0.0043427 | 0.0043426
coeffs for | h= | 0.00064252 | 0.00064268 | 0.00064264
ITS-90 | i= | 2.2433e-05 | 2.2553e-05 | 2.2535e-05
| j= | 1.9681e-06 | 1.9947e-06 | 1.9932e-06
| fo= | 1000 | 1000 | 1000
drift degrC/yr: |  | 0 | -0.0007 | +0.00047
Temperature | date: | 2000-06-24
# 03-2926 | g= | 0.0043698
| h= | 0.00064799
| i= | 2.3381e-05
| j= | 2.1484e-06
| fo= | 1000
Temperature | date: | 2001-03-21 |  | 2002-12-31 | 2004-05-15 | 2004-12-10
# 03-4115 | g= | 0.0043583 |  | 0.0043583 | 0.0043583 | 0.0043584
| h= | 0.00064523 |  | 0.00064527 | 0.00064536 | 0.00064564
| i= | 2.2859e-05 |  | 2.29e-05 | 2.2954e-05 | 2.314e-05
| j= | 1.9982e-06 |  | 2.0089e-06 | 2.0203e-06 | 2.0616e-06
| fo= | 1000 |  | 1000 | 1000 | 1000
drift degrC/yr: |  |  |  | +0.00003 | +0.00017 | +0.00026
Temperature | date: | 2002-03-14 |  | 2002-12-31
# 03-4109 | g= | 0.0043846 |  | 0.004384
| h= | 0.00064987 |  | 0.00064875
| i= | 2.3482e-05 |  | 2.2729e-05
| j= | 1.9523e-06 |  | 1.7861e-06
| fo= | 1000 |  | 1000
drift degrC/yr: |  |  |  | +0.00102
Temperature | date: | 2003-03-08 |  | 2004-05-25 | 2004-10-05
# 03P4322 | g= | 0.0044195 |  | 0.0044195 | 0.0044193
(ITS-90 coeffs) | h= | 0.00065479 |  | 0.00065467 | 0.00065433
| i= | 2.4111e-05 |  | 2.4027e-05 | 2.3804e-05
| j= | 1.9761e-06 |  | 1.9572e-06 | 1.9079e-06
| fo= | 1000 |  | 1000 | 1000
drift degrC/yr: |  |  |  | -0.00015 | -0.00018
Temperature | date: | 2003-03-07 |  | 2004-05-25 | 2004-10-05
# 03P4239 | g= | 0.0044146 |  | 0.0044146 | 0.0044146
| h= | 0.00065408 |  | 0.00065401 | 0.00065399
| i= | 2.3742e-05 |  | 2.3682e-05 | 2.3676e-05
| j= | 1.9085e-06 |  | 1.8943e-06 | 1.8944e-06
| fo= | 1000 |  | 1000 | 1000
drift degrC/yr: |  |  |  | -0.00018 | -0.00062
Temperature | date: |  |  | 2002-05-04
# 03-1086 | g= |  |  | 0.0048156
| h= |  |  | 0.00067502
| i= |  |  | 2.6032e-05
| j= |  |  | 2.0502e-06
| fo= |  |  | 1000
drift degrC/yr: |  |  |  | -0.0003
Temperature | date: | 12/03/2004 (NEW) |  |  | 2005-01-20
# 03-4397 | g= | 0.0043828 |  |  | 0.0043828
| h= | 0.00064745 |  |  | 0.00064752
| i= | 2.2363e-05 |  |  | 2.24e-05
| j= | 1.7932e-06 |  |  | 1.7995e-06
| fo= | 1000 |  |  | 1000
drift since last cal | degrC/yr: |  |  |  | 8e-05
Temperature | date: | 06/03/2004 (NEW) |  |  | 2005-01-20
# 03-4402 | g= | 0.0043692 |  |  | 0.0043692
| h= | 0.00064929 |  |  | 0.00064934
| i= | 2.2614e-05 |  |  | 2.2638e-05
| j= | 1.7781e-06 |  |  | 1.7822e-06
| fo= | 1000 |  |  | 1000
drift degrC/yr: |  |  |  |  | 8e-05
Note on C42: BvH: on new SBE9/11 Primary Temperature HIGHER ACCURACY THAN SBE-25 SENSORS?
Note on C50: BvH: on new SBE9/11 Primary Temperature HIGHER ACCURACY THAN SBE-25 SENSORS?

## SBE P-sensors
SBE-Psensors |  |  | see individual CTD's for SBE-19's
# 290442 | date: | 2000-11-15 | 2002-02-21
max=1024m | A0= | 787.8446 | 787.0138
| A1= | -0.20726 | -0.20721
| A2= | 2.7948e-07 | 2.6257e-07
| TC= |  | -140
#195334 | date: | 2001-04-27
max=1024m | A0= | 795.367
| A1= | -0.20741
| A2= | 2.8042e-07
TC=
# 290436 | date: | 2001-01-19 | 2002-02-11
max=6800m | A0= | 5285.741 | 5286.5012
| A1= | -2.7596 | -2.7587
| A2= | -1.205e-06 | -1.2983e-06
| TC= |  | 19
# 63507 | date: | 1996-01-11
10,000 m? | C1 | -44354.68
on SBE9-0443 | C2 | -0.18457
| C3 | 0.011927
| D1 | 0.039799
| D2 | 0
| T1 | 29.9357
| T2 | -0.00033362
| T3 | 3.7592e-06
| T4 | 1.5653e-09
| T5 | 0
| AD590M | 0.012833
| AD590B | -9.295
# 290464 | date: | 2001/2 ? | 2002-12-27
max= 1024 m? | A0= |  | 796.2542
| A1= |  | -0.20736
| A2= |  | 2.5193e-07
| TC= |  | -264
# 90559 | date: | 2002-10-29
max= 6800 m? | C1 | -44678.87
(fixed) | C2 | -0.628
on SBE9-0724 | C3 | 0.014524
| D1 | 0.037286
| D2 | 0
| T1 | 30.1586
| T2 | -0.00042482
| T3 | 3.8582e-06
| T4 | 4.3914e-09
| T5 | 0
| Slope | 1
| Offset | 0
| AD590M | 0.012854
| AD590B | -9.035
# 91164 | date: | 2004-03-18
max= 6800 m? | C1 | -46369.51
max= 10,000psi | C2 | -0.074006
on SBE9-0756 | C3 | 0.014342
| D1 | 0.034424
| D2 | 0
| T1 | 30.4004
| T2 | -0.0003058
| T3 | 4.4282e-06
| T4 | 1.9912e-09
| T5 | 0
| Slope | ?
| Offset | ?
| AD590M | ?
| AD590B | ?
Note on C49: BvH: on new SBE9/11, not interchangeable

## SBE-05 pumps
# 053373
# 053164 |  | on SBE9-0443 for Arctic-2002
# 051604 | IOS/OSAP | on SBE9-0443 for Arctic-2002
# 053610 |  | on new SBE9-0724
# 053615 |  | on new SBE9-0724
#053869 |  | on new SBE9-0756
#053871 |  | on new SBE9-0756

## SBE9-0756
SBE9 - 0756 |  | NEW March 2004 'for Arctic Use'
date | temperature | conductivity | pump | temperature | conductivity | pump | pressure | D.O. | transmissometer | fluorometer | turbidity | Bottom | Altimeter | Comments
| primary | primary | primary | secondary | secondary | secondary |  | sensor | Wetlabs | Seapoint | sensor | Contact Warning | sensor
Set-up as received March 2004:
2004-03-31 | 4397 | 2984 | yes | 4402 | 2992 | yes | 91164 | SBE43-615 | none | none | none | none | none | original setup
Set-up for August 2004 LSL (2004-16) as backup. Test cast only, not used afterwards.
2004-07-25 | 4397 | 2984 | yes | 4402 | 2992 | yes | 91164 | none | none | none | none | none | none | did not have to use
Set-up for August 2005 LSL (2005-04) as backup.
2005-06-03 | 4397 | 2984 | yes | 4402 | 2992 | yes | 91164 | none | none | none | none | none | none

## SBE9-0724
SBE9 - 0724 |  | NEW! from IOS for Arctic use
date | temperature | conductivity | pump | temperature | conductivity | pump | pressure | D.O. | transmissometer | fluorometer | turbidity | Bottom | Altimeter | Comments
| primary | primary | primary | secondary | secondary | secondary |  | sensor | Wetlabs | Seapoint | sensor | Contact Warning | sensor
Set-up as received March 2003:
2003-03-31 | 4322 | 2809 | 3610 | 4239 | 2810 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | 10133 |  | none | original setup
Set-up for August 2003 LSL (2003-21):
2003-07-15 | Same as original configuration??
Set-up for August 2004 LSL (2004-16):
2004-07-25 | 4322 | 2809 | 3610 | 4239 | 2810 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | none | yes | Datasonics | original setup with addition of WHOI altimeter
PSO-916T#640
Set-up for Sept-2004 LSL (2004-33):
2004-09-25 | 4397 | 2992 | 3610 | 4402 | 2984 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | none | yes | none | as original but T&C pairs from NEW 9/11 (SBE9-0756) system
Set-up for August 2005 LSL (2005-04):
2005-06-03 | 4322 | 2809 | 3610 | 4239 | 2810 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | none | yes | Datasonics PSA-916D #1161 | original setup now with first use of Arctic Altimeter

## SBE9-0443 (IOS)
SBE9 - 0443 | on loan from IOS/OSAP |  | for 2002-20, -21, -22, -23 (LAURIER Leg1-3 & LOUIS STLAURENT)
date | temperature | conductivity | pump | temperature | conductivity | pump | pressure | transmissometer | fluorometer | Comments
| primary | primary | primary | secondary | secondary | secondary
1996-07-31 | 2106 | 1748 | 1602 | 2062 | 1764 | 1604 | 63507 |  |  | original setup
---
2002-06-15 | 4044 | 2232 | 3164 | 4109 | 2676 | 1604 | 63507 | SeaTech# 139 | SeaPoint #2336 | also Touchdown Switch
Laurier leg1,2,3
SeaTech# 192D
deep casts LSL
Note on I13: BvH: used deep transm from OSAP for 2002-23, but it died after several casts, so swapped #139 back for shallower casts.

## SBE25-0293
SBE25-0293 |  |  |  | press rating
# 2524864-0293 | max.depth
| 3400 m |  | ARCTIC-2001
sensors | Temperature | Conductivity | Pump: | Pressure | P-range | Transmissom | Fluorometer | DO
type | SBE-03 | SBE-04 | SBE5-01 | SBE-29 |  | SeaTech | Seapoint | SBE-43
press rating | [5 empty] | 2000m | 6000m
ARCTIC-2001 | # 034044 | # 042232 | #053000 | # 290442 | 1024m | # 139 | # 2336
|  |  |  | # 290436 | 6800m
ARCTIC-2002 | #034115 | #042607 | #053100? | # 290464 | 1024m
|  |  |  | # 290436 | 6800m
ARCTIC-2003
ARCTIC-2004 | #034115 | #042607 |  | # 290464 | 1024m |  | # 2336 | # 430052
30x (5mg/l)
ARCTIC-2005

## SBE25-0307
SBE25-0307 | purchased 2001
#2525901-0307
max.depth 600m | Temperature | # 2926
| Conductivity | # 2607
| Pressure | # 195334 | max 1024 dbar
| Pump: SBE5-01 | #053000
THIS SBE-25 EXTENSIVELY MODIFIED BY MELLING/LINDSAY:
no longer standard modular but internal pump, sensors in-line for through-ice deployments

## SBE16-4056
SBE 16plus SEACAT
S/N | 16P25902-4056
1) Count, Temperature
Serial number | 4056
Calibrated on | 20-Apr-01
A0 | 0.0012433
A1 | 0.00026616
A2 | -6.0888e-07
A3 | 1.7815e-07
Slope | 1
Offset | 0
2) Frequency, Conductivity
Serial number | 4056
Calibrated on | 20-Apr-01
G | -0.96832
H | 0.13551
I | -0.00033778
J | 4.3232e-05
CTcor | 3.25e-06
CPcor | -9.57e-08
Slope | 1
Offset | 0
3) Count, Pressure, Strain Gauge
Serial number | 4056
Calibrated on | 25-Apr-01
PA0 | 0.11494
PA1 | 0.0015959
PA2 | 6.7871e-12
PTEMPA0 | -65.5077
PTEMPA1 | 52.4302
PTEMPA2 | -0.58815
PTCA0 | 525308.1
PTCA1 | 4.6884
PTCA2 | -0.068187
PTCB0 | 24.2797
PTCB1 | -0.00045
PTCB2 | 0
Offset | 0

## SBE16-4057
SBE 16plus SEACAT
S/N | 16P25902-4056
1) Count, Temperature
Serial number | 4057
Calibrated on | 20-Apr-01
A0 | 0.0012408
A1 | 0.00026678
A2 | -5.7575e-07
A3 | 1.7983e-07
Slope | 1
Offset | 0
2) Frequency, Conductivity
Serial number | 4057
Calibrated on | 20-Apr-01
G | -0.98226
H | 0.1364
I | -0.00051315
J | 5.6495e-05
CTcor | 3.25e-06
CPcor | -9.57e-08
Slope | 1
Offset | 0
3) Count, Pressure, Strain Gauge
Serial number | 4057
Calibrated on | 25-Apr-01
PA0 | 0.024922
PA1 | 0.0015895
PA2 | 6.8826e-12
PTEMPA0 | -70.2803
PTEMPA1 | 49.3556
PTEMPA2 | -0.040851
PTCA0 | 524442.6
PTCA1 | 3.036
PTCA2 | -0.097682
PTCB0 | 24.3084
PTCB1 | -0.000925
PTCB2 | 0
Offset | 0

## SBE19-1030
SBE19-1030 |  | (pump: SBE5-01) | [8 empty] | mod: 1998/05 new housing, new press 3K psi?
# 195891-1030 |  | #050478 | [8 empty] | mod sept-98: | new 5Kpsia pressure sensor
Factory cal
Temperature | cal: | 1992-02-20 | 1993-06-25 | 1994-03-16 | 1995-05-31 | 1995-10-31 |  |  |  | 1997-05-15 | date: | 1998-08-28 | date: | 1999-12-16 | 2001-03-23 | 2002-02-20
# 1030 | a= | 0.0036742 | 0.0036816 | 0.0036734 | 0.0036834 | 0.0036835 |  |  | G= | 0.0041813 | G= | 0.0041787 | G= | 0.0041786 | 0.0041809 | 0.0041783
| b= | 0.00058077 | 0.00058052 | 0.00058055 | 0.00057941 | 0.00058064 |  |  | H= | 0.00059305 | H= | 0.00058661 | H= | 0.00058602 | 0.00059202 | 0.00058514
| c= | 9.2078e-06 | 5.7046e-06 | 9.2139e-06 | 1.823e-06 | 7.4723e-06 |  |  | I= | 5.5332e-06 | I= | 3.3133e-07 | I= | -2.9748e-07 | 4.6952e-06 | -1.0628e-06
| d= | -5.995e-07 | -5.4104e-07 | -4.1614e-07 | -9.4383e-06 | -2.7471e-06 |  |  | J= | -1.2749e-06 | J= | -2.6466e-06 | J= | -2.8314e-06 | -1.5024e-06 | -3.0535e-06
| fo= | 2371.56 | 2341.4 | 2374.56 | 2333.47 | 2371.56 |  |  | Fo= | 1000 | Fo= | 1000 | Fo= | 1000 | 1000 | 1000
drift (degC/yr) | = | none (new) |  | 0.0015 | [7 empty] | 0.00134 |  | 0.00143 | -0.0038 | 0.00232
factory cal:
Conductivity | cal: | 1992-02-20 | 1993-06-25 | 1994-03-16 | 1995-05-31 | 1995-10-31 |  | 1997-05-15 |  | 1997-05-15 |  | 1998-08-28 |  | 1999-12-16 | 2001-03-23 | 2002-02-20
# 1030 | a= | 0.043492 | bad cal? | 0.0234 | -0.063445 | 0.29952 | G= | -4.1731 | G= | -4.1731 | G= | -4.1815 | G= | -4.1816 | -4.1846 | -4.1928
| b= | 0.45132 |  | 0.50216 | 0.75248 | -0.31391 | H= | 0.5052 | H= | 0.49813 | H= | 0.49927 | H= | 0.49933 | 0.49976 | 0.50151
| c= | -4.1579 |  | -4.187 | -12.3762 | 12.9003 | I= | 0.00068681 | I= | 0.0011809 | I= | 0.0011068 | I= | 0.0011022 | 0.0010729 | 0.00076572
| d= | 0.00052781 |  | 0.00038087 | -0.19057 | 0.28918 | J= | -2.1004e-06 | J= | -2.7492e-05 | J= | -2.4442e-05 | J= | -2.4174e-05 | -2.1087e-05 | -7.2566e-06
| m= | 2.1 |  | 3.3 | 2.3 | 2.3 | Cpcorr | -9.57e-08 | Cpcorr | -9.57e-08 | Cpcorr | -9.57e-08 | Cpcorr | -9.57e-08 | -9.57e-08 | -9.57e-08
[7 empty] | CTcorr | 3.25e-06 | CTcorr | 3.25e-06 | CTcorr | 3.25e-06 | CTcorr | 3.25e-06 | 3.25e-06 | 3.25e-06
drift (psu/mon) | = | none, new |  | 0.0001 | [7 empty] | 0.0026 |  | 0.002 | 0.0018 | 0.0001
Pressure | cal: | 1992-02-27 |  | 1994-03-09 |  |  |  |  | date: | 1997-05-12
# 154898 | A0 | 148.8465 |  | 148.925 |  |  |  |  | A0 | 143.0902
(300 psia) | A1 | -0.039069 |  | -0.039071 |  |  |  |  | A1 | -0.039057
| A2 | 3.4771e-08 |  | 3.2421e-08 |  |  |  |  | A2 | 3.3418e-08
Pressure | cal: | [8 empty] | 1997-05-19
# 177965 | A0 | [8 empty] | 499.8601
(1000 psia) | A1 | [8 empty] | -0.13045
| A2 | [8 empty] | 1.9117e-07
Pressure | cal: | [9 empty] | date: | 1998-06-04
# 184833 | A0 | [9 empty] | A0 | 1509.2799
(3000 psia) | A1 | [9 empty] | A1 | -0.39093
| A2 | [9 empty] | A2 | -3.0934e-09
Pressure | cal: | [9 empty] | date: | 1998-08-26 | date: | 1999-12-27 | 2001-03-27 | 2002-02-27
# 189759 | A0 | [9 empty] | A0 | 2501.659 | A0 | 2502.7156 | 2502.824 | 2502.9772
(5000 psi) | A1 | [9 empty] | A1 | -0.65204 | A1 | -0.65194 | -0.65197 | -0.65202
| A2 | [9 empty] | A2 | -5.9724e-08 | A2 | -1.6646e-08 | -1.751e-08 | -3.8302e-09
T-compensation | TC | [14 empty] | -221
SERVICE: |  |  |  | Factory cal. T,C,P | Factory cal. T,C | [6 empty] | Calibr T, C, P |  | Factory cal. T,C,P | Factory cal. T,C,P
[5 empty] | check & cal 300psi | [6 empty] | replaced 3Kpsi with 5Kpsi |  | replaced 6-pin conn | replaced 6-pin I/O
[5 empty] | install & cal 1000psi | [6 empty] | Pump SBE5 #050478: |  | on pump 050478 | and 6-pin AUX conn
[5 empty] | found spreadsheet: | [6 empty] | replaced 6-pin conn &
[5 empty] | 300psi test (IOS?) | [6 empty] | new pump interface cable
1995-05-26
88.5604
-0.026882
1.9057e-08
-2.3418e-12

## SBE19-2688
SBE19-2688 |  |  |  | repaired pump #SBE5-050478
#1920437-2688
Temperature | date: | 1999-02-02 | 1999-12-16 | 2001-03-31 | 2002-02-20 | 2004-12-15
# 2688 | G= | 0.004238 | 0.004237 | 0.0042425 | 0.0042349 | 0.0042364
| H= | 0.000618 | 0.00061518 | 0.00062824 | 0.00061001 | 0.00061361
| I= | 1.0711e-05 | 8.2355e-06 | 1.8565e-05 | 4.1161e-06 | 7.0516e-06
| J= | -4.4827e-07 | -1.1436e-06 | 1.5603e-06 | -2.2012e-06 | -1.4298e-06
| F0= | 1000 | 1000 | 1000 | 1000 | 1000
| drift/month = | none, new | 0.00218 | 0.00359 | -0.00331 | 0.00041
Conductivity | date: | 1999-02-02 | 1999-12-16 | 2001-03-31 | 2002-02-20 | 2004-12-15
# 2688 | G | -4.2223 | -0.0042219 | -4.2266 | -4.2313 | -4.2255
| H | 0.5052 | 0.50487 | 0.50557 | 0.50701 | 0.50544
| I | 0.00068681 | 0.00078607 | 0.00066489 | 0.00039076 | 0.00074027
| J | -2.1004e-06 | -7.713e-06 | -1.9877e-06 | 1.176e-05 | -6.4719e-06
| CP-corr | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CT-corr | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
| drift PSU/month = | none, new | 0.0006 | 0.0001 | -0.0003 | -0.0002
Pressure | date: | 1999-03-29 | 1999-12-28 | 2001-04-04 | 2002-02-23 | 2005-01-04
# 192787 | PA0 | 500.4465 | 500.8254 | 501.396 | 501.4606 | 502.1478
(1000psia) | PA1 | -0.13016 | -0.13019 | -0.13018 | -0.13015 | -0.13018
| PA2 | 1.3081e-07 | 1.2742e-07 | 1.2591e-07 | 1.2897e-07 | 1.1719e-07
T-compens | TC= |  |  |  | -70 | -70
note: 6-pin IO conn
damaged at Seabird
replaced by factory
battery bulkh re-glued

## SBE19+4132
SBE19+4132 |  | This SBE19plus is on loan from Jim Gower.
We have replaced the 100m pressure sensor with a 600m(?) and sensor calibrations Jan 2005: see 4132-05.CON
Temperature | date: | 2005-01-04
# 4132 | A0 | 0.0011941
| A1 | 0.00028429
| A2 | -3.7353e-06
| A3 | 3.3708e-07
| Slope | 1
| Offset | 0
Conductivity | date: | 2005-01-04
# 4132 | G | -0.96451
| H | 0.13547
| I | -0.00042622
| J | 4.8554e-05
| CTcor | 3.25e-06
| CPcor | -9.57e-08
| Slope | 1
| Offset | 0
Pressure | date: | 2004-12-30
# 4132 | PA0 | -0.68787
(1000psia) | PA1 | 0.0044295
| PA2 | -1.3836e-11
| PTEMPA0 | -72.2228
| PTEMPA1 | 48.4525
| PTEMPA2 | 0.076849
| PTCA0 | 520412.8
| PTCA1 | -13.8036
| PTCA2 | 0.3222
| PTCB0 | 24.711
| PTCB1 | 0.0016
| PTCB2 | 0
| Offset | 0

## SBE19-3112
On loan from Grant Ingram (UBC)
SBE19-3112
1925666-3112
Temperature | date: | 2001-03-22
# 3112 | G= | 0.004195
| H= | 0.00061401
| I= | 1.2935e-05
| J= | -4.9472e-07
| F0= | 1000
| drift/month = | none (new)
Conductivity | date: | 2001-03-22
# 3112 | G | -4.1164
| H | 0.49289
| I | 0.00053015
| J | 7.859e-06
| CT-corr | 3.25e-06
| CP-corr | -9.57e-08
| drift PSU/month = | none, new
Pressure | date: | 2001-03-23
# 165700 | PA0 | 149.5054
(300 PSIA) | PA1 | -0.039268
Max.200m | PA2 | 4.0474e-85
T-compens | TC= | 87
Pump
SBE 5T # 261

## SBE19-1294
SEABIRD SEACAT CALIBRATIONS (from IOS spreadsheet by Minkley)
serial number 1294
date | 1998-04-29 | 1999-06-17
temperature
g | 0.0042368 | 0.0042388
h | 0.00059866 | 0.00060264
I | 5.2131e-06 | 7.9157e-06
j | -1.4243e-06 | -8.0729e-07
f0 | 1000 | 1000
Conductivity
g | -4.0406 | -4.0336
h | 0.48217 | 0.48061
I | 0.0012708 | 0.0015497
j | -3.7073e-05 | -4.9584e-05
Cpcor | -9.57e-08 | -9.57e-08
Ctcor | 3.25e-06 | 3.25e-06
Pressure (quadratic Fit)
A0 | 496.4681 | 496.5765
A1 | -0.13028 | -0.13028
A2 | 1.2446e-07 | 1.2761e-07
TC | -44 | -44

## SBE37-0071
SBE-37 # 0459 |  | Humfrey's?
Temperature: | Cal. date: | 2002-11-26 | 2004-12-14
#0459 | a0 | (Lindsay?) | -6.3299e-05
| a1 |  | 0.00028965
| a2 |  | -3.5177e-06
| a3 |  | 1.8365e-07
drift since prev cal | (deg.C/yr) |  | -1e-05
Conductivity: | Cal. date: | 2002-11-26 | 2004-12-14
# 0459 | g |  | -1.0478
| h |  | 0.15201
| i |  | -0.00044553
| j |  | 5.6951e-05
| Cpcor |  | -9.57e-08
| Ctcor |  | 3.25e-06
| WBOTC |  | -8.079e-05
drift since prev cal | (PSU/mon) |  | -0.0005
RTC freq | Cal. date: | 2002-11-26 | 2004-12-14
ITS-90 scale | rtca0 |  | 0.99997
| rtca1 |  | 1.7725e-06
| rtca2 |  | -2.8914e-08

## SBE43 D.O.
ordered May 2001 - delivered end June - used in sea trials - see remarks
# 430052 | Cal. Date: | 2001-06-12 | 2001-08-06 |  | 2005-10-07
|  |  | repaired! |  | repaired!
Coefficients: | Soc= | 0.3557 | 0.3159 |  | 0.3322
| Boc= | 0.0138 | 0.0019 |  | 0
| Offset | -0.6129 | -0.6129 |  | -0.5987
| Tcor= | 0.0013 | 0.0003 |  | 0.0007
| Pcor= | 0.00015 | 0.00015 |  | 0.000135
|  | 2001-07-01 |  | Aug,2004 | Sept. 2005
Services performed: |  | large hysteresis in initial tests: returned unit from Dutch Harbor to factory for repair | factory replaced torn membrane, | returned to factory for repair (?), replaced bulkhead connector. Bad data in 2004-13 cruise!! | returned for repair: SPAR assembly replaced, no post-cruise cal, new cal above.
#430615 | New date: | 2004-03-09
Came with SBE9-0756
Coefficients: | Soc= | 0.3916
| Boc= | 0
| Offset | -0.4643
| Tcor= | 0.0011
| Pcor= | 0.000135
#43-0435 | New date: | 2003-03-18 |  | 2005-03-14
Came with SBE9-0724
Coefficients: | Soc= | 0.3581 |  | 0.4159
| Boc= | 0 |  | 0
| Voffset | -0.4804 |  | -0.4812
| Tcor= | 0.0015 |  | -0.0006
| Pcor= | 0.000135 |  | 0.000135
| Tau= | 0 |  | 0
Comments |  |  |  | Torn membrane replaced prior to calibration
#430071 | New date: | 2001-06-01 | Factory loaner used for Leg 1 of 2001-21
Coefficients: | Soc= | 0.4273
| Boc= | 0.0201
| Offset | -0.6275
| Tcor= | 0.0014
| Pcor= | 0.00015

## SBE23-0326 DO
SBE-23 YSI Dissolved Oxygen sensor
# 230326 | Cal. Date: | 1992-02-28 | 1993-03-30 | 1993-03-31 | 2001-04-10
|  |  | post-cruise | new diaphram | new diaphram
Sensor current | m= | 1.053e-05 | 1.053e-05 | 1.053e-05 | 1.053e-05
| b= | -1.053e-09 | -1.053e-09 | -1.053e-09 | -1.053e-09
Comp.Temp. | k= | 6.4851 | 6.4851 | 6.4851 | 6.4852
| c= | -2.17 | -2.17 | -2.17 | -2.17
Seasoft coeffs: | Soc= | 0.0652 | 0.3254 | 0.078 | 0.0783
( Oxfit ) | Boc= | -0.0517 | -0.4451 | -0.036 | -0.0156
| tcor= | -0.033 | -0.033 | -0.033 | -0.033
| pcor= | 0.00015 | 0.00015 | 0.00015 | 0.00015
| tau= | 2 | 2 | 2 | 2
| wt= | 0.85 | 0.85 | 0.85 | 0.85
| barom(mB)= | 1013.13 | 1013.037 | 1005.349 | 1014.015
| Twater(C)= | 23.34 | 21.6 | 21.85 | 4.495
| Tcomp(C)= | 23.35 | 21.73 | 21.97 | 4.263
| Isat(uA)= | 33.937 | 7.633 | 26.629 | 15.015
| Iair(uA)= | 35.074 | 7.37 | 27.293 | 25.955
| Izero(uA)= | 0.793 | 1.368 | 0.462 | 0.199

## Fluorometers
FLUOROMETERS
SeaTech | new: | 1994 ?
S/N S93 | unknown | set by internal jumper
SeaPoint | new: | 2001-03-01
S/N 2336 | gain: | 30 x, 0-5 ug/L
max.depth | offset: | 0
6000m
Note: a Y-cable was used in 2002
was modified 10 Aug 2002 on LSL to float (open) pin# 5 & 6
Fluoro range was 30X (i.e. pin 5&6 open)
SeaPoint | new: | 2001-03-01
S/N 2569 | gain: | 30 x, 0-5 ug/L
max.depth | offset: | 0
6000m
Wetlabs | date: |  | DGS modified the SBE19-1030 to power this unit, was used in 2005 for surveys of Okanagan Lake,
Wetstar Chla | Vblank= | 0.068 | Sept-'02: shipped to SEDNA for use on SBE19-3112 (UBC loan, Seabird checked that it was modified to take Wetstar fluorom)
WS3S521P | Scale fact= | 14
SeaPoint | new:
S/N 2569 | gain:
max.depth | offset:
6000m
SEAPOINT GAIN is controlled by cable: setting pin #5 and/or #6 either open (=5V) or short-to-ground (=0V) see manual.
Note on C8: BvH: RANGE SETTING: Y-cable used in 2002 on SBE9/11 connected Fluorometer AND Transmissom; modified 10 Aug 2002 to float pin 5&6 (30X)
Note on C16: BvH: GAIN is controlled by cable !

## Transmissometers
TRANSMISSOMETERS=SeaTech |  |  |  | (added April 2002 from separate spreadsheet)
serial | cal.date | V-air | lab | V-ref | blank | Comments | Guildline: |  | Seabird:
number | mm/dd/yy | voltage | cal. |  | voltage |  | Offset | Slope | M= | B=
Wet-Labs: |  | A | B | 90.2 | Z | 600 nm,25cm path,6000m
CST-662DR | 2003-03-20 | 4.86 | 4.86 | 4.772 | 0.054 | Tcal=19.9 Trans%= |  |  | 18.9019 | -1.0207
|  | Factory | Cal. |  | Cal | Pmax=? Pathlength? | Guildline: |  | Seabird:
| mm/dd/yy | V-air | V-air |  | V-block | comments | Offset | Slope | M= | B=
SeaTech: |  | A | B |  | Z
23 | 1980-08-06 | 4.79 |  | 0.003 |  | factory
23 | 1995-07-12 |  | 4.439 |  | 0.003 | JL
26 | 1980-08-06 | 4.79 |  | -0.007 |  | factory
26 | 1995-07-12 |  | 4.3 |  | 0.004 | field, JL
28 | 1980-08-06 | 4.74 |  | -0.007 |  | factory
28 |  |  | 4.63 |  | 0.002 | JL | -2.0475 | 102.3349
78 | 1981-06-03 | 4.77 |  | -0.001 |  | factory
78 | 1988-02-02 |  | 4.576 |  | 0.0017 | lab | -2.0848 | 104.2041
78 | 1989-03-20 |  | 4.5468 |  | 0.0007 | lab | -2.0982 | 104.8943
78 | 1990-03-16 |  | 4.5455 |  | 0.0008 | lab | -2.0988 | 104.9222
78 | 1992-05-11 |  | 4.5262 |  | 0.0011 | lab | -2.1077 | 105.3632
78 | 1994-04-01 |  |  |  |  | failed at 400m
81 | 1981-06-03 | 4.81 |  | -0.002 |  | factory
81 | 1988-02-02 | 4.8101 |  | 0.0007 |  | new factory value after repair
81 | 1989-03-20 |  | 4.8578 |  | 0.0008
81 | 1989-08-22 |  |  |  |  | hit by Russian, meter and trans not returned
Seabird CON file:
109 | 1982-05-05 | 4.748 | 4.748 | 0 | 0 | factory (Bon) |  |  | M= | B=
109 | 1997-10-30 |  | 4.392 |  | 0.001 | to SHEBA (on sheet from DGS) |  |  | 21.6261 | -0.021626
130 | 1982-11-12 | 4.734 | 0.001 |  |  | EDR 004547A
| 1990-04-09 |  | 4.622 |  | 0.0015
132 | 1982-10-23 | 4.725 |  | 0.002 |  | factory
132 | 1995-07-01 |  | 4.31 |  | 0.003 | JL | -2.1926 | 109.563
136 | 1993-10-06 | 4.75 |  | 0 |  | factory |  |  | Seabird CON file:
139 | 1982-12-02 | 4.74 | 4.74 | 0.003 | 0.003 | factory |  |  | M= | B=
139 | 1993-08-31 |  | 4.74 |  | 0.009 | RP |  |  | 20.0254 | -0.18023
139 | 1998-09-17 |  | 4.434 |  | 0.004 | Bon |  |  | 21.386 | -0.085544
139 | 2000-01-15 |  | 4.562 |  | 0.006 | Doug |  |  | 20.7946 | -0.12477
139 | 2001-04-23 |  | 4.554 |  | 0.006 | Doug |  |  | 20.8311 | -0.12499
163 | guess only: | 4.745 |  | 0.001 |  | need factory values!
163 | 2000-01-15 |  | 4.564 |  | 0.001 | Doug | from where? ---> |  | 20.7945 | -0.12477
|  |  | 4.564 |  | 0.001 |  |  |  | 20.7933 | -0.020793
175 | 1985-09-24 | 4.803 |  | 0 |  | factory
175 | 1988-08-01 |  |  |  |  | smashed by fishing boat
176 | 1985-09-25 | 4.728 |  | -0.001 |  | factory
176 | 1988-01-01 |  |  |  |  | flooded
197 | 1985-10-08 | 4.749 | 4.749 | 0 | 0 | factory - EDR 007333A
197 | 1988-02-02 |  | 4.6697 |  | 0.0007 | side rods look slightly bent | -2.034 | 101.6839
197 | 1988-09-07 |  | 4.6452 |  | 0.0012 | lab | -2.0447 | 102.21
197 | 1988-11-15 |  | 4.6336 |  | 0.0011 | lab | -2.0498 | 102.468
197 | 1988-11-22 |  | 4.637 |  | 0.0009 | lab | -2.0483 | 102.3969
197 | 1989-03-20 |  | 4.6283 |  | 0.0008 | lab | -2.0522 | 102.5915
197 | 1990-03-16 |  | 4.5724 |  | 0.0009 | lab | -2.0772 | 103.8436
197 | 1990-09-25 |  | 4.5476 |  | 0.0015 | lab | -2.0886 | 104.3974
197 | 1992-05-07 |  | 4.4863 |  | 0.0018 | lab | -2.1171 | 105.8175
197 | 1994-04-04 |  | 4.39 |  | 0.002 | JL | -2.1636 | 108.1344
197 | 1996-02-08 |  | 4.209 |  | 0.002 | lab, BGM | -2.2566 | 112.7845
197 | 1996-10-24 | 4.73 | 4.73 | -0.001 | -0.001 | factory calibration | -2 | 100
197 | 1997-02-07 |  | 4.74 |  | 0.005 | lab, Doug | 1.9958 | 99.6892
231 | 1992-02-17 | 4.734 |  | 0.0001 |  | factory
231
231 | 1993-03-29 |  | 4.661 |  | -0.0004 | lab, JL | -2.0313 | 101.5743
231 | 1994-08-13 |  | 4.61 |  |  | JL
231 | 1995-04-29 |  | 4.56 |  | 0 | field, JL | -2.0763 | 103.8158
231 | 1996-02-08 |  | 4.511 |  | 0.012 | lab, BGM, unstable with large zero offset | -2.0989 | 104.6916
240 | 1992-12-29 | 4.674 |  | 0 |  | factory
241 | 1986-12-31 | 4.754 |  | 0.001 |  | factory
241 | 1992-02-17 |  | 4.669 |  | 0.0025 | lab, JL, sent to Russia.
242 | 1987-01-01 | 4.765 | 4.765 | 0.001 |  | factory 25cm path |  |  | 19.9958 | 0
242 | 1993-06-21 |  | 4.82 |  | 0.0006 |  |  |  | 19.7701 | -0.011862
242 | 1994/08/00 |  | 4.76 |  | 0 |  |  |  | 20.0168 | 0
242 | 1995-01-23 |  | 4.76 |  | 0 | Ron Cooke (?) |  |  | 20.0168 | 0
242 | 96? note found… | [7 empty] | 22.58 | -0.027
254 | 1987-06-23 | 4.755 | 4.755 | 0.003 |  | factory
254 | 1989-03-20 |  | 4.754 |  | 0.0031 | lab | -2.0004 | 99.959
254 | 1990-03-16 |  | 4.7603 |  | 0.0029 | lab | -1.9978 | 99.8307
254 | 1990-09-30 |  | 4.7462 |  | 0.0018 | lab | -2.0037 | 100.1493
254 | 1992-05-11 |  | 4.7516 |  | 0.0022 | lab | -2.0014 | 100.0275
254 | 1993-10-16 |  | 4.73 |  | -0.002 | lab | -2.0106 | 100.5688
254 | 1996-02-08 |  | 4.708 |  | 0.004 | lab, BGM | -2.02 | 100.9175
571 | 1993-03-04 | 4.673 |  | 0 |  | factory
572 | 1993-03-04 | 4.674 |  | 0 |  | factory
572 | 1994-08-13 |  | 4.58 |  |  | field, JL | -2.041 | 102.0524
598 | 1994-03-21 | 4.662 | 4.662 | 0.002 | 0.002 | factory
598 | 1994-04-02 |  | 4.65 |  | 0.001 | lab, JL | -2.0052 | 100.238 | 20.0473 | -0.020047
598 | 1995-04-02 |  | 4.63 |  | 0.002 | JL, date is a guess only!!! | -2.0138 | 100.6509 | 20.1383 | -0.040277
598 | 1996-02-13 |  | 4.622 |  | 0.004 | lab, BGM | -2.0173 | 100.7847 | 20.1819 | -0.080728
598 | 1999-06-21 |  | 3.86 |  | 0.007 | lab,DGS | -2.4155 | 120.6081 | 24.1889 | -0.16932
598 | 2000-01-15 |  | 4.029 |  | 0.002 | Doug | -2.3142 | 115.6648 | 23.1438 | -0.046288
598 | 2001-01-16 |  | 3.922 |  | 0.002 | Doug | -2.3774 | 118.8204 | 23.7755 | -0.047551
182D | 1992-07-07 | 4.672 | 4.672 | 0 |  | factory, 5000 metres | -2 | 100
182D | 1994-04-06 |  | 4.5829 |  | 0.0005 | lab | -2.0389 | 101.934
182D | 1996-02-05 |  | 4.33 |  | 0 | lab, BGM, 28448,-103.11, WOCE ctd | -2.158 | 107.8984 | 28448 | -103 | 0.32937 | 0.0031978
182D | 1996-02-08 |  | 4.323 |  | 0.001 | lab, BGM | -2.1615 | 108.0515
182D | 1996-06-10 |  | 4.332 |  | 0 | lab, BGM | -2.157 | 107.8486
182D | 1996-07-01 |  | 4.305 |  | 0.0002 | lab, Doug | -2.1705 | 108.5206
182D | 1996-08-07 |  | 4.29 |  | 0 | lab, bgm, 28136, -137 | -2.1781 | 108.9044 | 28136 | -137 | 0.4424 | 0.0032292
182D | 1996-11-13 |  | 4.682 |  | -0.001 | factory calibration
182D | [5 empty] | factory rebuild
mm/dd/yy
192D | 1993-02-18 | 4.675 |  | 0 |  | factory, 5000 metres
192D | 1994-04-06 | 4.6967 | 4.6967 | 0.0003 |  | new factory value, 18/2/93 value not aged.
192D | 1994-06-30 |  | 4.6612 |  | 0 | JL | -2.0152 | 100.7616
192D | 1995-02-05 |  | 4.59 |  | 0 |  | -2.0465 | 102.3246
192D | 1995-04-27 |  | 4.559 |  | 0.0038 | 29936, -108 | -2.0604 | 102.9421
192D | 1995-07-01 |  | 4.52 |  | 0.001 | field, JL | -2.0782 | 103.8885
192D | 1996-02-05 |  | 4.433 |  | 0 | lab, BGM, 29132,-99.34,WOCE CTD |  |  | 29132 | -99 | 0.30922 | 0.0031234
192D | 1996-02-08 |  | 4.4478 |  | 0.0018 | lab, BGM | -2.1119 | 105.558
192D | 1996-06-06 |  | 4.258 |  | 0.002 | lab, BGM | -2.2061 | 110.2588
192D | 1996-07-18 |  | 4.36 |  | 0.002 | lab, Doug | -2.1544 | 107.6794
192D | 1996-08-07 |  | 4.36 |  | 0.001 | lab, bgm, 28652, -129, woce ctd | -2.1544 | 107.7009 | 28652 | -129 | 0.40912 | 0.0031715
192D | 1996-10-31 |  | 4.66 |  | 0 | factory checkup and calibration | -2 | 100
V=(A/B*(X-Z)
V=corrected output voltage
A=air calibration, factory
B air calibration, present
X=data value (output voltage measured in water
Z=zero offset with light path blocked
Note on E5: BvH: Tw - transm in pure water for 660 nm = 90.2 (to derive M,B)
For Guildline, slope = -2.0*(A/B)
offset = 20.0*(A/B*(5.00-zero offset)
For SBE-19: M = 20(A0 - Y0)/(A1 - Y1)
B = -M Y1
A0 & A1 are the original & present AIR CALIBRATION voltages
Y0 & Y1 are the original & present BLOCKED PATH voltages
For Guildline, slope = -2.0*(A/B)
offset = 20.0*(A/B*(5.00-zero offset)
For SBE-19: M = 20(A0 - Y0)/(A1 - Y1)
B = -M Y1
A0 & A1 are the original & present AIR CALIBRATION voltages
Y0 & Y1 are the original & present BLOCKED PATH voltages

## Turbidity
Turbidity Meter
Manufact: | Seapoint
S/N: | # 10133
range: | selected | range select by grounding Pin#5 &/or 6 (same as fluorometer)!
by cable !

## Altimeter
Altimeter
New: | 2005-03-31
Manufact: | Benthos
Model: | Datasonics PSA-916D
S/N: | 1161
range: | 0-6000m

## ICTD-1329
ICTD CALIBRATION INFO
S/N = 1329
VER = 2.4
CDATE = 21MAY96
A01 = 8.287852e-02
A501 = 3.743052e+01
A1001 = 7.482841e+01
B01 = 1.179096e-05
B501 = 1.909808e-05
B1001 = 3.347727e-05
C01 = 3.858695e-10
C501 = 1.037280e-08
C1001 = 1.091776e-08
A1 = 1.336248e-02
B1 = 9.487742e-01
C1 = -5.103178e-05
D1 = 4.515775e-07
A02 = 3.592097e+01
A502 = 1.567107e+01
A1002 = -4.478627e+00
B02 = -1.651400e-05
B502 = -1.823800e-05
B1002 = -2.638600e-05
C02 = -1.925820e-08
C502 = -1.977630e-08
C1002 = -1.400120e-08
A2 = -5.716996e-04
B2 = 1.001717e+00
C2 = -1.631159e-04
D2 = 3.604622e-06
A03 = -6.070145e-01
A503 = 1.444939e+01
A1003 = 2.955527e+01
B03 = 1.614800e-04
B503 = 2.103200e-04
B1003 = 8.756400e-05
C03 = -2.800560e-08
C503 = 1.952680e-07
C1003 = 1.781560e-07
A3 = -1.995102e-02
B3 = 1.050521e+00
C3 = -4.940318e-03
D3 = 1.108796e-04
A04 = 3.554498e+01
A504 = 1.530000e+01
A1004 = -4.763292e+00
B04 = -6.646200e-05
B504 = -1.238300e-05
B1004 = -3.009600e-05
C04 = -8.802370e-09
C504 = -1.024190e-08
C1004 = -8.208591e-09
A4 = 0.000000e+00
B4 = 1.000000e+00
C4 = 0.000000e+00
D4 = 0.000000e+00
A05 = 6.951330e+03
A505 = 3.515259e+03
A1005 = 6.139834e+01
B05 = -1.024400e-02
B505 = 5.197400e-04
B1005 = 1.160800e-02
C05 = 3.284700e-06
C505 = 7.119200e-07
C1005 = 5.978300e-06
A5 = 8.020587e-01
B5 = 1.000720e+00
C5 = -3.663700e-07
D5 = 3.187000e-11
PLTC = 4.000000e+02
TRTC = 1.000000e+02
XOFF = 0.000000e+00
XGAN = 1.000000e+00
CH1 = on, CH2 = on, CH3 = on, CH4 = off, CH5 = on, EXTD = off
DC1=on, DC2=off, DC3=on, DC4=on, DC5=off, DC6=off, DC7=off, DC8=off, Alt=off
Bytes/Scan = 16, Max Bytes/Scan allowed = 122

## MCTD-1357
WHOSE INSTRUMENT??
CALIBRATION INFO
original MCTD
S/N = 1357
VER = E.2
CDATE = 26JUN95
A01 = 7.604567e+01
A501 = 3.805856e+01
A1001 = 1.528660e-01
B01 = 1.195773e-05
B501 = 5.502455e-06
B1001 = 2.855086e-06
C01 = 2.815857e-09
C501 = 1.756495e-09
C1001 = 2.061360e-09
A1 = 9.251564e-03
B1 = 9.508290e-01
C1 = -8.058304e-05
D1 = 6.863835e-07
A02 = 3.914346e-01
A502 = 1.548280e+01
A1002 = 3.048322e+01
B02 = -3.294100e-06
B502 = 3.247900e-06
B1002 = 6.598900e-06
C02 = -9.607980e-09
C502 = -2.778820e-09
C1002 = 1.582310e-10
A2 = -2.134496e-02
B2 = 1.048333e+00
C2 = -4.658864e-03
D2 = 1.033809e-04
A03 = 4.556631e+00
A503 = 1.540935e+02
A1003 = 2.992850e+02
B03 = 8.327920e-04
B503 = -2.285100e-05
B1003 = -6.088090e-04
C03 = -3.486200e-07
C503 = -3.085400e-07
C1003 = -3.163300e-07
A3 = -2.507200e-01
B3 = 1.011834e+00
C3 = -9.197901e-05
D3 = 1.778900e-07
XOFF = 0.000000e+00
XGAN = 1.000000e+00

## MCTD-1534
Pmax = ???
CALIBRATION INFO | CALIBRATION INFO | Date: 21 Mar 1996 | Date: 17 May 1996 | Date: 15 Mar 1999
original MCTD | upgraded MCTD | new 1400 dBAR transducer | 7000 dBAR TRANSDUCER | with Psensor for 7000m | RETURNED TO FSI FOR CALIBRATION: Spring 2004
(later modified to S/N 1534) | (new electronics) | Pressure Calibration | Pressure Calibration |  | damaged by FEDEX (broken-off C-sensor)
S/N = 1344 | S/N = 1533 | S/N = 1534 | S/N = 1534 | S/N = 1534 | repair estimate ~US$ 4000.-
VER = E.2 | VER = H.7 | VER = H.8 | VER = H.8 | VER = I.0 | Asked FSI to return unit without repair for spare parts of other MCTD
CDATE = 15JUN95 | CDATE = 3APR96 | CDATE = 21MAY96 | CDATE = 17MAY96 | CDATE = 17FEB99
A01 = 7.619018e+01
A01 = 7.481707e+01 | A01 = 7.687679e+01 | A03 = -4.503574e+00 | A03 = 3.063934e+01 | A501 = 3.819031e+01
A501 = 3.740881e+01 | A501 = 3.846718e+01 | A503 = 7.940275e+02 | A503 = 3.755900e+03 | A1001 = 2.476492e-01
A1001 = 1.280668e-01 | A1001 = 2.628167e-01 | A1003 = 1.564395e+03 | A1003 = 7.257228e+03 | B01 = 7.901935e-06
B01 = 1.478333e-05 | B01 = 1.196468e-05 | B03 = 3.125000e-04 | B03 = -3.371500e-03 | B501 = 1.822803e-06
B501 = 5.057728e-06 | B501 = 5.900900e-06 | B503 = 4.421800e-03 | B503 = 3.095300e-02 | B1001 = -7.228598e-07
B1001 = 2.013860e-06 | B1001 = 6.645934e-06 | B1003 = 7.693600e-03 | B1003 = 5.611130e-02 | C01 = 2.239277e-09
C01 = 1.412444e-09 | C01 = -5.000634e-09 | C03 = 1.738800e-07 | C03 = 2.287600e-06 | C501 = 1.266642e-09
C501 = 4.282004e-10 | C501 = -3.119247e-09 | C503 = 9.459300e-07 | C503 = 8.474700e-06 | C1001 = 7.272759e-10
C1001 = 7.991649e-10 | C1001 = -3.288075e-09 | C1003 = 1.536800e-06 | C1003 = 1.314800e-05 | A1 = 2.321624e-02
A1 = 8.645451e-03 | A1 = 7.837162e-03 | A3 = -2.187000e-01 | A3 = -2.513260e+00 | B1 = 9.398204e-01
B1 = 9.586888e-01 | B1 = 9.228891e-01 | B3 = 1.001102e+00 | B3 = 1.004783e+00 | C1 = -3.802821e-05
C1 = -9.121145e-05 | C1 = -7.983520e-05 | C3 = -1.605100e-06 | C3 = -2.033500e-06 | D1 = 3.344590e-07
D1 = 7.672685e-07 | D1 = 6.977060e-07 | D3 = 6.579800e-10 | D3 = 1.925600e-10 | A02 = 6.706564e-01
A02 = 6.443921e-01 | A02 = -4.330795e-03 |  |  | A502 = 1.591774e+01
A502 = 1.588781e+01 | A502 = 1.527188e+01 |  |  | A1002 = 3.108448e+01
A1002 = 3.106772e+01 | A1002 = 3.054178e+01 |  |  | B02 = -2.317800e-05
B02 = -1.903600e-05 | B02 = -1.825400e-05 |  |  | B502 = -2.056000e-05
B502 = -1.631500e-05 | B502 = -1.721800e-05 |  |  | B1002 = -1.702900e-05
B1002 = -1.194100e-05 | B1002 = -1.483700e-05 |  |  | C02 = -2.821210e-10
C02 = -8.786180e-10 | C02 = 8.165930e-10 |  |  | C502 = -1.180450e-10
C502 = -4.897190e-10 | C502 = 1.496950e-09 |  |  | C1002 = -1.890760e-09
C1002 = 6.559269e-11 | C1002 = 1.955320e-09 |  |  | A2 = -2.016589e-02
A2 = -2.136674e-02 | A2 = -1.911568e-02 |  |  | B2 = 1.047491e+00
B2 = 1.049017e+00 | B2 = 1.047338e+00 |  |  | C2 = -4.637278e-03
C2 = -4.711407e-03 | C2 = -4.578197e-03 |  |  | D2 = 1.035693e-04
D2 = 1.043942e-04 | D2 = 1.018532e-04 |  |  | A03 = 3.063934e+01
A03 = 5.089466e+02 | A03 = 5.244594e+01 |  |  | A503 = 3.755900e+03
A503 = 4.173482e+03 | A503 = 8.477444e+02 |  |  | A1003 = 7.257228e+03
A1003 = 7.684304e+03 | A1003 = 1.621784e+03 |  |  | B03 = -3.371500e-03
B03 = -1.265490e-02 | B03 = -2.416600e-03 |  |  | B503 = 3.095300e-02
B503 = -6.529770e-02 | B503 = 1.287500e-03 |  |  | B1003 = 5.611130e-02
B1003 = -1.163910e-01 | B1003 = 3.583330e-03 |  |  | C03 = 2.287600e-06
C03 = -3.709300e-06 | C03 = 1.036400e-06 |  |  | C503 = 8.474700e-06
C503 = -2.094900e-05 | C503 = 1.774100e-06 |  |  | C1003 = 1.314800e-05
C1003 = -3.664200e-05 | C1003 = 2.349800e-06 |  |  | A3 = -2.513260e+00
A3 = 5.685000e+00 | A3 = 3.113400e-01 |  |  | B3 = 1.004783e+00
B3 = 1.008536e+00 | B3 = 1.003375e+00 |  |  | C3 = -2.033500e-06
C3 = -3.697700e-06 | C3 = -6.469000e-06 |  |  | D3 = 1.925600e-10
D3 = 3.451000e-10 | D3 = 2.844000e-09 |  |  | XOFF = 0.000000e+00
XOFF = 0.000000e+00 | XOFF = 0.000000e+00 |  |  | XGAN = 1.000000e+00
XGAN = 1.000000e+00 | XGAN = 1.000000e+00
Open Mode

## MCTD-1598
Pressure Calibration
3400?? dBAR TRANSDUCER
(5000 PSIA? ...not on FSI pr.cal printout)
MCTD
MCTD S/N = 1598
VER = ??
CDATE = 29SEP98
A03 = 168.069911
A503 = 1811.033911
A1003 = 3461.290496
B03 = -10428.738977D-06
B503 = 84312.562567D-07
B1003 = 23003.272361D-06
C03 = 38760.727267D-11
C503 = 27665.705167D-10
C1003 = 52381.759251D-10
(--different page in printout: )
A3 = -20.072216E-02
B3 = 10.125706E-01
C3 = -11.101592E-06
D3 = 21.629247E-10

## SBE19-1031 LostAtSheba
SBE19-1031 |  | Pump: SBE5-01
LOST at SHEBA - 1998/08 (with transmissometer and borrowed pump… )
box with manual, download cable, pump, etc. should have been returned to FWI?!
Temperature | date:
# 1031 | a=
b=
c=
d=
fo=
Conductivity | date:
# 1031 | a=
b=
c=
d=
m=
Pressure | date:
# 154898 | A0
(300 psia) | A1
A2
Pressure | date:
# 177965 | A0
(1000 psia) | A1
A2
Pressure | date:
# 184833 | A0
(3000 psia) | A1
A2
Pressure | date:
# 189759 | A0
(5000 psi) | A1
A2
SERVICE: