## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 18-Dec-2013 | Changed units for Flag:Oxygen:Dissolved channel. J.L. |
| 31-Oct-2005 | Added nutrient data to the rosette files. J.L. |

## PROCESSING NOTES

Cruise: 2005-07

Agency: OSD, PGC, CHS

Location: Barkley Sound/Effingham Inlet

Project: Multi-beam Surveys, Paleoclimate

Party Chief: Dallimore A.

Platform: Vector

Date: April 18, 2005 – April 22, 2005

Processed by: Germaine Gatien

Date of Processing: July 11, 2005 – August 2, 2005

Number of original CTD casts: 22

Number of CTD casts processed: 22

Number of original rosette casts: 22

Number of rosette casts processed: 22

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with Chelsea/Seatech transmissometer (#333DR), Altimeter OA-916D (#1024), a SeaBird dissolved oxygen sensor (#0766) and a Seapoint Fluorometer (#2356) with a 10X cable. The deck unit was a model 911 (#0424) and the logging computer was F602. The salinometer was an 8400 Autosal on loan from Guildline.

# SUMMARY OF QUALITY AND CONCERNS

The CTD daily log had no list of sensors used. For one rosette cast there was a log sheet with sample numbers but no indication of what specific sampling was done.

The dissolved oxygen data in the CTD files should be considered:

- ±0.4ml/l in the top 50m

- ±0.3ml/l from 50db to 200db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Note was made that the file for cast #4 was misnamed. The files with names 2005-07-0001 were renamed as 2005-07-0004.

The rosette log sheets have a few entries about bottles that misfired.

Nutrient data was not yet available.

The dissolved oxygen data were obtained; there were no flags or comments but there is no indication on the rosette sheets that any were required.

There was no chlorophyll sampling.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no errors found.

Configuration file 2005-07-0004.con was copied to 2005-07-ctd.con and this was used for all casts.

##### Conversion of Raw Data

All files were converted to CNV files.

A few casts were examined.

The PAR and SPAR, while included in the CON file, appear not to have been connected.

All other expected channels contain reasonable data. The altimetry is noisy but looks reasonable near the bottom. The pairs of conductivity sensors are reasonably close. For cast #26 the upcast is somewhat unlike the downcast, but during a stop for a bottle the upcast values drift close to the downcast for both temperature and conductivity sensors. The upcast has lower temperature and higher salinity, which fits the explanation that the offset in values is due to the rosette wake carrying deeper water with it.

The DO data includes very low values as this cruise sampled anoxic waters. In the past this has led to poor data after sampling anoxic waters, so editing may be required to remove unreliable DO data.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format and the extensions were changed to BOT.

There was a ROS files for cast #16 but according to the CTD log event #16 was not a CTD cast; there was a cast #17 that had been saved as #16. The file name was changed to #17.

All BOT files were plotted and no outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

All casts were rosette casts with frequent stops on the upcast. They are also shallow casts, so it is impossible to pick out a suitable section for the usual test for CELLTM setting. The results of cruise 2005-08 which used the same equipment will be applied to this data.

CELLTM was run using (0.02, 7) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

All casts for 2005-07 were shallow and little can be learned from the differences between sensors. The same sensors were used for 2005-08 which was run shortly after this cruise; the differences between the pairs of T and C sensors found during 2005-08 and the conclusions reached are shown below.

Results of 2005-08:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 370 | -0.006 | -0.00055 | -0.005 | Excellent |
| 74 | 370 | -0.004 | -0.00045 | -0.005 | ok |
| 74 | 1070 | -0.0005 | -0.0004 | -0.004 | ok |
| 116 | 370 | ~0 very noisy | -0.0005 | -0.0053 | ok |
| 116 | 1070 | -0.0005 very noisy | -0.00045 | -0.0046 | ok |
| 116 | 1480 | -0.0004 very noisy | -0.0004 | -0.0046 | ok |
| 119 | 370 | -0.0003 | -0.0005 | -0.0053 | ok |
| 119 | 1070 | -0.0003 | -0.0005 | -0.005 | ok |
| 119 | 1760 | -0.0002 | -0.00035 | -0.005 | ok |

The salinity differences are large but reasonably steady with time and pressure. The temperature and conductivity differences are small; there may be some drift with time or pressure, but it is not large enough to be clear.

Data was plotted for a few casts of 2005-07 and there are no obvious problems. The salinity difference at 175m looks a little smaller than found during 2005-08 with values around -0.0035 to -0.004psu. The temperature and conductivity differences are extremely noisy at those depths.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add an Event Number based on the last 4 digits of the file name, and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run, and the cruise track was plotted. The only problem was that the CTD cast that should be named as #17 was called #16. The files for that cast were renamed.

The average surface pressure is 2.1db, which is reasonable for the Vector. There was one cast with negative pressures at the end of the upcast. That data was examined and the negative values were at pressures very close to zero and with the pumps off, so this does not indicate a problem.

A few casts were checked to ensure the header got the right reading from the altimeter. The algorithm worked well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

Sample numbers were then added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

Dissolved oxygen files (*.add) created by the analyst had a flag channel.

The salinity data was received in an EXCEL spreadsheet. Sample numbers and station name were given, but not event numbers. Those were added, as was a quality flag column (there were no quality concerns.) One error in sample number was corrected. There was a sample #74 in the salinity analysis, but no sampling was indicated on the rosette sheet; the sample number was assumed to be correct, but checks will be made later to ensure the number is correct. The resulting file was saved as 2005-07 Salinity.csv. That file was converted to individual SAL files.

There was no CHL sampling. Nutrient data were not yet available.

The SAL and ADD files were merged with SAMAVG in two steps. A format error in one of the DO flags had to be fixed to enable merging. (Output: MRG1, MRG)

##### 11. COMPARE

Salinity

The salinity samples are all from above 200db so the large scatter is large. The two deepest bottles show differences of +.0122/+.0100 and +.0069/+.0045 with the CTD reading high for both pairs of sensors. There is a sufficiently large gradient at the level of those bottles to expect errors on the order of ±0.005psu. There is too much scatter to detect drift with time and the only outlier that really stands out is from only 10db in a strong gradient region.

The results confirm that sample #74 is probably labelled correctly.

Dissolved Oxygen

COMPARE was run. The best fit was with the differences between bottles and CTD versus CTD DOX. When outliers were excluded for fit residuals>0.3, the following fit was found:

CTD-BOT = 1.367 DOX-CTD - 0.018

During the last 4 uses of this sensor the results were:

CTD-BOT = 1.361 DOX-CTD + 0.025 (2005-08)

CTD-BOT = 1.463 DOX-CTD + 0.111 (2005-02)

CTD-BOT = 1.493 DOX-CTD + 0.056 (2005-01)

CTD-BOT = 1.359 DOX-CTD + 0.069 (2004-37)

The differences were fairly flat with time. There are no severe outliers. Because anoxic waters were sampled which are known to cause some drift in calibration, groups of casts were plotted together to see if the use of a single calibration equation would be a problem. There are variations, but nothing as striking as was seen in the 2005-01 data. It is difficult to compare DO data from different cruises, but this cruise has similar results to those of 2004-37, suggesting there has not been significant drift. (See 2005-07-dox-comp1.xls.)

Plots of CTD DO and Titrated DO versus CTD Salinity from the bottle files were examined. One bottle DO is slightly out of line with both the salinity and CTD DO (cast #32, sample #161, ~50db), so that sample was flagged “c” and a note put in the header.

##### Other Comparisons

Previous experience with these sensors – The primary sensors were used for 3 shallow casts of 2005-02 when no bottle comparisons were available. They were also used for the cruise which followed this one, 2005-08, when there were only 7 bottles, one of which was an extreme outlier. The differences between CTD and bottles were scattered, but suggest that the CTD primary salinity was high by about 0.006psu and the secondary was high by less than 0.001psu. That result was consistent with the differences between channels for 2005-08.

Historic ranges – All casts are in inlets or very close to shore, regions likely not well represented in the climatology. Salinity fell below the historic limit for many of the casts, but this is not considered significant.

Post-cruise calibration- A post-cruise calibration of the secondary conductivity sensor indicates a drift of +0.0007psu per month at conductivity = 3 S/m. If the drift were linear with time this would suggest that the salinity was high by about 0.0015 at the time of this cruise. If the drift is linear with use it would be high by no more than ~0.0004psu.

##### 13. SHIFT

Because there are no casts without stops on the upcast, it is difficult to test the alignment of the fluorescence and dissolved oxygen. The same equipment was used during 2005-08, when there were some non-rosette casts. The same parameters were chosen for this cast after tests to ensure they did not make things worse.

Fluorescence

A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Dissolved Oxygen

Tests were run using +120, +140, +160 and +180 records on a few casts. The choice of +140 records looked reasonable, though it was very hard to judge with all the stops on the upcast.

A shift of +140 records was applied to all casts. (Output: SHFO)

Conductivity

Tests on 3 casts showed there were slight improvements when the 2005-08 settings were applied.

All casts were put through SHIFT using -0.2s and -0.5s. (Output *.SHFC1 and SHFC2).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The descent rate was low near the bottom for many of these casts. Files were checked to ensure that the drop rate feature did not remove too much data in the range from 20db above the maximum pressure to 10db above the maximum pressure. No problems were found.

All DEL files were copied to EDT files.

##### DETAILED EDITING

There is no significant difference in noise or spikes in the two channels. The secondary salinity appears to be closer to the bottles for both this cruise and 2005-08, so the secondary sensors were selected for archiving.

Page plots were produced using (T1,S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required light editing, mostly at the surface and bottom.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

The salinity is probably within 0.001psu of the bottles; given the problems in the comparison for this cruise and for 2005-08, no recalibration is justified.

The Oxygen:Dissolved:SBE channel was recalibrated in the MRG and SAM files using the relationship found in section 11:  DOX(CORRECTED) = 1.367 DOX(ORIGINAL) -0.018

COMPARE was rerun and when only points with -0.3<residuals <0.3 are included, the average difference is -0.0004 ml/l. The edited downcast files, EDT, were also recalibrated.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Dissolved Oxygen comparison

The averaged downcast files were thinned and the SBE DO values were compared with the upcast bottle DO values. The fit against Bottle DO values was flatter than that versus pressure, with the following trendline when residuals were used to exclude outliers:

DOX-BOT = 1.004 * DOX-CTD – 0.1453.

An almost constant offset between the downcast CTD DO and upcast bottle DO has been seen from other cruises in the past, but usually when plotted against pressure. (See 2005-07-dox-comp3.xls)

- Final DO calibration

A second recalibration, 2005-07-recal2.ccf, was applied to the thinned downcast files based on the results in the previous section. The output files were then compared to the upcast bottles. The resulting fit against Bottle DOX gave an average difference of only 0.0002ml/lm. (See 2005-07-dox-comp4.xls.) The fit against pressure is not as good, with average differences of 0.007ml/l, but the differences are not large. The recalibration was then applied to the AVG files.

Based on scatter in the COMPARE plots the following errors are estimated for the CTD dissolved oxygen data:

- ±0.4ml/l in the top 50m

- ±0.3ml/l below that

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Pump Status, Descent_Rate, Altimeter, PAR, Surface PAR and Flag channels. (Output: *.REM)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and

unedited except that some records were removed in editing T and S

The dissolved oxygen data in the CTD files should be considered:

±0.4ml/l in the top 50m

±0.3ml/l from500db to 200db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Pump Status, Altimeter, PAR, Surface PAR and Flag for all casts.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, fluorescence and transmissivity profiles.

Cross-reference lists were produced for the final CTD and CHE files.

Particulars

4. File original name was 2005-07-0001, changed to 2005-07-0004.

16. Renamed as #17.

20. No record on rosette log sheet of what sampling was done.

Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-07 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 18 April 2005 End: 22 April 2005 |  |  |  |  |  |  |
| Location: Barkley Sound/Effingham Inlet |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Dallimore A |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-07

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 16/10/04 | Factory “ |  |  |
| Conductivity | 2102 | 15/10/04 | “ |  |  |
| Secondary Temp. | 2968 | 01/02/05 | “ |  |  |
| Secondary Cond. | 1729 | 03/02/05 | “ |  |  |
| Transmissometer | 333DR | 30/03/05 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 16/11/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |