## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2005-08

Agency: OSD, UW, NOAA

Location: Strait of Georgia, WCVI, Juan de Fuca Strait

Project: La Perouse, Juan de Fuca, BIO, Sill, ECOHAB, Aquaculture

Party Chief: Juhasz T.

Platform: John P. Tully

Date: April 25, 2005 – May 7, 2005

Processed by: Germaine Gatien

Date of Processing: July 11, 2005 – August 11, 2005

Number of original CTD casts: 109

Number of CTD casts processed: 109

Number of rosette casts: 28

Number of rosette casts processed: 28

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#333DR) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#0766) and a Seapoint Fluorometer (#2356) with a 10X cable were both mounted on the primary pump. The deck unit was a model 911 (#0424) and the logging computer was #FS02. The salinometer was an Autosal on loan from the factory. A thermosalinograph (SeaBird 21, S/N 2487) was mounted with a fluorometer.

# SUMMARY OF QUALITY AND CONCERNS

The log book was in good order.

There were a few problems with the rosette sheets. There were 6 bottles fired during cast 106, but there is no record of what sampling was done. There was another cast with multiple bottles and no rosette sheet, but note was made in the CTD log that only salinity sampling was done.

The secondary temperature sensor for the TSG (inlet temperature) gave bad data, so heating due to the ship was estimated from other TULLY cruises.

There was no loop sampling so the TSG salinity was recalibrated based on comparisons with the CTD.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 100m

- ±0.3ml/l from 100 to 400m

- ±0.1ml/l below 400m – 1000m

- the instrument is considered unreliable below 1000m, but errors appear to be less than 0.1ml/l in the 1000-1500m range.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Nutrient data were unavailable. Salinity data were obtained; these had no flag channel.

The dissolved oxygen data were obtained; flags and comments had been entered.

There was no chlorophyll sampling.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no errors found.

Configuration file 2005-08-0004.con was copied to 2005-08-ctd.con and this was used for all casts.

##### Conversion of Raw Data

The files were converted to CNV files. File 2005-08-0067 could not be converted. The error message referred to the header length being longer than the buffer. Examining a copy of this file in Ultraedit and Notepad showed there are records but they are full of null values. The DAT file on the CD-ROM backup and on Tom Juhasz’s computer are also empty of real data. Doug Anderson was able to find the original files on the computer on which it was originally acquired. Those files appear to be fine.

A few casts were examined.

All expected channels contain reasonable data.

The descent rate is quite low near the bottom – care will be needed in the use of DELETE. Perhaps the layer for which it is not applied should be increased to 15 or 20m. Consider this later.

The down and upcast temperature and conductivity channels are reasonably close, though the two channels vary more on the upcast than the downcast, and there is an offset in the upcast data in high gradient regions suggesting poorer flow to the sensors. We expect wake problems on the upcast.

Transmissivity looked fine with good agreement between downcast and upcast.

The dissolved oxygen had the usual sort of offset; there are very low values at the bottom of some casts which may cause problems with the sensor. Editing may be required for some casts.

The fluorescence was occasionally very noisy and sometimes off-scale, but generally looked ok.

The altimetry is often extremely noisy but looks excellent near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s. Cast #67 could not be converted (as noted above) but a copy of the file from the shipboard computer was obtained and converted.

All files were then converted to IOS HEADER format. CLEAN was used to add event numbers to the header. (Output: BOT)

All BOT files were plotted and no significant outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5	Points per block = 50

##### CELLTM

Tests were run on 3 casts using settings (0.01,7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5). The best results overall were with (0.02, 7) though the differences were very small among the last 5 choices.

CELLTM was run using (0.02, 7) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 370 | -0.006 | -0.00055 | -0.005 | Excellent |
| 74 | 370 | -0.004 | -0.00045 | -0.005 | ok |
| 74 | 1070 | -0.0005 | -0.0004 | -0.004 | ok |
| 116 | 370 | ~0 very noisy | -0.0005 | -0.0053 | ok |
| 116 | 1070 | -0.0005 very noisy | -0.00045 | -0.0046 | ok |
| 116 | 1480 | -0.0004 very noisy | -0.0004 | -0.0046 | ok |
| 119 | 370 | -0.0003 | -0.0005 | -0.0053 | ok |
| 119 | 1070 | -0.0003 | -0.0005 | -0.005 | ok |
| 119 | 1760 | -0.0002 | -0.00035 | -0.005 | ok |

The salinity differences are large but reasonably steady with time and pressure. The temperature and conductivity differences are small; there may be some drift with time or pressure, but it is not large enough to be clear.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run. No errors were found.

The average surface pressure is 1.4db, which is a little lower than usual. A few casts with low surface readings were examined and data in the top 1db have what look like near-surface values; the pumps are not turned on until 2 or 3db which looks about right. The pressure sensor was recalibrated recently. No further calibration is necessary.

A few casts were checked to ensure the header got the right reading from the altimeter and the values look good.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

There were 9 bottles fired during cast #144 that have no sample numbers assigned. Numbers 901 through 909 were assigned as place holders, no analysis results were found. If salinity results should turn up for those samples, these numbers should either be changed in the ADDSAMP file and SAM and SAMAVG files be re-derived, or those sample #s can be added to the salinity files.

There is no rosette sheet for cast #106 – there were 6 bottles fired there. It is likely that only salinity was sampled, but that is not noted in the CTD log.

Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data was provided in EXCEL format. Channel names were changed to standard format and columns were added for event number and Flag:Salinity:Bottle. The file was saved as 2005-08 Salinity.csv. The files were converted to individual SAL files.

The bottom line of the ADD file for cast #82 contained unreadable information; there was a bottle fired, but no sampling done, so this line was removed from the file. The same was true for one line in each of  casts #111 and 122.

Nutrient data were not yet available.

The SAL and ADD files were merged with SAMAVG in two steps. (Output: MRG1, MRG)

##### 12. COMPARE

Salinity

There were only 7 salinity bottles available. The differences varied from +0.001 to +0.0085psu for the primary, and from -0.004 to +0.0032 for the secondary. The deepest value was out of line with the others for both sensors. The salinity value was flagged “c” and a comment added to the header. The average of the 6 remaining bottles is +0.0060 and +0.0008psu for the primary and secondary, respectively, and the trend-lines are reasonably flat with pressure, though the scatter is large.

Dissolved Oxygen

COMPARE was run. As usual, the best fit was with the differences (CTD DOX-BOT DOX) versus CTD DOX. When outliers were excluded by using fit residuals as indicators, the following fit was found:

CTD-BOT = 1.361 DOX-CTD + 0.0253

During the three most recent uses of this sensor the results were:

CTD-BOT = 1.367 DOX-CTD  - 0.018 (2005-07)

CTD-BOT = 1.493 DOX-CTD + 0.056 (2005-01)

CTD-BOT = 1.359 DOX-CTD + 0.069 (2004-37)

The results are similar to those of 2004-37 when a moderate range of DO values was experienced. Cruise 2005-01 was very complicated with four different recalibration schemes used for different casts; sampling anoxic waters was an issue. 2005-07 also sampled some anoxic waters but the effect was not as noticeable. For 2005-08, the lowest DO bottle value was 0.017ml/l. When values of DO <0.17ml/l were excluded from 2005-07, the relationship was

CTD-BOT = 1.352 DOX-CTD + 0.048 (2005-07 – excluding low DO)

It is not a simple matter to compare different cruises, but there does not seem to be any significant drift in the calibration of this sensor since late 2004. (See 2005-08-dox-comp1.xls.)

Plots of CTD DO and Titrated DO versus CTD Salinity from the bottle files were examined. The only severe outlier other than one that was flagged “d” by the analyst was sample #50, cast #73, at 250db. That sample was also a severe outlier in COMPARE so a flag “d” was added to it, along with a comment in the header. There were other outliers, but they are all from the top 30db where there are large DO gradients.

##### Other Comparisons

Previous experience with these sensors – The primary sensors were used for 2005-07 just before this cruise and for 3 casts during 2005-02 in February. There are no bottle comparisons for the latter. For 2005-07 the bottles were all from above 200db. Compared to the lowest bottle at 197db the CTD primary salinity was high by 0.006psu and the secondary salinity high by 0.0005psu.

Historic ranges –Almost all data fell within the historic ranges. Exceptions were small excursions towards low salinity and high temperatures near shore and near the bottom. These may reflect real variations, or waters not previously sampled, but do not look like evidence of poor calibration.

Post-cruise calibration – A post-cruise calibration of the secondary conductivity cell was obtained. It shows that there was a drift of +0.0007psu per month at conductivity = 3S/m. If the drift is linear with time this would suggest that the salinity was high by about 0.002psu. If the drift is linear with use the drift would vary from 0.0004psu to 0.0023psu through the cruise; this assumes that I am aware of all cruises for which this sensor was used in 2005, and that may not be the case.

##### 14. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were about 1.5s or higher. However, at the depth at which there is a significant fluorescence signal, the ascent rate was generally much more uneven than the descent rate so much of the offset is probably due to upcast-only problems. A shift of +24 records (1s) is the shift that has been used in most other cruises, so it was chosen for this cruise too. (Output: SHFFL)

Dissolved Oxygen

Tests were run using +120, +140, +160, +180 records on a few casts. As well as making the trace offset similar to those of temperature we also judge by how the DO looks during a bottle stop. We expect the DO to settle to the correct value by the end of the bottle stop. That value should be roughly half way between the up and down values in motion. The results vary with DO gradient, so there is no clear RIGHT answer. However, the best choice overall appears to be to advance the DO channel by +140 records, which was also the choice made for this equipment for 2005-09 later in May 2005. This shift was applied to all casts.

Conductivity

Tests were run on a few casts to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.5 for the primary and -0.2 for the secondary. All casts were put through SHIFT twice using -0.5s for the primary and -0.2s for the secondary. (Output: *.SHFC1 and SHFC2).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

There are some casts with descent rate slowing significantly more than 10db above the maximum pressure. A few were checked to ensure that too much data had not been removed by DELETE; no cases were found where the drop rate feature was activated inappropriately. Each file will be examined during the editing step and if there are problems they will be detected then.

##### DETAILED EDITING

There is no significant difference in noise levels between the pairs of sensors. Since the secondary sensors appear to produce salinity that is closer to the bottles, that pair were selected for archiving.

Page plots were produced using (T1, S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts requiring editing, mostly from the surface and bottom; only casts #114 and 116 required heavy editing. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

There is not a lot of calibration information for salinity, but what we have shows that the secondary salinity is high. From the bottles we estimate that it is high by about 0.0008psu and from the post-cruise calibration by about 0.0013psu at the mid-point of the cruise. So 0.001psu will be subtracted from the salinity.

From section 12 we have the following equation for recalibration of DOX:

CTD DOX (corrected) = 1.3605*CTD-DOX + 0.0253

File 2005-08-recal1.ccf was prepared to apply the recalibrations to salinity and dissolved oxygen. That was applied first to the SAM and MRG files.

COMPARE was rerun for secondary salinity and the average difference (excluding the same outlier as in the first comparison) was -0.00015psu. (See 2005-08-sal-comp2.xls.)

COMPARE was rerun for dissolved oxygen and when only points with -0.3<residuals <0.3 are included, the average difference is -0.001ml/l. (See 2005-08-dox-comp2.xls.)

The edited downcast files, EDT, were then recalibrated.

##### Special Fluorometer Processing

There was no titrated chlorophyll data, so no special files were prepared for A. Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure		Averaging interval = 1.000

Minimum bin value =   .000	Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison

The averaged downcast files were thinned and the CTD DO values were compared with the upcast bottle DO values. The differences were fairly flat when plotted against the titrated values, the CTD reading high by about 0.08ml/l. For 2005-07 with the same instrument this offset was to 0.15ml/l but it sampled much shallower water. When plotted against pressure there is a fairly convincing trendline, but it is not as tight as that versus DOX-BOT. (See 2005-08-dox-comp3.xls)

21. Final DO Calibration

File 2005-08-recal2.ccf was prepared to adjust the CTD DOX channel using:

CTD DOX (corrected) = 0.9966 * CTD DOX – 0.0784

It was applied to the THN1 and COR1 files. COMPARE was rerun and the results show a good correspondence between downcast CTD DOX and upcast bottles. This recalibration was not applied to the bottle files. (See 2005-08-dox-comp4.xls.)

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 100m

- ±0.3ml/l from 100 to 400m

- ±0.1ml/l below 400m – 1000m

- the instrument is considered unreliable below 1000m, but errors are less than 0.1ml/l for the deep water of this cruise.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Altimeter and Flag channels. (Output: *.REM)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and

unedited except that some records were removed in editing T and S

The dissolved oxygen data in the CTD files should be considered:

•	±0.5ml/l in the top 100m

•	±0.3ml/l from 100 to 400m

•	±0.1ml/l below 400m – 1000m

•	the instrument is considered unreliable below 1000m, but errors are

believed to be no more than 0.01ml/l in the 1000-1500m range.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag for all casts.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared using the edited data and displaying T and S. Profile plots were produced showing temperature, DO, fluorescence and transmissivity.

Cross-reference lists were produced for the final CTD and CHE files.

##### 25. Thermosalinograph Data

a.) Checking calibrations

There was 1 file containing TSG data. A report was printed for the con file and the calibrations were checked for temperature and conductivity. No errors were found. The primary sensors were recalibrated in Dec. 2004, the secondary temperature in August 2004. When last calibrated the drift in primary salinity was found to be +0.0008psu per month and the temperature drift about -0.0012Cº. If drift has continued at the same rate, we would expect the temperature to be low by about -0.007Cº (which would lead to a salinity value that is too high by 0.006psu) and a salinity error due to conductivity drift on the order of +0.01psu. However, the conductivity drift between the 2003 and 2004 calibrations was of the opposite sign, so we should not assume it will be the same for this deployment.

TSG data has been processed from 2 other cruises that occurred after the latest recalibration.

b.) Converting to IOS Headers and adding position headers and time channels.

The data was converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, temperature:Difference, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

Note that UPloy0 was the actual channel name and that it contains the flow rate.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

Time-series plots were produced and no severe problems noted.

There is only one spike in each of the primary temperature and salinity channels. The secondary temperature shows no such spike so pad values were entered into the primary temperature and salinity channels for record #3877 of file 2004-25-0001.atc.

c.)  Checking Time Channel

The times and positions of the first and last CTD casts were found from the CTD Daily Log and those times found in the ATC file. The positions were compared and found to agree very well. As there has been no problem with the timing of this instrument in the past year, this was considered a sufficient check.

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are no obvious problems this time, so this step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 For a first comparison the differences were plotted versus time and they were very large, on the order of 2 to 3Cº. In the past the difference has been on the order of 0.2Cº. The primary temperature was very noisy while the intake temperature was quite smooth.

- TSG vs CTD The CTD data after DELETE and metre-averaging was thinned to reduce the files to a single point at or near 4db and exported to a spreadsheet. This step was done before the CTD data was edited, but it turned out that there was little editing applied at 4db, so this should not be a significant source of error. (2005-08-atc.xls)

The ATC file was opened in EXCEL and reduced to a single header line and the channels of interest (Date, Time, T0, T1, Sal, Flow, Fluor, Lat, Long). All records were removed EXCEPT those that correspond to the time of a CTD cast. The two files were then combined as 2005-08-TSG-CTD-comp.xls.

Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. The remote TSG temperature was higher by from 1 to 5Cº. The primary TSG temperature was higher than the CTD temperature but the differences were very noisy. For casts #37 through 50 the differences were fairly quiet, at about 0.3Cº. Picking out 8 stations with low gradients near the surface, the TSG primary temperature was higher than both CTD temperature channels by about 0.36Cº. The TSG intake temperature was higher than the CTD temperature by about 3.0Cº. The remote TSG temperature appears to be bad. This was reported to Doug Anderson, who says there are a number of possible causes for this problem. A quick glance at the data from 2005-12 which followed this cruise, suggests the remote temperature sensor data was ok then, so the sensor itself is probably not the cause of trouble.

Using the same 8 casts, the TSG salinity was found to be lower than the uncalibrated CTD salinity by about 0.036psu, but after the CTD salinity is recalibrated, that would be 0.035psu.

The same data was used to compare the fluorescence from the two instruments. The fluorescence from the TSG was very low, with CTD FL = 14.2 * TSG Fl – 1.13. This is similar to results from other cruises.

- Loop Bottle Comparisons There were no loop samples.

f.) Calibration History

The TSG was recalibrated in December 2004. It was used for 2005-01 during which there was a lot of noise in the salinity comparison, and concerns about the Portasal performance. The surface gradients were very low so vertical matching of TSG and CTD data was not critical. The temperature was found to be high by about 0.2Cº and the salinity low by about 0.03psu ±0.01psu. The same recalibration was found appropriate for 2005-02. The temperature range was about 7 to 9ºC for those cruises.

It is believed that the salinometer malfunctioned for the analysis of 2005-01 and 2005-02 bottles. The 2005-01 loop bottles were analyzed in two separate sessions with different errors. Applying the estimated errors to the results leads to estimates that the TSG was low by 0.042 or 0.034psu. So using 0.04psu is reasonable, but should be considered ±0.01psu.

During 2004 the TSG was used for 2 September cruises during which the T adjustment was -0.13Cº. The near-surface temperatures were in the 10 to 18ºC range.

Conclusions

The intake temperature is bad, so cannot be used to deduce the effects of heating due to the ship. The best we can do is to apply the results of other Tully cruises using the same equipment, but the amount of heating is somewhat dependent on the intake temperature. So a comparison was made of the near-surface temperature range during all the relevant cruises

During this cruise the surface temperatures were in the 9 to 13ºC range. Given that the temperatures for this cruise are about midway between the February and September cruises, an offset of -0.16Cº seems appropriate.

The TSG Fluorescence is very noisy, but appears to be low by a factor of about 14.

f.) Editing

There are a few large spikes in T and S, with individual temperature values out by at least 1Cº and as much as 6Cº. CTDEDIT was used to interpolate T and S for 10 such records. (Output: ED1)

g.) Recalibration

CALIBRATE was used to apply offsets of -0.16Cº and +0.035psu to Temperature:Primary and Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Temperature:Secondary, temperature:Difference, Conductivity:Primary, Flag, UPloy0 (flow rate)

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, add the depth of sampling to the header and change the fluorescence channel name to FLUORESCENCE and the corresponding units to VOLTS.

The sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

Particulars

45 – Weights on; no heave compensator

52 – Turbid at bottom

65 – Heave compensator used

67 – Heave compensator used; file contains only null data

70 – Jelly fish; lot of cleaning done afterwards

109 – According to log saved as 108 but no files found for either 108 or 109. Actually saved as 107. Name changed to 109.

114 – 9 salinity bottles at 500m

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-08 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 25 April 2005 End: 07 May 2005 |  |  |  |  |  |  |
| Location: SoG/JdeF/WCVI |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Juhasz T. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-08

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 16/10/04 | Factory “ |  |  |
| Conductivity | 2102 | 15/10/04 | “ |  |  |
| Secondary Temp. | 2968 | 01/02/05 | “ |  |  |
| Secondary Cond. | 1729 | 03/02/05 | “ |  |  |
| Transmissometer | 333DR | 30/03/05 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 16/11/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2005-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 4/12/04 | Factory |  |  |
| Conductivity | 2487 | 4/12/04 | “ |  |  |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab WetstarFluorometer | ? | ? | ? |  |  |
| Flow Meter | ? | ? | ? |  |  |