## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27_May_2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2005-09

Agency: OSD

Location: B.C. Inlets

Project: Inlets, Multibeam etc.

Party Chief: Spear D

Platform: VECTOR

Date: 25 May 2005 – 29 May 2005

Processed by: Germaine Gatien

Date of Processing: 11 July 2005 –9 August 2005

Number of original CTD casts: 28

Number of casts processed: 28

Number of rosette casts: 8

Number of rosette casts processed: 8

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlabs CSTAR transmissometer (#333DR), a Seabird Dissolved Oxygen Sensor (#0766), Altimeter (S/N 1024) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911+ (#0424). The salinometer was a Autosal on loan from Guildline.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Log Book was in good order. Details about what sampling was done was missing from some of the rosette log sheets.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 200m

- ±0.25ml/l from 200m to 500m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

There was no chlorophyll sampling. The titrated dissolved oxygen data were obtained with flags and comments added by the analysts.

The nutrients were not ready at the time of CTD processing.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There was an error in the pressure calibration date and an inappropriate pressure offset entered. In previous use a setting of +0.4db has been used, but in the con file -0.6db was used. A con file was prepared with +0.4db and saved as 2005-09-CTD.con.

The history of the conductivity, dissolved oxygen and pressure sensors was found.

##### Conversion of Raw Data

The raw data were converted using configuration file 2005-09-ctd.con. A few casts were examined and all expected channels are present. The two channels track well, especially during the downcast. The up and downcasts are reasonably similar. The secondary conductivity is noisier than the primary.

The altimetry was noisy but there is a clear signal near the bottom. Transmissivity looks reasonable. Fluorescence was off-scale near the surface for some casts and there is some spiking at depth. The dissolved oxygen shows the usual evidence of hysteresis.

The descent rate was generally very steady.

Rosette files were converted using a start time of -2s and duration of 5s.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.

##### CELLTM

Tests were run on a few casts with (alpha, 1/beta) set to (0.1, 7), (0.01, 9), (0.02, 9), (0.03, 9), (0.02, 7) and (0.03, 7) and (0.0245,0.5). The changes were very slightly, but (0.03, 9) looked best overall. CELLTM was run using (0.03, 9) on both conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 4 | 510 | -0.0001 | +0.00012 | +0.0014 | Steady, moderate |
| 10 | 521 | +0.0005 | -0.00002 | -0.0008 | “ |
| 21 | 474 | +0.00035 | +0.00015 | +0.0012 | “ |
| 28 | 450 | ~0 X Noisy | ~+0.0002XN | +0.002 | Noisy, moderate |

The differences in temperature were highly variable. The salinity differences, while also variable, are never large. None of these casts were very deep and a lot of variability is expected in the area, so these results are not evidence of poor performance of sensors.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers to the headers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files. CLEAN was run to add event numbers and the output files were named BOT.

All BOT files were plotted and no significant outliers were found.

##### Checking Headers

The cruise track was plotted and no problems noted.

A header summary, header check and cross-reference listing were produced. The only discrepancy was in one latitude – the log is presumed to be wrong since the bottom and stop positions are close to those in the headers.

The average surface pressure is 1.5db which is reasonable for the Vector.

The altimetry header data was exported to a spreadsheet. A few casts were examined and all values are reasonable.

##### 10. SHIFT

Conductivity

Tests were run on two casts with various shifts of conductivity and the best results overall were with an advancement of -0.2 records to the primary and -0.5 records for the secondary.

All casts were put through SHIFT using those settings.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between the two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent + ascent rate to find the shift (in seconds) to remove that offset. For this cruise the results varied from +1.3s to +1.5s. The only depths at which the offset is clear had significantly variable upcast descent rate so the results are approximate. The value that has been used for most cruises in the past is +1s (+24 records); that was applied to this data as well.

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. Tests were done shifting the DO by +80 to +160 records and the best results were found using +140 records. In recent cruises using this instrument, shifts of from +120 to +160 records have been used. A shift of +140 records was applied to all casts.

##### 11. BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. The addsamp file was edited to add sample numbers taken from the rosette sheets. There is a note on the rosette log sheet for cast #2 that the bottle #s did not match the log entries; a bottle was fired at 220db that is not noted on the sheet. The confirmed pressures are close to what is actually in the bottle file. The addsamp file was used to add sample numbers to the BOT files with output named *.SAM. These were then averaged to create SAMAVG files. One line was removed from the SAMAVG file for cast #26, because, as noted on the rosette log, there was no sample from the bottle at 10db.

The ADD files had a few problems – cast #5 is missing. There was a repeat titration of one sample which must have caused a problem in the conversion. Cast #16 has an ADD file but no data. The formats were fixed in the OXY files for those two casts and new ADD files created using the ADD CHANNEL program. Sheila Toews agreed with these changes. There are no flags. The ADD file for cast #21 had to be rearranged in order of sample number.

The salinity analysis spreadsheet was saved in CSV format. There were some results with station name, but no sample number. That information was found from the rosette log sheets and transferred to the CSV file. There was no flag channel, so one was added along with a comment column. Notes from the salinity analysis sheet indicate that there was no liner for 4 of the samples; those were flagged “c”. Sheila Toews agreed with these flags being applied. The CSV file was then converted to SAL files. The SAL file for cast #19 had to be rearranged in order of sample number.

They were then merged with the ADD and SAL files in two steps. (Output: MRG1 and MRG)

The nutrient data was not yet available.

##### 12. COMPARE

Salinity

COMPARE was run. There were only 14 bottles, of which 4 were from above 200m and 4 had no liners. The latter were all outliers in COMPARE. When those 8 bottles were excluded the primary salinity was found to be higher than the bottles by an average of 0.0022psu and the secondary was high by 0.0032psu. The deepest bottle shows the primary to be high by just 0.0003psu with the secondary high by 0.002. There is no evidence of time or pressure dependence, but with so little data this is not conclusive. (See 2005-09-sal-comp1.xls.)

The flags for the 4 bottles without liners were changed from “c” to “d” and the comments updated.

Dissolved Oxygen

COMPARE was run and plots were prepared of differences between CTD and titrated values versus pressure, DO and cast number. There were no severe outliers in COMPARE. Excluding 6 points with residuals greater than 0.4 produced a reasonably tight fit with

DOX (Titrated) = 1.3115* DOX (CTD) + 0.2403

When used in April 2005 for cruise 2005-08 the relationship was

DOX (Titrated) = 1.361* DOX (CTD) + 0.0253

There is no obvious evidence of time dependence. (See 2005-09-dox-comp1.xls)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

All the DEL files were copied to *.EDT.

##### DETAILED EDITING

The primary sensors were chosen for archiving since there is less noise and the calibration is slightly better than for the secondary.

Page plots were produced. These were used to guide the editing.

The descent rate was very steady for most casts. All casts required light editing to remove records from the surface, and in a few cases to remove records corrupted by shed wakes; salinity was cleaned lightly.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensor was used for 2004-29 when it was thought to be low by about 0.003psu, and for two cruises after that when there were insufficient calibration samples to determine accuracy.

Historic Ranges – Local climatology is only available for a few casts. The temperature was above the historic maximum for cast #26 between 100 and 150m, but this is not likely to be due to calibration problems; this cast was close to shore and is probably not represented well in the climatology.

Factory calibrations – The primary conductivity sensor was recalibrated in August 2005 at which time there had been a drift downwards by about 0.001psu. The secondary temperature and conductivity sensors were recalibrated just before this cruise; the secondary temperature sensor is one that has drifted more than average over the past few years despite repairs done to it.

##### Initial Recalibration

There were only 6 useful bottles from below 200m; those suggest that the primary salinity is high by at most 0.002psu, but the deepest bottle suggests the error is much smaller than that. The post-cruise calibration indicates a drift that would result in salinity low by about -0.001psu. No recalibration will be applied to the salinity.

File 2005-09-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of section 12.

Corrected DOX = 1.3115* DOX (CTD) + 0.2403

This was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for SAL and DOX and the corrections were found to be made correctly. (See 2005-09-sal-comp2.xls and 2005-09-dox-comp2.xls.)

### Final Dissolved Oxygen Comparison

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but in the past there have sometimes been significant errors due to response-time problems. To check for such errors the downcast CTD data (after SHIFT and CALIBRATE) was compared with the bottles from the upcast. A set of downcast files were prepared by bin-averaging (0.25db bins) the recalibrated downcast files and thinning the data to the depths generally used for bottles. COMPARE was run comparing those files with the bottle data in the MRG files. The fit of differences versus pressure was very flat when severe outliers were excluded. The CTD is higher than the bottles by an average of +0.055ml/l. (See 2005-09-dox-comp3.xls.)

##### Special Fluorometer Processing

There was no chlorophyll sampling so no special files were prepared for Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the full COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

(Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

File 2005-09-recal2.ccf was prepared to subtract 0.055ml/l from the Oxygen:Dissolved:SBE channel.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared.

Separate profile plots were prepared with Temperature, Transmissivity, DOX, and Fluorescence versus pressure.

##### FINAL CTD files steps (REMOVE, HEADEDIT and CALIBRATE)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag channels were removed from all casts.

Oxgyen:Dissolved:SBE was calculated in umol/kg and that channel added to the files.

HEADER EDIT was used to fix formats and units and to add the following comment using file 2005-09hdr.txt:

Transmissivity: The data are unedited except where records were

removed in editing temperature and salinity.

Fluorescence:URU:Seapoint - The data are unedited except as above.

The dissolved oxygen data are unedited except as above.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 200m

- ±0.25ml/l from 200m to 400m

The Standards Check routine was run and HEADER EDIT adjusted and rerun until no further problems were found. These files were named CTD.

A cross-reference listing was produced.

The sensor history was updated.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird.

SORT was used to put the data in order of increasing pressure.

REMOVE was used to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag for cast #1.

REMOVE was used to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag for casts #2-68.

Oxgyen:Dissolved:SBE was calculated in umol/kg and that channel added to the files.

HEADER EDIT was run to fix channel names and formats and to add a standard comment including an explanation of the quality flags. The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The output files were named CHE.

A cross-reference listing was produced.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2005-09

Dates: Start: May 25, 2005 End: May 29, 2005

Location: B.C. Inlets Vessel: Vector

Party Chief: Spear D

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585               Cruise ID#:		2005-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 01/02/05 | Factory |  | Factory |
| Conductivity | 2173 | 19/08/04 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 2424 | 07/04/05 | “ |  |  |
| Transmissometer | 333DR | 30/03/05 | IOS |  |  |
| Dissolved Oxygen | 0766 | 16/11/04 | Factory |  |  |
| Fluorometer | 2229 |  |  |  |  |
| Altimeter | 1024 |  |  |  |  |
| Pressure Sensor | 77511 | 30/09/99 | Factory |  |  |