## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected Salinity:Bottle lost during addition of HPLC. S.H. |
| 19 January 2021 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2005-10

Agency: OSD

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG - JdeF

Party Chief: Masson D.

Platform: Vector

Date: April 11, 2005 – April 16, 2005

Processed by: Germaine Gatien

Date of Processing: 18 October 2005 – 7 November 2005

Number of original CTD casts: 68 	Number of CTD casts processed: 67

Number of rosette casts: 22		Number of rosette casts processed: 21

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#333DR) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#0766) and a Seapoint Fluorometer (#2356) with a 10X cable were both mounted on the primary pump. The deck unit was a model 911 (#0508) and the logging computer was #FS03. The salinometer was an Autosal on loan from the factory.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book was in good order.

One salinity sample appears to be mislabelled and one had no sample number on the label, or could not be read by the analyst. The rosette sheet lacked any evidence of salinity sampling for that cast.

Some of the nutrient samples appear to have been mislabelled.

Salinity should be considered ±0.002psu.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 100m

- ±0.2ml/l from 100 to 200m

- ±0.1ml/l below.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. There is mention of problems with time; the computer differed from the NMEA data – the computer screen shows both times.

Nutrient data were obtained with flags and comments.

Salinity data were obtained; these had no flag channel, but the salinity analysis sheets were available. The dissolved oxygen data were obtained; flags and comments had not been entered. These were added by Doug Anderson in files named ADD.

The chlorophyll data (with flag channel and comments) was obtained.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no errors were found.

Configuration file 2005-10-0004.con was copied to 2005-10-ctd.con and this was used for all casts.

The data for two casts were missing, but were found by Doug Anderson in the wrong folder on the acquisition computer.

##### Conversion of Raw Data

The files were converted to CNV files

A few casts were examined.

All expected channels were found and contain reasonable data except for cast #68. That cast was mentioned in the log as having the syringes left on. The temperature, salinity, fluorescence, dissolved oxygen and transmissivity all look bad. This cast will not be processed.

Two other casts (33 and 67) had notes in the log about syringes being left off the CTD; it is unknown if this means between casts or if the operator was just being careful about removing the syringes for the cast and making note of the fact. The data from the casts before and after these comments look ok, but should be examined carefully later.

The descent rate is excellent for some casts and terrible for others. It was noted that stops at some casts were 30s or more, but for some casts there were some very short stops and in at least one case the CTD never came to a complete stop.

The down and upcast temperature and conductivity channels are reasonably close, though the two channels vary more on the upcast than the downcast.

Transmissivity looked fine with good agreement between downcast and upcast.

The dissolved oxygen had the usual sort of offset; there are very low values at the bottom of cast #1 which may cause problems with the sensor.

The fluorescence was occasionally very noisy and sometimes off-scale, but generally looked ok.

The altimetry is often extremely noisy but looks excellent near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format. CLEAN was used to add event numbers to the header. (Output: BOT)

All BOT files were plotted and no significant outliers were found, though the secondary salinity does look a little noisier than the primary.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

There were few casts deep enough, without stops for bottles and with a steady descent rate to use for testing these parameters. Cast #48 was put through CELLTM using settings (0.01,7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7) and (0.03, 9). The best results overall were with (0.02, 7) though the differences were very small. During 2005-08 which used the same equipment and was run shortly after this cruise the same setting was found to be best.

CELLTM was run using (0.02, 7) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 170 | -0.0002 | -0.00023 | -0.002 | Mod., steady |
| 11 | 170 | -0.0004 | -0.00035 VN | -0.003 XN | Mod, steady |
| 48 | 170 | ~0 X Noisy | -0.0004 | -0.004XN | Mod, steady |
| 48 | 300 | ~0 X Noisy | -0.0005 | -0.004 | Mod, steady |
| 52 | 170 | ~0 | -0.00045 | -0.0045XN | Mod, steady |
| 52 | 300 | ~0 | -0.00045 | -0.0045 | Mod, steady |

The salinity differences vary with time but show no drift with pressure. The upcast differences are extremely noisy. The time variation is not completely clear. When cast #31 was examined it was found that just before a stop at 200m the differences were extremely noisy, then as the CTD stopped they dropped to about 0.0025psu. Then they went up to ~0.007psu, then settled again to about 0.0035psu after 20s, at which point the CTD started up again and the differences became extremely noisy again. This suggests that the differences may reflect local gradients and pump irregularities more than drift in calibration. One pump may work better than the other, a feature that would be exaggerated when the descent rate is irregular and when the local gradients are large.

The temperature and conductivity differences are small, but extremely noisy at times. There may be evidence of some drift with time, but it is difficult to judge that from such shallow casts. The results are very close to those from the middle of the 2005-07 and 2005-08 which immediately followed this one. There was no evidence of significant time or pressure difference during those cruises.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run. The only inconsistencies found were between time in the log and time in the header. There are notes in the log book that the computer time and NMEA time disagreed by one hour. The headers sometimes match the times in the log book, and at other times do not. The question is which time to trust.

A few cases were found when a watch change appears to have taken place, based on initials in the log book. In all cases the times that make sense are those taken from the headers, i.e. NMEA time. This is what we expect to be correct. There are two different times that can be read from the computer screen, one from NMEA and one from the computer itself. The log entries were probably based on computer time sometimes and NMEA at other times. The times in the files are fine.

The average surface pressure is 1.9db which is reasonable for the Vector.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Cast #68 had a rosette file, but the log indicates there was no sampling. The syringes were left on the CTD, so the file will not be processed further.

Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data was provided in EXCEL format. There was no flag channel. The salinity values were linked to other columns, so they were copied using Paste Special to preserve the values. Channel names were changed to standard format and columns were added for event number and Flag:Salinity:Bottle, and event numbers were added. The file was saved as 2005-10 Salinity.csv. There are three samples shown on the analysis sheet for cast #7, but no such sampling appears on the Rosette sheet. Two of those bottles were labelled as sample #14 and one has a “?” for the sample number. All are said to be from BOT 2. Looking at the salinity values compared to those of the CTD rosette salinity, it appears that one of the samples called #14 should be #13, from the bottom, and one does look like it is #14. The “?” sample looks like a duplicate of #14. The sample number for the bottle that looks like Bottle #1 was changed to #13, and an average of the other two salinity values was entered for sample #14. Notes were put in the comments including the two values that went into the average, and a “c” flag was entered for both samples. This should be pursued when COMPARE is run. The files were converted to individual SAL files.

There were two ADD files for cast #13 due to a computer glitch and the first file has a comment that 3 bottles were missing – they were actually in the 2nd file. The two files were merged into a single file and the comment that 3 bottles were missing was edited. The quality flags were in the wrong column.

A spreadsheet was obtained with the data and flag channels. The nutrient samples were stored frozen and later analysed at IOS by Mellissa Hennekes and Wendy Richardson. A few problems were found in the files:

- For cast #22 there was a sample labeled as a duplicate of sample # 68, but the values suggest it might be a duplicate of #66. The values from the questionable sample were entered as a comment for sample #68 and the original entry removed from the spreadsheet. Janet Barwell-Clarke agreed to this and the following changes and asked that a flag “c” be added to the PO4 flag channel with the comment that it looked low.

- A similar situation was found for cast #46, sample #147 duplicate. Again the values were entered as a comment for the other sample from that bottle and the original entry was dropped from the spreadsheet.

- Cast #66, sample #264 was treated in the same way.

The channel names were edited to standard format, samples with no analyses were removed, files were re-ordered on sample number and converted to NUTS files.

The extracted chlorophyll data was in spreadsheet format and included a flag channel. The only comment applied to all samples, so it was entered once for each cast in the comments column: “Average of two samples is reported.” (Output: CHL)

The SAL, CHL, ADD  and NUTS files were merged with SAMAVG in four steps. (Output: MRG1, MRG2, MRG3 and MRG) (An extra line had to be removed from 2005-10-0066.samavg; this was a bottle that was fired but no samples were taken and no sample # assigned.)

##### 12. COMPARE

Salinity

There was a lot of scatter in the results which is not terribly surprising given many were from fairly shallow water, but it does make it difficult to have faith in the results. When 8 outliers were excluded the fit of differences was quite flat with the primary CTD salinity being high by about 0.001psu and the secondary low by 0.0025psu. When differences were plotted against file pair number the primary differences are quite flat, but there might be a drift in the secondary with the difference from bottles increasing.

Given the huge scatter, the six deepest bottles were examined in detail to see which might be considered reliable. The differences given are for the primary channel only.

- #1 – 172m: The stop was at the bottom and was very long and there was no motion of the CTD during the stop. The CTD appears to have been slowly sinking. The standard deviation of the CTD salinity was 0.0002. Local salinity gradient was 0.001psu/m. Difference = -0.0001psu.

- #16 – 200m: The descent rate was extremely noisy. The stop was long enough, but the CTD was moving a lot during the stop. The standard deviation of the CTD salinity was 0.0005. Local salinity gradient measured during the downcast was -0.0001psu/m. Difference = -0.002psu.

- #28 – 348m: There was almost no stop, the CTD slowed to near zero, but immediately took off again. There is slight evidence of shed wakes, but the gradient was fairly low. It looks like the CTD salinity might have gone down a little if there had been a longer wait. The standard deviation is 0.0003. Local gradient +0.0005psu/m. Difference = +0.005psu.

- #31 – 200m: There was a stop of about 20s. The temperature was still changing when the CTD started upwards, and the descent rate was very noisy. The standard deviation was 0.0016. Local gradient -0.003psu/m. Difference = 0.008psu.

- #44 – 300m: The wait was short, maybe 15s, but the gradient was low and T and S seem to have equilibrated. Standard deviation was 0.0002. Local salinity gradient -0.0009psu/m. and difference = -0.002psu.

- #53 – 250m: The wait was only about 10s, but again the conditions were quiet, so the CTD may have been in equilibrium. Standard deviation in CTD salinity was 0.0003. Local salinity gradient -0.004psu/m. Difference = -0.0007psu.

The most reliable data would appear to be from cast #1 with a long stop and little salinity variation. Two other casts where we expect little variability with time would be casts #28 and 31, but they have higher local gradients at the depths of those particular bottles. Cast #28 had almost no stop and #31 shows a lot of variability; both are in fairly high salinity gradients.

Casts #16, 44 and 53 are in active mixing zones, so while the local gradients are lower the temporal variability is higher. The data indicate that the primary CTD salinity was low by 0.002, 0.002 and 0.0007psu respectively, while the secondary salinity was low by 0.004, 0.005psu and 0.006psu. Those results are lower than the average found in COMPARE. But if the rosette did not have time to flush properly and it contains water of salinity higher than the ambient salinity, that would make the CTD appear to be reading low. However, for cast #16 the salinity was decreasing with pressure around 200m so the opposite would apply, so this may not be a valid explanation.

The results of the tests of differences reported in section 7 for cast #31 during the stop at 200m suggests a longer wait before firing the bottles would achieve better results. When the bottle was fired the differences were about -0.006psu, but a few seconds later they were down to -0.0035 and might well have gone lower, but the CTD then started upwards.

The only extreme outlier was from the surface bottle of cast #1. Examining the full CTD file shows extremely high salinity variability at the surface, so the bottle results will not be flagged.

Dissolved Oxygen

COMPARE was run. A few outliers were found and investigated.

- Cast #13, Sample #45 was an extreme outlier. On the log sheet there is a note that there was a bad endpoint, and no value was entered, so the value was replaced in the MRG file with a pad value. Original value was entered as a comment in the header.

- Cast # 13, Sample #34 was an outlier, and there was no value entered in the rosette sheet, so this was also replaced by a pad value. Original value was entered as a comment in the header.

- Cast # 63, Sample #259 was an outlier, and there was no value entered in the rosette sheet, so this was also replaced by a pad value. Original value was entered as a comment in the header.

- Cast #66, sample #274 was an outlier, but from very close to the surface, so it was not flagged.

As usual, the best fit was with the differences (CTD DOX-BOT DOX) versus CTD DOX. When the outliers mentioned above were excluded, the following fit was found:

CTD-BOT = 1.3723 DOX-CTD + 0.040

Because the CTD sampled anoxic waters during the first cast the fit of the first few casts is expected to be different. Using a fit with all casts included and gradually removing the early casts suggests that by cast #7 the differences are similar to the later casts. Using the first two casts for which there are bottles, namely casts #1 and 3 the following fit was found:

CTD-BOT = 1.3988 DOX-CTD - 0.1129

The fit for all bottles except those from the first 2 casts, excluding outliers was:

CTD-BOT = 1.3705 DOX-CTD + 0.049

The first fit should be applied up to cast #3 and the second for cast #7 onwards; the casts between those two have no DO calibration samples, so based on Line P experience it was guessed that cast #5 would be a good point to go to the general fit.

During two recent Line P cruises, a negative offset was found after anoxic sampling and the calibration drifted back to one with a positive offset. So this behaviour is similar.

During the most recent uses of this sensor the results were:

CTD-BOT = 1.359 DOX-CTD + 0.069  (2004-37)

CTD-BOT = 1.493 DOX-CTD + 0.056  (2005-01)

CTD-BOT = 1.367 DOX-CTD  - 0.018  (2005-07)

CTD-BOT = 1.361 DOX-CTD + 0.0253 (2005-08)

The only one which has a negative offset, was another with anoxic sampling, but in that case there was insufficient sampling to enable splitting the recalibration.

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. No further outliers were found.

##### Other Comparisons

Previous experience with these sensors – The primary sensors were used for 2005-02, -07 and -08 but the only useful bottles were 6 from 2005-08 which indicated that the primary was high by 0.006psu. The secondary sensors were used during 2005-08 and were found to be high by about 0.0009psu. But these differences were based on very little information.

Historic ranges – Most of the data fell within local climatology limits, but there were two areas where there were excursions from those ranges. Towards the northern shore of the Strait of Georgia the salinity was lower than the minimum around 40db. And near the mouth of Juan de Fuca Strait the salinity was low and the temperature sometimes high. Given the nature of the area it is assumed that these are real variations and not evidence of calibration problems.

Post-cruise calibration – There were post-cruise calibrations of both conductivity sensors and the secondary temperature sensor. If the drift were linear with time then the primary conductivity would be high by about 0.0035psu, whereas if linear with the number of casts on which it was deployed, it is likely to be high by ~0.0015psu. The secondary would be considered high by ~0.0015psu using the time approach, but it may be less than that since this was its first use after calibration. The temperature error would have a negligible effect.

This would lead us to expect differences on the order of .001 to 0.002psu. During the Saanich Inlet cast when the descent rate was very steady the differences are about 0.002psu, with the primary higher than the secondary. But we see differences much larger than that during most casts. In COMPARE the secondary is farthest from what we expect based on the post-cruise calibration, so perhaps the secondary pump is not working as well.

##### 14. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were from 0.8 to 2s. A shift of +24 records (1s) is the shift that has been used in most other cruises, so it was chosen for this cruise too. (Output: SHFFL)

Dissolved Oxygen

Tests were run using +120, +130, +140, +150 and +160 records on a few casts. As well as making the trace offset similar to those of temperature we also judge by how the DO looks during a bottle stop. We expect the DO to settle to the correct value by the end of the bottle stop. That value should be roughly half way between the up and down values in motion. The results vary with DO gradient, so there is no clear RIGHT answer. However, the best choice overall appears to be to advance the DO channel by +140 or +150 records. The former was the choice made for this equipment for 2005-09 and 2005-10 which immediately followed this cruise. So +140 records was applied to the DO channel for all casts.

Conductivity

Tests were run on a few casts to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. There was very little change but the best settings proved to be -0.5 for the primary and +0.2 for the secondary. All casts were put through SHIFT twice using -0.5s for the primary and -0.2s for the secondary. (Output: *.SHFC0 and SHFC1).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salinity

Maximum Surface Pressure (relative): 10.00 Minimum Salinity: 5

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

##### DETAILED EDITING

It is not obvious which sensors to archive. The primary sensors are a little less noisy than the secondary, and the COMPARE results are closer to the post-cruise calibration than those of the secondary, but the post-cruise calibration and previous use suggest the secondary is better. A few casts were examined in areas where little temporary variation was expected and the salinity during upcast bottle stops seemed closest to the downcast if the primary sensors were used. This may be an issue of the better pump, rather than the better calibration.

Page plots were produced using (T0, S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing. The following casts required heavy editing: 23-25, 28-31, 33, 56-59.

Note was made of the editing details in the headers of the relevant files.

An adjustment was made to the Altimeter entry in the header of cast #37 to reflect that the bottom 3db of data had been removed from the file.

##### Initial Recalibration

The results of the salinity comparison are scattered and some of the data are considered unreliable; the history is not very useful. The primary looks better than the secondary in COMPARE, but the secondary looked better during 2005-08. The post-cruise calibration suggests both are a little high, with the secondary probably better. No recalibration will be applied and salinity should be considered ±0.002psu.

From section 12 we have the following equation for recalibration of DOX:

CTD DOX (corrected) = 1.3988*CTD-DOX -  0.1129 (Casts 1-4)

CTD DOX (corrected) = 1.3705*CTD-DOX + 0.0490 (Casts #5-end)

File 2005-10-recal1.ccf was prepared to apply the recalibrations of dissolved oxygen. That was applied first to the SAM and MRG files.

COMPARE was rerun for dissolved oxygen and when only points with -0.3<residuals <0.3 are included, the average difference is -0.006ml/l. (See 2005-10-dox-comp2.xls.)

The edited downcast files, EDT, were then recalibrated.

##### Special Fluorometer Processing

All files were put through CLIP to produce files  with data to 100db only for the use of A. Peña. Those files were then bin-averaged (1/4db bins), put through REMOVE to remove extraneous channels and HEADEDIT to fix formats and channel names. The final files (FCTD) were saved in a separate directory.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure		Averaging interval = 1.000

Minimum bin value =   .000	Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison

The averaged downcast files were thinned and the CTD DO values were compared with the upcast bottle DO values. When plotted against pressure and removing outliers identified by residuals, a fairly flat trendline is found with an offset of about 0.062ml/l with the CTD reading high. When plotted against CTD DO values there is a similar results. (See 2005-10-dox-comp3.xls)

21. Final DO Calibration

File 2005-10-recal2.ccf was prepared to adjust the CTD DOX channel using:

CTD DOX (corrected) = CTD DOX – 0.062

It was applied to the THN1 and COR1 files. COMPARE was rerun and the results show a good correspondence between downcast CTD DOX and upcast bottles. This recalibration was not applied to the bottle files. (See 2005-10-dox-comp4.xls.)

The dissolved oxygen data in the CTD files should be considered:

- ±0.6ml/l in the top 100m

- ±0.2ml/l from 100 to 200m

- ±0.1ml/l below 200m

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE [umol/kg], Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to re-derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and

unedited except that some records were removed in editing T and S

The dissolved oxygen data in the CTD files should be considered:

•	±0.6ml/l in the top 100m

•	±0.2ml/l from 100 to 200m

•	±0.1ml/l below 200m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Oxygen:Dissolved:SBE [umol/kg], Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to re-derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared using the edited data and displaying T and S. Profile plots were produced showing PAR, DO, fluorescence and transmissivity.

Cross-reference lists were produced for the final CTD and CHE files.

Particulars

1. Problems with altimeter – problem was in con file, data is fine.

33. Comment in log: “Syringes off CTD” - data looks ok.

46. Transmissivity went to 0 at 81m (bottom of cast).

67. Log note:  “Syringes left off CTD” but data looks ok.

68. Syringes left on CTD. Bad Temp, Sal, FL, DO and Trans data.  Delete cast.

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-10 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 11 April 2005 End: 16 April 2005 |  |  |  |  |  |  |
| Location: SoG/JdeF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Masson D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-10

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 16/10/04 | Factory “ |  |  |
| Conductivity | 2102 | 15/10/04 | “ |  |  |
| Secondary Temp. | 2968 | 01/02/05 | “ |  |  |
| Secondary Cond. | 1729 | 03/02/05 | “ |  |  |
| Transmissometer | 333DR | 30/03/05 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 16/11/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| PAR | 4656 | 11/02/03 |  |  |  |
| Surface PAR | 16504 | 02/01/04 |  |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |