## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 15 November 2005 | TSG Fluorescence recalibrated using manufacturers 2001 calibration |

## PROCESSING NOTES

Cruise: 2005-11

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Mackas D.

Platform: John P. Tully

Date: May 23, 2005 – May 31, 2005

Processed by: Germaine Gatien

Date of Processing: September 15, 2005 – 10 October 2005

Number of original CTD casts: 44    Number of CTD casts processed: 44

Number of original bottle casts: 22   Number of bottle casts processed: 22

Number of Thermosalinograph files: 4

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#732DR), a SeaBird DO sensor #0047 (on the primary pump) and a Seapoint Fluorometer (#2356) with a 10X cable (on the secondary pump.) The deck unit was a model 911 (#0508). The logging computer was FS01. The salinometer was an Autosal on loan from the factory. A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a fluorometer.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order, with notes entered about problems encountered. There was a cruise report from the chief scientist which was very helpful. There was no list of scientific personnel in either the log or the cruise report; the chief scientist is not always available when the processing is being done, and it is useful to know who else is likely to have information about the cruise.

There is a continuing problem of formatting errors in the titrated dissolved oxygen files. The flags were all in the wrong column. There should be only one space skipped between the DO value and the flag.

Fluorescence and transmissivity are nominal and unedited except that some records were removed in the editing of salinity and null transmissivity data was removed from part of cast #24. The fluorescence was occasionally off-scale.

The dissolved oxygen data in the CTD files should be considered

- ±0.6ml/l from 0 - 200m

- ±0.2ml/l from 200 – 300m

- ±0.1ml/l from 300 - 1300m

- data from below 1300m is considered unreliable by the manufacturer, but the errors appear to be less than 0.1ml/l for this data set.

The CTD calibrations used at sea for the thermosalinograph were correct for the first file, but there was then a reboot and the files used thereafter had entries from an out-of-date calibration.

There is an 11-hour gap in the TSG data between the first two files due to time/NMEA not  updating.

The Thermosalinograph salinity may be low; the results of 2005-12 were used to recalibrate because the comparison for 2005-11 showed so much scatter. When more cruises using this equipment have been processed, the recalibration should be revisited.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Together with the cruise report from Dave Mackas this gave an excellent summary of problems encountered during the cruise. Note was made on May 28th of a drift in the temperature differences and an apparent blockage in the secondary system; both apparently cleared. On May 30th it was noted that the NMEA time was out by 13 minutes; positions were also not exact when compared to the navigation computer.

There was mention in the rosette log of wrong station data for casts 24 and 31 – this appears to apply to the DO files only and should not be a problem.

The CTD sampled low oxygen waters in the first cast. This may affect the DO sensor on subsequent casts. This needs to be considered when recalibrating the data.

Dissolved oxygen data were obtained; flags and comments had been added.

The nutrient data was obtained (for both the rosette casts and the loop) which included flags and comments.

The salinity data was obtained, but without flags due to problems interfacing the computer with the borrowed Autosal.

Titrated chlorophyll data was obtained.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and an error was corrected in the transmissivity calibration; the corrected file was named 2005-11-ctd.con.

##### Conversion of Raw Data

All data were converted using file 2005-11-ctd.con.

A few casts were examined and all expected channels are present. The pairs of sensors look reasonably alike during the downcasts, but the upcast secondary data is a little odd. The fluorescence is unusually noisy which may be an indication of trouble with the secondary pump. The transmissivity has a reasonable shape and the dissolved oxygen looks offset between upcast and downcast in the usual way.

Rosette files were converted using a start time of -2s and duration of 5s.

The rosette files were then converted to IOS SHELL files. All ROS files were plotted and no significant outliers were found. All IOS files were renamed as BOT files.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

The same configuration was used for 2005-12. Tests were run on 2 casts with (alpha, 1/beta) set to (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03,9) and (0.0245, 9.5). For the primary the best setting was (0.03, 9), the same as found for 2005-12. For the secondary the best choice was (0.01, 9) for one cast and (0.02, 7) for the other; the results for 2005-12 were a little different though little change was noted for that cruise. It is unlikely that the secondary will be selected for archiving.

CELLTM was run using (0.03, 9) for the primary and (0.01, 9) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 25 | 925 | 0.002 | 0.00015 | ~0 noisy | High |
| 25 | 1800 | 0.0009 | 0.00018 | 0.0012 | High |
| 31 | 925 | 0.0015 | 0.0002 | 0.0008 noisy | High |
| 31 | 1800 | 0.001 | 0.00022 | 0.0017 | High |
| 49 | 925 | 0.0018 | 0.00025 | 0.001 | High |
| 59 | 925 | 0.0022 | 0.00027 | 0.001 | High |

There is some pressure dependence in the temperature and salinity differences, but not much in the conductivity differences. It is unusual to see significant dependence in the temperature differences. It was noted at sea and during 2005-12 when the temperature differences were similar, but the salinity and conductivity differences were larger. There is some evidence of time dependence in conductivity and salinity but not temperature. The upcast differences are extremely noisy.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and led to the discovery that the time/position headers were scrambled in file 2005-11-0081. This was corrected in the CNV file. It was then reconverted and put through CLEAN. No further errors were found.

The header summary was run and the entries compared with the CTD log. Three files had errors in the station names which were corrected. There is an entry in the log for cast #81 indicating that there were discrepancies between the computer times and those from the navigation computer. The latter were used for the log entries for that cast and those match the header record well. So it seems likely that the wrong times and positions were due to a problem with the SeaSave display that does not affect the file headers.

The cruise track was plotted and no problems found.

The average surface pressure is 2.1db, with a range of 0.3 to 5.8db. One cast was found with a surface pressure of -0.05db on the upcast. The transmissivity makes it clear that this was very near the surface, though the salinity still looked like in-water values. However, the pumps were not on, so this is not too surprising given the water was well mixed above 5db. The pressure sensor was calibrated in October 2004 and an offset is included in the configuration file based on that calibration.

A few casts were examined on screen. The upcasts are much noisier than the downcasts with larger differences between sensors. The transmissivity and fluorometer traces look reasonable, though the latter goes off-scale sometimes. The same thing was noted during 2005-12.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

Sample numbers were then added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

As mentioned earlier, the salinity file had no flags due to a problem interfacing a computer and a borrowed salinometer. There was a comment column with no entries and the analyst indicates that no flags were needed. A flag column was added and event numbers were added to another column. The file was then converted to individual SAL files.

Dissolved oxygen files (*.add) created by the analyst had a flag channel. The files were put through SORT (on sample number) since there were a few entries out of order. (Output: ADDSORT) There were many formatting errors in the ADD files. The flags were in the wrong column, there were extra characters in the opening line of cast #1 and the final “s” was missing from “*comments” in one case. All these errors had to be fixed before the data could be merged into the bottle files. For cast #31 sample #73 was entered as 0, but should be -99 since the analyst noted the titration went wrong and could not be restarted; a flag “d” was entered. There were 6 files which required correction due to problems with flask volume; the OXY files were corrected and saved as DOXX files which were then manually edited to remove the original DO channel. The headers were not updated and in one case there were still 2 DO channels in the headers. That file could not be processed in IOS SHELL without further editing. That was done and all 6 of the files were put through CLEAN to fix the headers. These were named ADD.

The nutrient spreadsheet was simplified and saved as QF2005-11nuts.csv. Flag comments had been added. Header names were edited to standard names and pad values were entered in spaces where no sample existed. The file was then converted to NUTS files. These were put through SORT to rearrange by sample number. (Output: NUT).

Extracted chlorophyll data became available after this stage and was added later.

The SAL, ADDSORT and NUT files were merged with SAMAVG in three steps. (Output: MRG1, MRG2, MRG)

DO was plotted against CTD Salinity. There were no significant outliers. A note in the rosette log by the analyst noted an odd DO profile for cast #38. It does look odd at the 150db bottle. The SBE DO does not show the same shape, nor do the temperature and salinity. Cast #31 has a similar look at 150db. It is not clearly bad, and the SBE DO has a significant time response delay so might miss such a feature.

##### 12. COMPARE

Salinity

COMPARE was run. There were few bottles and a lot of scatter.

Primary Salinity: When points above 100db and those with differences greater than 0.02psu were excluded, the average indicates that the CTD was low by 0.0001psu, but there was a lot of scatter and the trendline was not very flat. When records were selected to achieve a flat trendline the average showed the CTD to be high by 0.0005psu. Whatever approach was taken, the average showed the CTD salinity to be close to the bottles.

Secondary salinity: Once again there was a lot of scatter. The average difference varied somewhat depending on the method selected. When records above 100db and differences greater than 0.02 were excluded, the CTD appeared to be high by 0.001psu. When points were chosen to achieve a flat trendline the average showed the CTD to be high by 0.0015psu.

Both pairs of sensors are producing salinity that is reasonably close to the bottles, but the scatter is huge so the salinity should be considered ±0.002psu. The primary appears to be closer than the secondary and no recalibration is justified.

During 2005-12 there was also a lot of scatter. The primary was closer to the bottles and no recalibration was applied.

Dissolved Oxygen

COMPARE was run using pressure as the reference variable. The best fit was found between differences versus DO value.

For cast #31 the bottle at 150db, which was noted as a little odd when DO was plotted against salinity, is only a mild outlier in COMPARE. It will not be flagged.

For cast #35 3 points were flagged “c” by the analyst. Two of these were changed to “d” because they were severe outliers in COMPARE. The other sample looked ok in COMPARE so was not changed.

The analyst noted a problem with the O2 profile from some of the bottles from cast #38. In COMPARE the first 2 bottles do look out of line. A “c” flag was entered for those two bottles.

All casts were plotted individually. There are great variations from cast to cast due to different ranges of DO values, but there appears to be a fairly gradual change during the first 6 casts, and after that the fit settles down. Anoxic waters were sampled during cast #1; it is known that the SBE DO sensor does not do well under those conditions and it may take some time to return to a normal calibration. During 2005-01 when similar calibration drift was noted after sampling anoxic waters, a variable calibration was used. For this cruise it looks like 2 fits will do a reasonable job of calibration.

For casts #1 to 24 with the exclusion of a few outliers, the following fit was found:

CTD-BOT = 1.2142 DOX-CTD - 0.0407

For casts #31 to 93 with the exclusion of a outliers, the fit is:

CTD-BOT = 1.2662 DOX-CTD - 0.0108

During the cruise that followed (2005-12) the fit was:

CTD-BOT = 1.3423 DOX-CTD - 0.0143

(See 2005-11-dox-comp1.xls.)

##### 14. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles are usually examined to determine the difference between the vertical offsets of the temperature and fluorescence traces. However, for this data the upcast fluorescence does not look reliable. So the value that has been used for most other cruises, +24 records will be applied. (Output: SHFFL)

Conductivity

When the same instruments were used for 2005-12, shifts of +0.5s and -0.5s were found appropriate for the primary and secondary conductivity channels respectively. Tests were run on 2 casts using similar shifts and T-S plots were prepared to compare the results. Settings of from +0.4s to +0.6s worked well for the primary and from -0.4 to -0.6s for the secondary.

All casts were put through SHIFT using +0.5 records for the primary conductivity and -0.5 records for the secondary. (Output *.SHFC1 and SHFC2).

Dissolved Oxygen

During 2005-12 the dissolved oxygen channel was advanced by 150 records. That value was found appropriate for cast #49 of this cruise as well. All casts were shifted by +150 records. (Output: *.SHFO)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None.

##### DETAILED EDITING

The primary sensors were chosen for the archive. Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

The salinity was noisy near the surface. The descent rate was often noisy and low so that there was much corruption by shed wakes.

All casts required some editing. Casts requiring heavy editing were: 3, 6, 16, 20, 24-35, 42, 43, 48-67, 70-75.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – Both sensors were used during 2005-12 which followed this cruise; the primary salinity was found to be low by 0.0027psu and the secondary high by 0.0022psu, but there was a lot of scatter. The deepest bottles suggested that the primary salinity was close to the bottles and no recalibration was applied.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The only excursions of salinity from the historic ranges were in shallow water in Queen Charlotte Sound; these are probably due to the paucity of data from those sites. There were many cases of temperatures above the maxima, mostly between 20 and 70db, though for a few casts this extended to the surface. The post-cruise report mentions anomalously high surface temperatures. It is believed that these excursions represent real conditions and do not reflect errors in calibration.

##### Initial Recalibration

No recalibration will be applied to the salinity. The differences from bottles have a lot of scatter, and the average difference suggest that the CTD is within 0.001psu. When the sensors are next calibrated the drift should be examined to see if recalibration is advisable.

An initial recalibration was run on the MRG and SAM files using file 2005-11-recal1.ccf to apply the following correction to the dissolved oxygen:

Casts 1-24:   CTD-BOT = 1.2142 DOX-CTD - 0.0407

Casts 25-94: CTD-BOT = 1.2662 DOX-CTD - 0.0108

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the calibration was applied to the EDT files as well. (See 2005-11-dox-comp2.xls.)

##### Special Fluorometer Processing

The COR files were clipped to 100db. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

Files were bin-averaged to 0.5m bins for just the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences were fairly flat when plotted versus pressure and a few outliers were excluded. On average, the downcast DO data appears to be high by about 0.007ml/l. A second recalibration was applied to the downcast files only (AVG, THN1 and CLIP) to subtract 0.007ml/l. (Output: COR2, THN2 and CLIPCOR)

A final run of COMPARE was done to ensure the recalibration was done properly and it looks fine. (See 2005-11-dox-comp3.xls and 2005-11-comp4.xls.)

The clipped files were put through BIN AVERAGE (0.25db bins), REMOVE and HEADEDIT and named as *.FCTD and saved for Angelica Peña.

The SAMCOR files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Dissolved Oxygen:Voltage:SBE,Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (umol.kg) was added and the channels reordered to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.6ml/l from 0 - 200m

- ±0.2ml/l from 200 – 300m

- ±0.1ml/l from 300 - 1300m

- data from below 1300m is considered unreliable by the manufacturer,

but the errors appear to be less than 0.1ml/l for this data set.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared using the edited data. Cast #24 was edited to remove null transmissivity values between 110 and 125db and the comment section was edited. Profile plots were made displaying DO, fluorescence and transmissivity data.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

At this point the titrated chlorophyll data was received. It was merged with the MRGCLN files to create MRGCLN2 files. A quick comparison shows the titrated values to be about half those of the fluorometer.

REMOVE was run to remove Scan_Number, Flag, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump and Descent_Rate. (Output: *.MRGREM)

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together. (Output: CHE1)

HEADER EDIT was run to fix formats and units, to change the agency name and to add a comment about quality flags and analysis methods. Standards Check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### 24. Thermosalinograph Data

a.) Checking calibrations

There were 4 files containing TSG data. The first file contained up-to-date calibrations, but the others had the primary temperature and conductivity calibrations from 2003 although a calibration was done in December 2004. A configuration file with the latest calibrations was saved as 2005-11-TSG.con.

There is a note in the Daily Log that some TSG data was written to 2005-08-0002.hex, but that file was not found. It is assumed that the data file had already been  renamed. TSG data was not gathered from about 0000 to 1100 UTC on May 26 due to problems with the time/NMEA not updating.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Temperature:Difference,  Conductivity:Primary, Salinity:T0:C0, Fluorescence:URU:Wetlabs, Time:Julian, UPloy0, Latitude and Longitude.

The CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

Time-series plots were produced and a few spikes in T, S and Fluorescence were noted. There is a noteworthy shift in T, S and C values at about 6:50UTC on May 25, while the ship was in Juan de Fuca Strait. Just before that the temperature difference shifted, but went back to normal values just as the other shifts were seen. This is presumably because the change was sensed first by the secondary temperature. The temperature rose by 3 Cº in just 10 minutes, and most of that change occurred within 5 minutes. The change in the remote temperature came about 4 minutes before it was seen in the lab, so that gives a measure of the time in transit. There were loop samples before and after that shift, and a quick check does not show any remarkable change in the comparison with the TSG values. So this sudden change appears to be real.

The temperature differences are much different from those noted in 2005-08 when the intake temperature was bad; for this cruise the differences look reasonable.

c.)  Checking Time Channel-

The CTD data, after editing and metre-averaging, was thinned to reduce the files to a single point at or within 1db of 4db and exported to a spreadsheet.

The TSG files were opened in EXCEL and reduced to the times when CTDs were run. Those files were combined in a spreadsheet. The positions were compared and were very close, with average differences for latitude of 0.00003º and longitude of 0.0001º and no difference was greater than 0.0006º. This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence. (See 2005-11-TSG-CTD-comp.xls)

d.) Alignment check

Recent uses of this equipment showed no alignment problems and there is no indication of such trouble, so this step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 When the differences between intake temperature and lab temperature were plotted for the last and largest TSG file, there was a lot of noise particularly for the near-shore sections. The average difference was -0.195Cº, but in a quieter section it was -0.18Cº. The differences were larger for the first file (~0.23Cº), and similar for the second file (~0.20Cº) and slightly larger for the third file (~0.21Cº). This is more likely to reflect geographic variations than calibration drift.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. There was a lot of scatter so the averages do not look very useful.  (See 2005-11-TSG at CTD times.xls)

The TSG primary temperature was higher than the CTD by an average of 0.4Cº, but the median differences is 0.21Cº. Again using the median we find the remote temperature to be ~0.009Cº higher than the CTD.

The TSG salinity was lower than the CTD salinity by an average of 0.15psu and the median difference suggests it is low by about 0.09psu.

The fluorescence from the CTD was higher than that of the TSG by a factor of roughly 8.5.

- Loop Bottle Comparisons There were 13 loop bottles all taken while the ship was moving, and most during the first two days. The TSG was not operating when 3 of the bottles were taken. The TSG files were averaged using a 5-point running average. Then the salinity value was recorded from the time indicated in the log book for the loop sample. The differences for 2 of the bottles are very large. In one case the bottle value is obviously bad (>34psu) and in another there was a large gradient at the time of sampling, so a small difference in time may have led to a large error. The average from the remaining 8 bottles indicates that the TSG is low by 0.12psu. This is higher than expected. There is a lot of noise with values from -0.02 to -0.20psu. There is no suggestion of time dependence. (See 2005-11 Salinity-loop-comparison.xls.)

- Calibration History

The TSG was recalibrated in December 2004 and the remote temperature sensor in August 2004. It was used for 2005-01 and 2005-02 but there were some concerns about the performance of the Portasal when the loop samples were processed. The TSG was lower than the loop samples by about 0.02psu for those two cruises, but those results are believed to be low by at least 0.01psu due to salinometer errors, which would imply the TSG was low by at least 0.03psu. For 2005-08 there were no loop samples. Cruise 2005-12 which immediately followed this cruise, the TSG was found to have temperatures that were high by 0.20Cº and salinity that was low by 0.05psu. There were few loop samples and a lot of scatter.

Conclusions

The TSG intake temperature is in reasonably good agreement with the CTD. The results of this cruise are reasonably close to the one that followed it in suggesting that the primary temperature is high by about 0.20Cº due to warming in transit to the lab.

The salinity is more problematic. The estimates based on the loop samples suggest it is low by 0.12, while the comparison with the CTD shows it to be low by about 0.09. The cruises before and after this one suggest that it is low by between 0.035 and 0.05psu, but there are problems with all these comparisons. The loop samples were mostly taken in active mixing regions with significant surface gradients. The comparisons with the CTD are likely a little better, a fact confirmed by the intake temperatures being fairly close to those of the CTD.

Applying a correction of +0.5psu to TSG salinity would be consistent with the 2005-12 results. It may be a little low, but if so, both cruises could be recalibrated if a better comparison becomes available.

The TSG Fluorescence appears to be low by a factor of about 8.5.

f.) Editing

The time-series plots were examined and T and S were smoothed for a few points; comments were entered in the headers where appropriate. The output files were named *.edt.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.20 Cº and +0.05psu to Temperature:Primary and Salinity:T0:C0. After calibration the two temperature channels were compared and are closer.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Primary, Flag, UPloy0 (flow rate) and Temperature:Difference.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, the location to WCVI, add the depth of sampling to the header and change the fluorescence channel name to FLUORESCENCE and the corresponding units to VOLTS.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

The TOB files were plotted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

1.   Deck unit changed during initial test. Switched from FS03 Shuttle Case to FSO1.

H2S noted in bottles.

3.   Bottle #7 did not fire.

16. Bottle #7 did not trip. Adjusted after cast.

24. No sample from bottle #11. Error in station data in DO file.

31. Error in station data in DO file.

35. Blockage noted in secondary CTD system – believed to have cleared out on board.

On upcast overshot at both 800 and 600m and returned to the proper depth.

37. Large temperature difference noted 0.003 deg C.

38. O2 data had weird profile but titrations were ok. Check against CTD.

43. Temperature difference looked ok.

53. Bottle #17 did not trip. Spring replaced after cast.

61. Bottle #15 empty. That bottle was adjusted after cast.

75. Rosette touched bottom.

80. Incinerator may have been operating during sampling.

81. NMEA time found to be slow by 13 minutes. Positions don’t agree exactly with navigation computer. Data from navigation computer used in log book. But headers look ok so this is assumed to be a SeaSave display problem only.

93. No sample from bottle #7.

# Institute of Ocean Sciences    CRUISE SUMMARY

| Cruise ID#: 2005-11 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 23 May 2005 End: 31 May 2005 |  |  |  |  |  |  |
| Location: West Coast Vancouver Island |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Mackas D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19Mar05 | Factory “ |  |  |
| Conductivity | 3038 | 03Mar05 | “ |  |  |
| Secondary Temp. | 2023 | 07Apr05 | “ |  |  |
| Secondary Cond. | 2399 | 07Apr05 | “ |  |  |
| Transmissometer | 732DR | 22May05 | Factory |  |  |
| SBE 43 DO sensor | 47 | 16May05 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2005-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 4/12/04 | Factory |  |  |
| Conductivity | 2487 | 4/12/04 | “ |  |  |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab/wetstarFluor. | WS3S- 713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? | ? |  |  |