## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 24-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. The dates and times in the cdiac file prepared by Lizette Beauchemin were incorrect so the times from John Page's CSV file were instead. A copy of John’s CSV file was added to the Excel file in a separate worksheet. |
| 28-Oct-2013 | Merged DMS data to bottle casts from Mike Arychuk’s spreadsheet file located in the \DOC directory. For information on methods see file "Method post 1999 for Dimethylsulfide Analysis.doc” in directory \\OSD_Data_Archive\Cruise_Data\Documents\Analysis Reference Papers\ |
| 27 June 2006 | CHL data added to CHE files for 6 casts; Geographic Area corrected. |
| 9 June 2006 | Fixed CHE files to include DO Flag channel; removed unwanted comments from headers. |
| 19-May-2006 | Header edit to edit headers and add Comments to CTD files |
| 19-Jan-2006 | Added loop data to the archive. |
| 15 November 2005 | TSG Fluorescence recalibrated using manufacturers 2001 calibration |
| 10-Oct-2005 | Corrections to comments in CHE file and CHL merged. |

## PROCESSING NOTES

Cruise: 2005-12

Agency: OSD / UVic

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: May 31, 2005 – June 18, 2005

Processed by: Germaine Gatien

Date of Processing: July 15, 2005 – 31 August 2005

Number of original CTD casts: 59

Number of CTD casts processed: 59

Number of bottle casts: 27 full rosette casts, 31 with 5m Niskin, 1 fired just for water

Number of bottle casts processed: 58

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#732DR), a SeaBird DO sensor #0047 (on the primary pump) and a Seapoint Fluorometer (#2356) with a 10X cable. The deck unit was a model 911 (#0508) and the logging computer was a Shuttle X. The salinometer was an Autosal on loan from the factory. A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a fluorometer.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in good order. Additional notes from the chief scientist were very helpful.

The CTD calibrations used at sea for the thermosalinograph were not from the latest recalibration.

There were many format errors in the ADD files (dissolved oxygen); some occurred when flags were added to the dissolved oxygen bottle data files, others when comments were entered. Also where data was missing either that line can be dropped from the file, or -99 should be entered as a value. The comments were not entered correctly.

The Thermosalinograph salinity is considered ±0.02psu.

The dissolved oxygen data in the CTD files should be considered

- ±0.6ml/l from 0 - 200m

- ±0.2ml/l from 200 – 400m

- ±0.1ml/l from 400 - 1300m

- data was removed from 1300m down since it is considered unreliable there.

A test was made of the effect on salinity data of putting a fluorometer or dissolved oxygen sensor on the pumps. The results suggest that there is no degradation in the salinity when those sensors are pumped, but temperature and salinity in the test casts were highly variable which could mask a small effect.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Together with the notes from Marie Robert this gave an excellent summary of problems encountered during the cruise.

Dissolved oxygen data were obtained; flags and comments had been added. There are some formatting problems. The nutrient data was obtained which included flags and comments. The salinity data was obtained, but without flags due to problems interfacing the computer with the borrowed Autosal. No titrated chlorophyll data was available at the time of processing.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and an error was corrected in the transmissivity calibration; the corrected file was named 2005-12-ctd.con.

##### Conversion of Raw Data

All data were converted using file 2005-12-ctd.con.

A few casts were examined and all expected channels are present.

Rosette files were converted using a start time of -2s and duration of 5s.

The rosette files were then converted to IOS SHELL files. All ROS files were plotted and no significant outliers were found. All IOS files were renamed as BOT files.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on 4 casts with (alpha, 1/beta) set to (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03,9) and (0.0245, 9.5). For the primary the best setting was (0.03, 9), though (0.03, 7) and (0.0245, 9.5) were very similar. For the secondary the best choice was (0.01, 7) with no change almost as good.

CELLTM was run using (0.03, 9) for the primary and (0.01, 7) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 13 | 2000 | 0.0007 | 0.00045 | 0.0047 | V.High, fairly noisy |
| 32 | 2000 | 0.0008 Noisy | 0.00041 | 0.0042 | V.High, V.noisy |
| 52 | 2000 | 0.0007 | 0.00042 | 0.0046 | High, fairly steady |
| 52 | 4260 | 0.0004 | 0.00048 | 0.0054 | High, fairly steady |
| 66 | 2000 | 0.00075 | 0.00046 | 0.0049 | V.High, steady |
| 75 | 2000 | 0.00075 | 0.0005 | 0.0052 | High, steady |
| 81 | 2000 | 0.001 | 0.00052 | 0.0052 | High, fairly steady |

There is some pressure dependence in the temperature and salinity differences, but not in the conductivity differences. It is unusual to see significant dependence in the temperature differences. It was noted at sea. The conductivity and salinity differences show slight evidence of time dependence after cast #52, but the casts with the larger differences are quite close to shore so sampling waters with quite different ranges and gradients. These differences will have to be investigated further when salinity bottle data is available.

An initial study was made of the differences between cast #48 with fluorometer and DO sensor connected and #50 with the fluorometer and DO sensor disconnected. It is difficult to pick out comparable sections of the data with quiet descent rate and similar T and S gradients. There is no obvious difference between the two casts. The differences between downcasts and upcasts were larger for the secondary sensors than for the primary for both casts. In general, cast #48 was noisier, but this appears to be due to descent rate irregularity.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run and small corrections were made to two station names to match those in the CTD Daily log.

The cruise track was plotted and no problems found.

The average surface pressure is 1.8db, which is a little lower than usual for the Tully, but the salinity values look appropriately low. The pressure sensor was calibrated in October 2004 and an offset is included in the configuration file based on that calibration. The mixed-layer calculation shows that surface salinity bottles might be useful for many casts between #15 and #31.

A few casts were examined on screen. The upcasts are much noisier than the downcasts with larger differences between sensors. The transmissivity and fluorometer traces look reasonable, though the latter goes off-scale sometimes.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. The bottles for cast #5 were removed since they were not assigned sample numbers; 9 were for sample prep and 1 was for loop calibration.

Sample numbers were then added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

As mentioned earlier, the salinity file had no flags due to a problem interfacing a computer and a borrowed salinometer. There was a comment column with no entries and the analyst indicates that no flags were needed. A flag column was added and event numbers were added to another column. The file was then converted to individual SAL files.

Dissolved oxygen files (*.add) created by the analyst had a flag channel. There are some formatting problems, with extra spaces entered in most lines with flags, one flag in the wrong column and where no sampling was done the entries were blank, but need to be -99. The comments were entered incorrectly – the word comment needs to be preceded by an *. Each ADD file was examined and corrected as needed.

There are two ADD files with no CTD cast to match. These are files 99 and 100. The data is from extra sampling during casts #17 and 48 for the use of a particular scientist. The data will not be archived.

The nutrient spreadsheet was simplified and saved as 2005-12-nuts.csv. The flag comments from the headers were transferred to the comment column as appropriate. That file was then converted to NUTS files. These were put through SORT to rearrange by sample number. (Output: NUT)

Final extracted chlorophyll data were not available at the time of processing.

The SAL, ADD and NUT files were merged with SAMAVG in three steps. (Output: MRG1, MRG2, MRG)

DO was plotted against CTD Salinity. There were 2 outliers. The one for cast 11 had already had the DO Winkler value flagged “d” by the analyst. The one for cast #52 is only a mild outlier so was not flagged.

Cast #5 has a BOT file but the bottles were fired only to collect water and no sample numbers were assigned. This file will not be processed further. According to the notes made by the chief scientist there were two other casts for which bottles were only fired for sampling. However, sample numbers were assigned and salinity samples analyzed for all the other casts, so the MRG files will be processed.

##### 12. COMPARE

Salinity

COMPARE was run. When data above 250db and 1 outlier were excluded the primary salinity was found to be low by an average of 0.0027psu and the secondary was high by 0.0022psu; there was some pressure dependence in both pairs of sensors. There also appeared to be some time dependence, but since most of the shallowest casts were late in the cruise, it is difficult to say which is significant – pressure or time. So a few individual casts were examined more closely.

- Cast #38 had little pressure dependence and the primary sensors were very close to the bottles. The secondary sensors were high by about 0.005psu.

- Cast #52 had a lot of scatter in the differences and many points had to be excluded to get a reasonably flat trendline. When that was done the differences were similar to those found when all casts were included.

- For cast #65 many bottles were fired at 3000db. When only 1 of those was included (one with a value that was roughly equal to the average of all values except one outlier) the differences are very flat with pressure. The primary is low by about 0.002psu and the secondary is high by about 0.003psu.

When only the data from 3000db downwards (again using a single 3000db value for cast #65) there is no pressure dependence and the primary is low by about 0.0009psu and the secondary is high by 0.0042psu. A final check of time dependence was made by plotting values at 3000db (again, just one for cast #65) and the fit is quite flat; there is only data from 3000db in the middle part of the cruise.

Gradually reducing points by using residuals produces smaller and smaller differences between the primary salinity and the bottles.

Finally tests were done to see if there was a problem with some Niskin bottles. Bottles #1 and 8 show the most extreme outliers. However, when the shallowest bottle fired is excluded for each cast most of those outliers disappear. However those two bottles are the worst outliers during the multiple firing at 3000db during cast #65.

Dissolved Oxygen

COMPARE was run using sample number as the reference variable. The best fit was with the differences between bottles and CTD versus CTD DOX. First, data from below 1200db were excluded from the fit and then outliers were excluded (differences >2ml/l and then residuals >0.2); the fit found was:

CTD-BOT = 1.3436 DOX-CTD - 0.0123

When more points were excluded based on residuals being >0.1, the fit changed only slightly to:

CTD-BOT = 1.3423 DOX-CTD - 0.0143

When plotted against file pair number there is a lot of scatter and no clear trend.

During 2005-01 it was noted that the calibration changed through the cruise, probably in response to sampling hypoxic waters. While hypoxic conditions were not seen for this cruise, it was thought wise to examine a few individual casts to see if there was any significant change. Plots were made of casts 11, 17, 39, 72 and 76. The plots fall into 2 groups with the slope of the first 2 casts being somewhat lower than for the last 3. The first two casts also differed in the range of values sampled – they had a larger proportion of values in the 3-5ml/l range, and they have far fewer bottles. The latter 3 casts had very few values in that range. Plots were next made by picking out groups of casts to determine when the change occurs. Most of cast #2 and the shallower bottles from casts #11 and 17 differ from the later casts. The only other samples that are at all like them are one at 5m in Rivers Inlet and the the two shallowest bottles for cast #92 north of Smith Sound.  Attempts were made to identify whether the calibration was really different, or if there were just more errors in the comparison. The casts with the different calibration are close to shore, but not all casts close to shore look like them. There appears to be more small-scale variability for those casts, which would increase the errors, but those should not be systematic. Nonetheless, there is insufficient information to recalibrate them separately since there are too few points to identify outliers well and it is hard to put faith in the fits. So the fit will be based on the full cruise. (See 2005-12-dox-comp1.xls.)

Two outliers were identified, samples 37 and 469. The former had already been flagged “d” by the analyst so a note was added to the comments confirming that it was bad in COMPARE. The latter sample was changed from flag “c”, as assigned by the analyst, to flag “d” and an explanation was added to the comments.

##### 14. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found were around 1s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Conductivity

Tests were run on 2 casts using shifts between -0.8s and +0.8s and T-S plots were prepared to compare the results. A setting of +0.5s worked best for the primary and -0.5s for the secondary.

All casts were put through SHIFT using +0.5 records for the primary conductivity and -0.5 records for the secondary. (Output *.SHFC1 and SHFC2).

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Values from +80 to +180 were tried and the best overall match of features was with a choice of +150 records. In previous use of this instrument settings of +120 and +140 have been used.

All casts were shifted by +150 records. (Output: *.SHFO)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None.

##### DETAILED EDITING

The primary sensors were chosen for the archive. Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing. Casts requiring heavy editing were: 4-11, 26-33, 45-46.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensors were new and the secondary sensors have not been used since their latest recalibration.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The temperature was a little above the historic maximum around 50m for cast #1. For cast #75 the salinity was a little low around 500m and for #76 the temperature was high and the salinity was low between 400 and 1000db. Those casts are in an area that may not be well represented in the climatology. Casts #87 through 89 have temperatures that are high in the top 100m. These excursions from the climatology are thought to be real, not indications of the calibration being off. High temperatures have been reported in 2005 in the top 50m based on ARGO float data.

##### Initial Recalibration

No recalibration will be applied to the salinity. The differences from bottles have a lot of scatter, and the deepest bottles suggest that the CTD is within 0.001psu. When the sensors are next calibrated the drift should be examined to see if recalibration is advisable.

An initial recalibration was run on the MRG and SAM files using file 2005-12-recal1.ccf to apply the following correction to the dissolved oxygen:

CTD-BOT = 1.3423 DOX-CTD - 0.0143

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the calibration was applied to the EDT files as well. (See 2005-12-dox-comp2.xls.)

##### Special Fluorometer Processing

The COR files were clipped to 100db. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

Files were bin-averaged to 0.25m bins for just the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences were fairly flat when plotted versus pressure or DO values when only values above 1250m were included and outliers were excluded. On average, the downcast DO data appears to be high by between 0.035 and 0.06ml/l. A second recalibration was applied to the downcast files only (AVG, THN1 and CLIP) to subtract 0.05ml/l. (Output: COR2, THN2 and CLIPCOR)

A final run of COMPARE was done to ensure the recalibration was done properly and it looks fine. (See 2005-12-dox-comp3.xls and 2005-12-comp4.xls.)

The clipped files were put through BIN AVERAGE (0.25db bins), REMOVE and HEADEDIT and named as *.FCTD and saved for Angelica Peña.

The SAMCOR files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added and the channels reorder to put the two SBE DO channels together. CLEAN was run to remove SBE DO values for pressures >1300db. (Output:CLN2)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.6ml/l from 0 - 200m

•	±0.2ml/l from 200 – 400m

•	±0.1ml/l from 400 - 1300m

•	data was removed from 1300m down since it is considered

unreliable there.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD. Later, another run of HEADEDIT was used to fix the Geographic area.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data. Profile plots were made displaying DO, fluorescence and transmissivity profiles.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers and to put pad values in the dissolved oxygen channel for pressures > 1300db. REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag. (Output: *.MRGREM)

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together. (Output: CHE1)

HEADER EDIT was run to fix formats and units, to change the agency name and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. Later, a second run of HEADEDIT was used to fix the geographic area description. (Output: CHE)

The chlorophyll data is not yet available but the comments were included in the header because it is presumed they will be added later.

##### Study of effect of pumped fluorometer and DO sensor on salinity

As discussed in section 7 the fluorometer and dissolved oxygen sensor were removed for cast #50 which was a repeat of cast #48. Cast #52 was also at the same site, but much deeper. The three casts were compared in the upper 200m to see if removing the sensors from the pump improved the salinity data. There is no evidence that it made a difference, but there is evidence that both temperature and salinity changed notably through the time taken for the 3 casts (~6.5 hours).

This test might be more useful if run in deeper water with lower variability. However, the reason for running this test was because of frequent problems with noisy salinity, and that feature has recently been narrowed down to a particular CTD system which was not used for this cruise.

##### 24. Thermosalinograph Data

a.) Checking calibrations

There were 3 files containing TSG data. All con files were identical. A report was printed for the first con file and the primary temperature and conductivity calibrations were found to be from 2003 although a calibration was done in December 2004. A new configuration file with the latest calibrations was prepared and saved as 2005-12-TSG.con.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

Time-series plots were produced and no large spikes noted.

The temperature differences are much different from those noted in 2005-08 when the intake temperature was bad; for this cruise the differences look reasonable.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, was thinned to reduce the files to a single point at or within 1db of 4db and exported to a spreadsheet. The TSG files were opened in EXCEL and reduced to the times when CTDs were run. Those files were combined in a spreadsheet. The positions were compared and were very close, with average differences for both latitude and longitude of 0.0001º and no difference was greater than 0.0003º. This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence. (See 2005-12-TSG-CTD-comp.xls)

d.) Alignment check

Recent uses of this equipment showed no alignment problems and there is no indication of such trouble, so this step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 When the differences between intake temperature and lab temperature were plotted for the first TSG file, there was a lot of noise particularly for the near-shore section. The average difference was -0.185Cº. The differences between the two temperature channels were plotted for times when the ship was stopped for a CTD using file 2005-12-TSG-CTD-comp.xls. The differences are very noisy towards the end of the cruise when the ship was in Rivers Inlet and there are a few other outliers. This is presumed to be due to higher surface temperature gradients plus problems of keeping the ship on station in narrow quarters leading to corruption of surface data. If all data are included, the temperature from the lab is higher than the remote temperature by an average of 0.310Cº. When differences were examined for casts #4 to #66 only the average was 0.180Cº which is very close to the average for TSG file #1.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. Averages were calculated using casts #4 to #66. (See 2005-12CTD-tsg-comp.xls)

The TSG primary temperature was higher than the CTD by about 0.20Cº and the remote TSG temperature was higher by about 0.008Cº. The TSG salinity was lower than the CTD salinity by about 0.07psu. The excellent agreement between the TSG intake temperature and the CTD at 4m is probably due to the existence of a very well mixed surface layer. The large difference between the salinity readings is surprising since in previous use the salinity of the TSG was thought to be low by about 0.002 to 0.004psu.

Using a calculation of mixed layer depth based on salinity variation, casts were selected that were well mixed to at least 10m. Then editing notes were checked and any cast with significant editing around 4db was removed. The averages were once again calculated. The primary TSG temperature was still found to be high by 0.20Cº, and the intake temperature high by 0.006Cº, so the choice of casts appears to be appropriate. The salinity is now found to be low by 0.059psu. (See TSGvsCTD-good-conditions.xls.)

The same data was used to compare the fluorescence from the two instruments, and the fit based on casts 3-48 is CTD = 12.2 * TSG – 1.4 and the average ratio of CTD fluorescence to that of TSG is about 6.

- Loop Bottle Comparisons There were only 6 loop bottles. The salinity values were compared with the TSG files and there is a lot of scatter with values from -0.3 to +0.2psu. The average of those values indicates that the TSG is low by an average of 0.04psu, which is close to the results of 2005-08 when it was low by 03psu. There is a lot of noise with values from +0.04 to -0.08psu but no suggestion of time dependence. (See 2005-12-loop-comp.xls.)

- Calibration History

The TSG was recalibrated in December 2004. It was used for 2005-01 and -02, but there were some concerns about the performance of the Portasal when the loop samples were processed. The TSG was lower than the loop samples by about 0.02psu for those two cruises, but those results are believed to be low by at least 0.01psu due to salinometer errors, which would imply the TSG was low by at least 0.03psu. For 2005-08 there were no loop samples.

Conclusions

The TSG intake temperature is in good agreement with the CTD. The primary temperature should be lowered by 0.20Cº to account for warming in transit to the lab.

The salinity is lower than the CTD by from 0.05 to 0.07psu depending on what data is included in the comparison and lower than the loop samples by about 0.04psu. There were few samples and a lot of scatter. Applying a correction of +0.05psu to TSG salinity leaves an uncertainty of ±0.02psu when the history, loop comparison and CTD comparison are considered.

The TSG Fluorescence appears to be low by a factor of about 6.

f.) Editing

The time-series plots were examined and no data was found to be clearly bad. No editing was done.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.20 Cº and +0.05psu to Temperature:Primary and Salinity:T0:C0. After calibration the two temperature channels were compared and are closer.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Primary, Flag, UPloy0 (flow rate) and Temperature:Difference.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, the location to North-East Pacific, add the depth of sampling to the header and change the fluorescence channel name to FLUORESCENCE and the corresponding units to VOLTS.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

5 – Bottles fired, but no sample numbers assigned – 9 for sample prep, 1 for loop sample

26. Computer crashed twice at start of cast

39. Went from 2000 up to 1950, returned to bottom to fire bottom bottle

43.  6 bottles tripped, 5 for sample prep (no samp #), 1 for loop sample (samp #204)

50. DO and FL sensors unplugged and dummied off. Compare to cast #48 and 52.

52. DO and FL reconnected

71 & 73. Log note that FL was noisy

89. 7 bottles closed at 20m for John Dower, no sample #, 1 loop sample (samp #453)

# Institute of Ocean Sciences    CRUISE SUMMARY

| Cruise ID#: 2005-12 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 31 May 2005 End: 18 June 2005 |  |  |  |  |  |  |
| Location: Gulf of Alaska |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Robert M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-12

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19Mar05 | Factory “ |  |  |
| Conductivity | 3038 | 03Mar05 | “ |  |  |
| Secondary Temp. | 2023 | 07Apr05 | “ |  |  |
| Secondary Cond. | 2399 | 07Apr05 | “ |  |  |
| Transmissometer | 732DR | 22May05 | Factory |  |  |
| SBE 43 DO sensor | 47 | 16May05 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2005-12

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 4/12/04 | Factory |  |  |
| Conductivity | 2487 | 4/12/04 | “ |  |  |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab/wetstarFluor. | WS3S-713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? | ? |  |  |