## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected Salinity:Bottle lost during addition of HPLC. S.H. |
| ### 12 Jan 2021 | ### Added HPLC Data. S.H. |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2005-14

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG-JdeF

Party Chief: Masson D.

Platform: VECTOR

Date: 6 June 2005 – 10 June 2005

Processed by: Germaine Gatien

Date of Processing: 19 October 2005 – 15 November 2005

Number of original CTD casts: 74 including 1 upcast	Number of casts processed: 73

Number of rosette casts: 22				Number of rosette casts processed: 22

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlabs CSTAR transmissometer (#333DR), a Seabird Dissolved Oxygen Sensor (#0766), Altimeter (S/N 1024) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911+ (#0424). The salinometer was a Autosal on loan from Guildline.

# SUMMARY OF QUALITY AND CONCERNS

The CTD Daily Log Book was in good order.

The Rosette Log sheets had some errors and one sample number was used twice.

There was a discrepancy between the Rosette Log sheet and the salinity analysis sheet.

Many of the dissolved oxygen files had the flags in the wrong column.

The rosette stops were very short compromising the quality of the bottle comparisons.

The dissolved oxygen data in the CTD files should be considered:

- ±0.6ml/l from   0 -  50m

- ±0.2ml/l from  50 –  200m

- ±0.1ml/l below 200m

Based on an analysis of data from the past several years it is believed that the secondary pump was not working well during this cruise. Confirming this, there was a lot of noise in the secondary temperature and conductivity that lines up with minima in the descent rate even when that rate was high.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

There are errors in the syntax of some of the file names (extra 0s). These were corrected.

##### Preliminary Steps

The Log Book was obtained.

The salinity data were received in spreadsheet format with comments but no flag channel.

Chlorophyll data were received in spreadsheet 2005-14 chlarc.xls with flags and comments.

The titrated dissolved oxygen data were obtained in individual ADD files with flags and comments added by the analyst.

The nutrients were received in spreadsheet format with flags and comments added by the analysts.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There was an error in the pressure calibration date and an inappropriate pressure offset entered. In previous use a setting of +0.4db has been used, but in the con file -0.6db was used. A con file was prepared with +0.4db and saved as 2005-14-CTD.con.

Note that one file was named 2005-14-0051a – this was later changed to 2005-14-0595. After conversion the former file was dropped from the CTD processing stream since it contains only upcast data.

The history of the conductivity, dissolved oxygen and pressure sensors was found.

##### Conversion of Raw Data

The raw data were converted using configuration file 2005-14-ctd.con. A few casts were examined and all expected channels are present. The up and downcasts are generally similar, but there are a lot of spikes in the secondary conductivity and temperature that appear to be associated with minima in the descent rate even when that rate is not particularly low. There are spikes in the primary as well, but not as significant. For cast #14 there is a very large excursion that might indicate something obstructing the flow. A study done in August found serious problems with the secondary pump on system #0585 which probably explains these observations.

The altimetry was noisy but there is a clear signal near the bottom.

Transmissivity looks reasonable.

The fluorescence is unusually noisy at depth; the noise looks like that in the secondary conductivity so is probably another sign of problems with the secondary pump. The descent was generally fairly quiet and high.

Because the stops are believed to be very short for this cruise, a study was made to determine what window should be set for the Rosette files. A variety of settings were chosen (varying the width of the window and whether it started before or after firing) and then the files were converted to IOS files and averaged on the Bottle Number channel. Plots were examined to see what the average descent rate was and also to look at the scatter in the unaveraged files. Some settings were found to be poor. Others were ok, but showed a tendency to slightly positive descent rates, especially at depth. Examination of the full CTD files shows that throughout the stop the CTD slowly moved deeper. In order to achieve a zero average descent rate, the window would have to cover the time when the CTD just started upwards. This is neither practical nor wise since there might be shed wake corruption. So choosing the window with minimum noise and low, positive values at depth will have to do. The best overall results when tested on 2 casts were with start time of -2s and duration of 5s, which is the usual setting applied. The results with that setting showed more scatter than with some others, but a lower average descent rate.

Rosette files were converted with a start time 2s before firing and a 5s window. Cast #51/51a is said to be a CTD cast in the CTD Daily Log and a rosette cast in the Rosette Log. No rosette file was created in conversion, so the Daily Log is presumed to be correct. The spreadsheets and ADD files were adjusted to show the samples as being from cast #49 and the ADD file was renamed 2005-14-0049.ADD. Note that the data said to be from #51 have the correct station name for cast #49.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.

##### CELLTM

Tests were run on two few casts with (alpha, 1/beta) set to (0.1, 7), (0.01, 9), (0.02, 9), (0.03, 9), (0.02, 7) and (0.03, 7). The best setting was (0.03, 9) for one cast and (0.02 ,9) for another, with not much difference between the two. When these sensors were used during 2005-09 (0.03, 9) was selected. CELLTM was run using (0.03, 9) on both conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Two casts were plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 260 | XN | +0.0002 | XN | Steady, high |
| 83 | 260 | -0.0006XN | ~+0.0005XN | +0.001 | Steady, high |

The differences were highly variable which is not surprising if one pump was not working well. The salinity differences, while also variable, are never large.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers to the headers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files. CLEAN was run to add event numbers and the output files were named BOT. All rosette files were then plotted to look for outliers. Cast #63 was edited using CTDEDIT to clean primary salinity lightly for Niskin #1. The output was copied to 2005-14-0063.bot.

##### Checking Headers

Based on notes in the log the files named as #44 were renamed #45, and the event number was changed in the header of the CLN file for that cast.

The cruise track was plotted and no problems noted. The header check turned up no errors.

A header summary was produced and two station names were wrong. Those were corrected and the cruise track plotted again.

The average surface pressure is 2.9db which is fairly deep for the Vector. While this could be a sign that the offset was a little too high, during cast #30 there is data from 0.6db with near-surface values, but definitely no sign of out-of-water values, so the pressure appears to ok.

The altimetry header data was exported to a spreadsheet. A few casts were examined and all values are reasonable.

##### 10. SHIFT

Conductivity

Tests were run on two casts with various shifts of conductivity and the best results overall were with an advancement of -0.2 records to the primary and -0.5 records for the secondary.

All casts were put through SHIFT using those settings.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between the two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent + ascent rate to find the shift (in seconds) to remove that offset. For this cruise the fluorescence was extremely noisy making an estimate difficult, and results varied from +1s to +2s. A test shift using +24 records (the value that is generally used) improved the data from cast #52. There was a distinctive feature in the upcast secondary temperature that appears to be reflected in the upcast fluorescence, and the shift brought those to the same level lending a little confidence to the +24 setting.

SHIFT was run on all casts using +24 records.

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. Tests were done shifting the DO by +120 to +160 records and the best results were found using +140 near the surface and +160 at depth. Since a lot of these casts are relatively shallow the setting of +140 was applied. In recent cruises using this instrument, shifts of +140 records was applied.

All casts were shifted using +140 records.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. The addsamp file was edited to add sample numbers taken from the rosette sheets. It was discovered that sample #43 was assigned twice (casts #9 and 12); it was changed to #943 for Niskin #14, cast #9. This was changed in the ADD file and in the CHL and nutrient spreadsheets.  The analysts and chief scientist were informed of these changes. The addsamp file was used to add sample numbers to the BOT files with output named *.SAM. These were then averaged to create SAMAVG files.

The nutrient file received from the analyst was simplified, header names were changed to standard format, cast 51/51a was renamed 51, and the data was sorted by sample number. It was then converted into individual NUTS files.

The headers of the chlorophyll spreadsheet were changed to standard names and a channel was added for event numbers. The spreadsheet was saved as 2005-14-chl.csv and was then converted to individual CHL files. There is a general comment, which will be added to the header comment file to be appended later.

The salinity was received in spreadsheet format without a flag channel column, 2005-14-sal.xls. An event number channel was added, using the station names to establish cast #. A discrepancy was found between the Rosette sheet for cast #19 and the salinity analysis sheet and spreadsheet. The salinity sample is said to be #69 from Niskin #2 on the Rosette Log sheet, but is noted as sample #68 from Niskin #1 on the analysis sheet. A quick check suggests that the salinity value looks in agreement with the sample being from Niskin #1, but this should be examined more carefully when COMPARE is run.

There was only one comment in the spreadsheet and this referred to a cracked bottle from which there was no sampling. A value of “0” had been entered for that sample. The Salinity:Bottle value was changed to  the pad value, -99.  Sheila Toews confirmed that no other flags were needed. A flag channel was added and “i” entered for the cracked bottle. The edited spreadsheet file was saved as 2005-14-sal.csv. The CSV file was then converted to SAL files. The SAMAVG files were then merged with the SAL, CHL, ADD and NUTS files in four steps. (Output: MRG)

##### COMPARE

It is known that the stops were short during this cruise. To investigate how big a problem this might be, plots were made of average descent rate versus bottle number in the SAMAVG files. There were a number of bottles for which the average descent rate was >0.1 or <-0.1m/s. The standard deviation in the CTD files may be a useful discriminator in the comparisons of salinity. Unfortunately the natural variability of fluorescence makes it less useful there. The SBE dissolved oxygen time constant is long so the stops would not have been long enough for much equilibration to take place.

Salinity

COMPARE was run. There were 22 bottles and a lot of scatter. It is known that the stops were short during this cruise and outliers were associated with large standard deviation in the CTD salinity channel, so that was used as the standard for excluding samples from the fit. When standard deviation of 0.0022 was chosen as the maximum allowed, both channels were found to be quite flat with pressure. The primary was found to be low by an average of 0.0005psu and the secondary high by 0.0026psu. When the maximum standard deviation was gradually reduced the fit became flatter, but the average difference did not vary much. There is a slight suggestion of time dependence, but since both channels vary in almost exactly the same way this is not likely to be calibration drift. The scatter is large, the stops were short and the types of waters sampled varying greatly from one area to another, so it would not be wise to read too much into the comparison.

No flags were added to outliers because they were all associated with noisy CTD data.

The question of whether the sample from cast #19 was from Niskin #1 or Niskin #2 was investigated. The bottle salinity is closer to the bottom bottle, but it was an outlier anyway so was not included in the comparison. The salinity analysis sheet is assumed to be correct, but a flag “c” was added along with a note of explanation since there remains some doubt.

Dissolved Oxygen

COMPARE was run and plots were prepared of differences between CTD and titrated values versus pressure, DO and cast number. There were two severe outliers in COMPARE. Sample #73 from cast #19 had already been flagged by the analyst. A note was added to the analyst’s comment. Sample #184 from cast #55 was flagged “c” based on being an outlier in COMPARE and in a plot of DO versus salinity and a note was added to the header to that effect.

Excluding 6 points with residuals greater than 0.4 produced a reasonably tight fit with

DOX (Titrated) = 1.4103* DOX (CTD) + 0.0742

In two other recent uses of this equipment the relationship found was:

DOX (Titrated) = 1.3115* DOX (CTD) + 0.2403 (2005-09)

DOX (Titrated) = 1.361* DOX (CTD) + 0.0253 (2005-08)

The differences from the other cruises are likely due to the narrow range of DO values sampled during this cruise (1.3-5.4ml/l) and the fact that the stops were too short for the SBE sensor to equilibrate as much as it usually does. There is no obvious evidence of time dependence.

Since the stop at the bottom is generally longer than any others, a plot was made of the differences in the bottom bottle of each cast. The fit found was:

DOX (Titrated) = 1.4248* DOX (CTD) + 0.0385

The errors at the bottom tend to be of the opposite sign since the CTD has usually just gone through decreasing DO, though the gradient is low at these depths so the error should be small anyway; the range of DO values is low so little could be deduced from the comparison. (See 2005-14-dox-comp1.xls)

Plots of DO versus CTD salinity turned up a few outliers; all the significant outliers had already been flagged. The only one not flagged was only a mild outlier in both the DO vs CTD plot and in COMPARE. It will not be flagged.

Chlorophyll versus Fluorescence

COMPARE was used to look at the differences between the CTD Fluorescence and the extracted chlorophyll from rosette samples. There was a lot of scatter, but when the fit was forced to go through the origin the slope was 1.09 and when 5 samples with chlorophyll > 5 were excluded the slope was 0.97. Three of the bottles from which the high chlorophyll samples came were from 2db, one from 6db and one from 10db. These bottles were in high gradient areas and stops were typically less than 10s so it is quite likely that the wake effects had not completely settled down.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00       Minimum Salinity: 5

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

All the DEL files were copied to *.EDT.

##### DETAILED EDITING

The primary sensors were chosen for archiving since there is some significant noise in the secondary data and a known problem with the secondary pump. The primary calibration looks better as well.

Page plots were produced. These were used to guide the editing.

The following casts required heavy editing: 6-9, 15 and 21. The following cast required no editing: 28.
The rest of the casts required lightly editing only, mostly due to small one-sided spikes in salinity and noise at the top and bottom of the casts.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – Both pairs of sensors were used for 2005-09 when the primary salinity was thought to be high by about 0.002psu and the secondary by about 0.0032psu, but there were only 6 bottles used in the comparison. The primary sensors were selected for the archive and no recalibration was applied.

Historic Ranges – All salinity data fell within the historic ranges. The temperature was well above the historic maximum for most casts in the Haro Strait area, especially between 20db and the bottom, and slightly above normal for the southern Strait of Georgia. The chief scientist believes the data. The climatology is somewhat out of date.

Factory calibrations – The primary conductivity sensor was recalibrated in August 2005 at which time there had been a drift downwards by about 0.0001psu per month. That would imply a drift downwards of 0.0012psu over the 12 months between calibrations. The secondary temperature and conductivity sensors were recalibrated just before cruise 2005-09 in April; the secondary temperature sensor is one that has drifted more than average over the past few years despite repairs done to it.

##### Initial Recalibration

The post-cruise calibration indicates a drift that would result in salinity low by about 0.001 to 0.0012psu. The comparison with bottles shows the salinity to be low by 0.0005psu, but there was a lot of scatter in the comparison. Previous use of the sensor indicated salinity was high by about 0.002, but that was based on very few bottles. No recalibration will be applied to the salinity; it is likely within 0.001psu.

File 2005-14-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of section 12.

Corrected DOX = 1.4103* DOX (CTD) + 0.0742

This was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for DOX and the corrections were found to be made correctly. There is little pressure or time-dependence after the recalibration. (See 2005-14-dox-comp2.xls.)

##### Special Fluorometer Processing

There was chlorophyll sampling so special files were prepared for Angelica Peña. The COR1 files were clipped to 100db. They were then put through BINAVG (0.25db bins), REMOVE and HEADEDIT to correct formats and channel names and named FREM.

The SAMCOR1 files, for those casts with chlorophyll sampling, were put through REMOVE and HEADEDIT and named BOF.

The BOF and FREM files were saved for Angelica Peña but will not be archived.

A median filter, fixed size=11, was applied to the fluorescence channel in the full COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

(Output: FIL)

### Final Dissolved Oxygen Comparison

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but in the past there have sometimes been significant errors due to response-time problems. To check for such errors the downcast CTD data (after SHIFT and CALIBRATE) was compared with the bottles from the upcast. A set of downcast files were prepared by bin-averaging (0.25db bins) the recalibrated downcast files and thinning the data to the depths generally used for bottles. COMPARE was run comparing those files with the bottle data in the MRG files. In the past the fit was fairly constant with pressure, but in this cast it is pressure-dependent and DO-dependent. When points were excluded with residual >0.4 the following fit was found:

DOX corrected = 1.0006 *Pressure -0.2268

This correction is -0.2268 at the surface and only -0.0018 for the deepest samples. In previous cruises simple offsets were applied in the range of -0.005 to -0.1ml/l. (See 2005-14-dox-comp3.xls.)

CALIBRATE was run using file 2005-14-recal2.ccf to recalibrate the thinned files using the equation above and COMPARE was rerun on those files. The results show that the calibration was done correctly. (See 2005-14-dox-comp4.xls,)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

File 2005-14-recal2.ccf was prepared to recalibrate the Oxygen:Dissolved:SBE channel in the AVG files.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared.

Separate profile plots were prepared with PAR, Transmissivity, DOX, and Fluorescence versus pressure.

##### FINAL CTD files steps (REMOVE, HEADEDIT and CALIBRATE)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE [umol/kg], Altimeter and Flag channels were removed from all casts.

Oxgyen:Dissolved:SBE was calculated in umol/kg and that channel added to the files. REORDER was used to place the two Oxygen:Dissolved:SBE channels together.

HEADER EDIT was used to fix formats and units and to add the following comment using file 2005-14hdr.txt:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data is also unedited except as noted above.

It was recalibrated based on bottles, but the calibration is less

reliable than usual because the rosette stops were very short.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.6ml/l from   0 -  50m

•	±0.2ml/l from  50 –  200m

•	±0.1ml/l below 200m

The Standards Check routine was run and HEADER EDIT adjusted and rerun until no further problems were found. These files were named CTD.

A cross-reference listing was produced.

The sensor history was updated.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird comments from the headers.

SORT was used to put the data in order of increasing pressure.

REMOVE was used to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag.

Oxgyen:Dissolved:SBE was calculated in umol/kg and that channel added to the files.

REORDER was used to arrange the two Oxgyen:Dissolved:SBE channels together.

HEADER EDIT was run to fix channel names and formats and to add a comment (2005-14-bot-hdr.txt) including an explanation of the quality flags and a comment that applies to all the CHL sampling.

The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The output files were named CHE.

A cross-reference listing was produced.

Particulars

4. PAR values high – error discovered in con file. Fixed for cast #5.

9. Sample #43 (Niskin #14) was renamed #943 because sample #43 was also used for cast #12.

19. salinity sample said to be from Niskin #1 (sample #68) on analysis sheet and Niskin #2 (Sample #69) on Rosette sheet. Analysis sheet probably correct.

45. This file was originally saved as 44, but log note indicates it should be 45.

47. Jelly fish tentacles

51/51a. The upcast was named 51a and later renamed 0519. This was not a rosette cast, so upcast likely not needed.

69/70. File originally named 69, changed to 70.

71. Jelly fish parts over equipment. Downcast data looks ok. No bottles.

78. Rosette hit bottom. Following cast looks ok.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2005-14

Dates: Start: June 6, 2005 End: June 10, 2005

Location: SoG – JdeF Vessel: Vector

Party Chief: Masson D

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585               Cruise ID#:		2005-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 01/02/05 | Factory |  | Factory |
| Conductivity | 2173 | 19/08/04 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 2424 | 07/04/05 | “ |  |  |
| Transmissometer | 333DR | 30/03/05 | IOS |  |  |
| Dissolved Oxygen | 0766 | 16/11/04 | Factory |  |  |
| Fluorometer | 2229 |  |  |  |  |
| PAR | 1024 | 11/02/03 |  |  |  |
| Surface PAR | 16504 | 02/01/04 |  |  |  |
| Altimeter | 1024 |  |  |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |