## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2005-22

Agency: OSD

Location: WCVI/QCS

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: September 5, 2005 – September 11, 2005

Processed by: Germaine Gatien

Date of Processing: October 13, 2005 – 16 December 2005

Number of original CTD casts: 63	Number of CTD casts processed: 63

Number of bottle casts: 33		Number of bottle casts processed: 31

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#732DR), a Benthos Altimeter (#1024), a SeaBird DO sensor #0766 (on the primary pump) and a Seapoint Fluorometer (#2356 on the secondary pump) with a 10X cable. The deck unit was a model 911 (#0508) and the logging computer was an FS03. The salinometer was a Guildline model 8400B Autosal (serial number 68572).

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order.

The pressure sensor looks like it may be drifting slowly; the error is likely very small now, but occasional tests are recommended for this particular CTD since it has shown a tendency to drift. This is not unexpected in an older CTD. It was recalibrated in October 2004.

Fluorescence went off scale occasionally.

There were severe problems in some of the salinity data from the thermosalinograph. There may have been a leak in a valve that allowed fresh water in from a hose that is used to flush the loop. Many values were replaced with pad values. The salinity in sections without the severe problems still differs from CTD salinity by more than the post-cruise calibration can explain. Recalibration of salinity was based on the results of 2005-21 and comparison with CTDs; there were no loop samples. There is considerable doubt about the calibration so salinity should be considered ±0.1psu.

The dissolved oxygen data in the CTD files should be considered

- ±0.6ml/l from 0 - 75m

- ±0.4ml/l from 75 – 125m

- ±0.2ml/l from 125 - 400m

- ±0.1ml/l from 400 - 1200m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Dissolved oxygen data were available in individual files, but there were no flags or comments entered.

The nutrient, chlorophyll and bottle salinity data were obtained in spreadsheet form; flags and comments had been added.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no errors were found.

##### Conversion of Raw Data

All data were converted using file 2005-22-0001.con.

A few casts were examined and all expected channels are present.

Rosette files were converted using a start time of -2s and duration of 5s.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and BOT was used for the extensions. All BOT files were plotted and no significant outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

The same equipment was used during 2005-21 which immediately preceded this cruise, so the same settings were used for CELLTM. One cast was checked to ensure that the data were improved by this step and they were.

CELLTM was run using (0.02, 9) for the primary and (0.03, 9) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Two deep casts were plotted to check for agreement between the pairs of T and C sensors. The differences were extremely noisy for the upcasts. The upcast differences were extremely noisy. The following values are rough estimates from downcast data:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 31 | 1000 | ~0 noisy | 0.00015 | 0.0018 | Steady, low |
| 31 | 1950 | -0.0007 | 0.00008 | 0.0016 | Steady, low |
| 59 | 1000 | -0.0004 | 0.00013 | 0.0018 | Noisy, low |
| 59 | 1950 | -0.0008 | 0.00006 | 0.0017 | Noisy, low |

The pressure dependence in the conductivity and salinity does not look significant.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and a header error in cast #51 was found; scrambled Seabird time/position headers were fixed and the file reconverted. The header check was rerun and no further errors found.

The track plot was produced and looks reasonable.

The header summary was run and station names were checked against the log as well as a random sampling of positions and times to ensure that the GPS was working well. There was one discrepancy in station name between the log and the headers, and the station name was changed in the headers of both the CTD and bottle files as it was clearly wrong.

The average surface pressure is 1.9db, which is a little lower than usual for the Tully; during 2005-12 it was 1.8db and for 2005-21 it was also 1.9db, so either the method of deployment may have changed slightly or the pressure may be drifting.  A few casts were checked for which the surface pressure was close to or less than zero; the salinity values do not look as low as expected. For cast #23 the CTD was started on deck, though the pumps were not turned on until it was in the water. Judging by the conductivity and transmissivity, it looks like the CTD entered the water when the pressure was approximately -0.12db. The pressure sensor was calibrated in October 2004 and an offset is included in the configuration file based on that calibration. Further drift is expected in an older pressure sensor like this one. No recalibration will be applied at this time since the error is probably very small and the evidence of drift is weak. Regular testing of this pressure sensor is recommended. Recalibration can be applied later if testing shows it necessary.

The altimeter readings from the headers were exported to a spreadsheet and a few casts were checked to see that the values were reasonable. No bad values were found, but as expected the header reading is higher than the minimum by about 1db, because it is the median of values in the bottom 2db.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. The entries for casts #75 and 95 were removed from the file because they had no sample numbers assigned. Samples were taken for Akash Sastri from UVic; those bottle files will not be processed further. Sample numbers were added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

The bottle salinity data were provided in spreadsheet form with flag and comment channels added. The file was saved as 2005-22sal.csv. That file was then converted to individual SAL files.

Since there were no flags in the individual dissolved oxygen files, the analyst provided a spreadsheet with flags and comments added (2005-22 hyd.xls). This was simplified and reordered on sample number and saved as 2005-22 hyd.csv. The spreadsheet was converted to individual ADD files. The file named as #17 was renamed as #18.

The nutrient spreadsheet was simplified and saved as 2005-22-nuts.csv. That file was then converted to NUTS files. These were put through SORT to rearrange by sample number. The file named as cast #17 was renamed as #18. (Output: NUT)

Extracted chlorophyll data were obtained in spreadsheet form. The spreadsheet was simplified and saved as 2005-22chl.csv. There were problems with the acetone during the analysis and all values were flagged “d”. A note was added to the 2005-22-bot-hdr.txt mentioning the problem and referring to another document that contains a full explanation. The spreadsheet was then converted to individual CHL files. The file named as #17 was renamed as #18.

The SAL, CHL, ADD and NUT files were merged with SAMAVG in four steps. (Output: MRG1, MRG2, MRG3, MRG)

##### 12. COMPARE

Salinity

COMPARE was run. There were three extreme outliers – cast #2 at 47db, cast #32 at 75db and cast #92 at 59db.

When data above 150db were excluded the primary salinity was found to be high by an average of 0.0002psu and the secondary was high by 0.0021psu; there was a lot of scatter but the trendlines were quite flat. There is no noticeable time dependence. When only the data from 400db downwards are included, the primary is high by about 0.0005psu, but the range is from +0.004 to -0.002psu. Removing a few more outliers makes little difference in the average.

The following were significant outliers even for shallow data:

- Cast #2 at 47db – There was a huge shed wake when the CTD stopped. The salinity had settled at the level of the CTD, but the bottle might contain some of that water, explaining the low salinity. No flag will be added, as the value may well reflect the contents of the bottle. A note was added to the header about the sample.

- Cast #32 at 75db – This salinity sample is way out of line; the value of 31.7865 is unbelievable at this depth. The sample was flagged “d”. It looks like it might have come from one of the near-surface bottles.

- Cast #92 at 59db – The stop at the bottom was marked by a lot of motion in the CTD, and the salinity never achieved equilibrium. However, the bottle salinity value is much lower than can be explained by shed wakes, with values like those seen at 23db during the downcast. The nutrient values also look out of line, resembling those from the bottle at 30db. Perhaps this bottle either closed early, or did not close fully allowing some leakage of shallower water into it. The sample was flagged “d” for all bottle samples after Frank Whitney and Janet-Barwell Clarke confirmed that the Dissolved Oxygen and Nutrient values were also outliers.

Dissolved Oxygen

COMPARE was run using sample number as the reference variable. The best fit was with the differences between bottles and CTD versus CTD DOX; this fit was unusually tight. When data from below 1200db were excluded from the fit, and all points flagged by the analyst and a few other outliers, the fit found was:

CTD-BOT = 1.4191 DOX-CTD + 0.0128

When this sensor was used during 2005-21 the fit found for all casts except the first three was:

CTD-BOT = 1.4255 DOX-CTD + 0.1063

The lowest value of DO sampled during the cruise was ~0.15ml/l so there are no worries about the sensor being affected by hypoxic conditions. The differences against time show too much scatter to identify any time dependence. (See 2005-22-dox-comp1.xls.)

No further outliers were identified beyond those noted by the analyst.

Fluorescence versus Chlorophyll

COMPARE was run to look at the chlorophyll versus fluorescence relationship. The chlorophyll values were much lower than fluorescence. There are known problems in the chlorophyll analysis, and all data were flagged “d” by the analyst.

##### 14. SHIFT

Fluorescence

A shift of +24 records (1s) was applied as has been done for all recent cruises. One cast was checked before and after this step and the results were satisfactory, with the resulting offset between downcast and upcast data looking much closer to that seen in the temperature signal. (Output: SHFFL)

Dissolved Oxygen

As the same equipment was used during 2005-21 the same values were used, then a few casts were examined to see if the results were satisfactory which they were. All casts were shifted by +140 records, this required 2 runs since there were 2 DO channels. (Output: *.SHFOX1 and SHFOX2)

Conductivity

Again SHIFT was run using the results for 2005-21 and then a few casts were checked to ensure the data were improved, and they were.

All casts were put through SHIFT using -0.4 records for the primary conductivity and +0.5 records for the secondary. (Output *.SHFC1 and SHFC2).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None.

The descent rate was low for many of these casts, but probably not low enough that too much data will be removed.

##### DETAILED EDITING

The primary sensors were chosen for the archive. Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing. Casts requiring heavy editing were: 50, 57-69, 79, 87-91.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensors have been used during 2005-11, 2005-12 and 2005-21 when they were found to produce salinity that was low by 0.0001, low by 0.0027 and high by 0.0006psu respectively. There was a lot of scatter in the first two cruises. The primary sensors were chosen for the archive for each of those cruises. The secondary sensors were used during 2005-21 when the salinity was found to be high by 0.0027psu.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The temperatures were higher than the historic maxima in the top 50m for casts 24, 25 and 57 which were all just offshore of the 200m isobath. Deep temperatures were just above or very close to the range maxima at casts 28, 31, 79 and 83, which were the casts furthest offshore. Salinity was a little high around 50m for casts 70, 73, 74 and 84. These excursions from the climatology are thought to be real, not indications of the calibration being off. High temperatures have been reported in 2005 in the top 50m based on ARGO float data. Higher temperatures near the surface and at mid-depths were noted for some casts during 2005-21 and around 50m during 2005-17.

##### Initial Recalibration

No recalibration will be applied to the salinity. The differences show a lot of scatter, but the average suggests that the CTD is well within 0.001psu of the bottles.

An initial recalibration was run on the MRG and SAM files using file 2005-22-recal1.ccf to apply the following correction to the dissolved oxygen:

CTD-BOT = 1.4191 DOX-CTD - 0.0128

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the calibration was applied to the EDT files as well. (See 2005-22-dox-comp2.xls.)

##### Special Fluorometer Processing

The COR files were clipped to 100db. (Output: CLIP) Those files were averaged with 0.25m bins.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

Files were bin-averaged to 0.25m bins for just the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences were slightly depth-dependent with higher values near the surface and at higher DO values. On average, the downcast DO data appear to be high by about 0.13ml/l. A second recalibration was applied to the downcast files only (AVG, THN1 and CLIPAVG) using equation:

DOX COR  = 0.9946 DOX CTD - 0.1136

(Output: COR2, THN2 and CLIPCOR)

A final run of COMPARE was done to ensure the recalibration was done properly and it looks fine. (See 2005-22-dox-comp3.xls and 2005-22-comp4.xls.)

The clipped files were put through BIN AVERAGE (0.25db bins), REMOVE and HEADEDIT and named as *.FCTD and set aside for the use of Angelica Pena.

The SAMCOR files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE (volts/kg), Oxygen:Dissolved:SBE (umol/kg), Status:Pump, Descent_Rate and Flag. The DO channel in umol/kg had not been recalibrated, so it was removed and then recalculated by running the CHANGE UNITS routine. REORDER was run to get the two DO channels together. CLEAN was run to ensure the DO pad values are all -99. (Output:CLN2)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.6ml/l from 0 - 75m

- ±0.4ml/l from 75 – 125m

- ±0.2ml/l from 125 - 400m

- ±0.1ml/l from 400 - 1200m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data. Profile plots were made displaying Temperature, DO, fluorescence and transmissivity profiles.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE (volts), Oxygen:Dissolved:SBE (umol/kg), Status:Pump, Descent_Rate and Flag. (Output: *.MRGREM)

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units, to change the agency name and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### 24. Thermosalinograph Data

a.) Checking calibrations

There were 3 files containing TSG data. All con files were identical and they were the same as the con file used 2005-21. A report was printed for the first con file.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

Time-series plots were produced and only one large spike noted in salinity in file #3. There are some odd features in salinity that look as though the flow to the conductivity cell might have been irregular. Problems were noted during 2005-21 with the same equipment, though in that case the flow meter showed that flow did vary. In this case the flow looks ok.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within 1db of 4db and exported to a spreadsheet. The TSG files were opened in EXCEL and reduced to the times when CTDs were run. Those files were combined in a spreadsheet matching times. The positions were compared and were very close, with average differences for both latitude and longitude of 0.0001º and no difference was greater than 0.0004º, so the clock would appear to be accurate. This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence. (See 2005-22-TSG-CTD-comp.xls) (Cast #57 was later added to the spreadsheet although there was no data from 4m available, because this was an extremely well-mixed cast at a point in the cruise where there were grave doubts about the quality of the TSG data.)

d.) Alignment check

Recent uses of this equipment showed no alignment problems and there is no indication of such trouble, so this step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 When the differences between intake temperature and lab temperature were plotted there was a lot of noise. The average differences were -0.199, -0.163 and -0.182Cº for the three TSG files. The quietest section was at the beginning of the second file and the average was then -0.136Cº. Looking at file 2005-22-TSG-CTD-comp.xls, we find that during stops for CTDs, the temperature from the lab is higher than the remote temperature by an average of 0.192Cº, and the median value is 0.178Cº. When only CTD casts that were well-mixed to 8db or deeper were included the average difference was 0.143Cº. During 2005-21 which immediately preceded this cruise the heating due to the ship was found to be 0.14Cº. Given that the differences reflect both real time variability and heating due to the ship, the value from quiet section is most reliable and is close to the results of 2005-21.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. (See 2005-22CTD-tsg-comp.xls)

The results were surprising. Most notable are the salinity differences which are very unusual. The differences are noisy but low up until about cast #22 when they become somewhat larger. Then after cast #49 they become much larger with the difference at cast #60 being 11.3psu. The differences then returned to much lower values, with a few outliers of up to 1psu. The TSG salinity is low by an average of 0.01psu between casts 67 and 87.

The TSG fluorescence is higher than that of the CTD by about 60%. However, we know that the calibration used in conversion was wrong. A correction (offset -0.05, slope 0.913) was applied to the data and when 3 outliers were excluded (CTD FL>11) the TSG is high by about a third.

The temperature differences are also noisy, though not as extreme as the salinity. At cast #48 the TSG temperature is high by 1.2Cº, but this is close to shore and may be caused by high variability. Between casts 67 and 87 the TSG intake temperature is higher by 0.08Cº.

To investigate further a graph was prepared with 3 differences plotted together versus cast number: the differences between the TSG sal and CTD sal, TSG temp and CTD temp, and TSG intake temp and TSG lab temp. The following observations were made:

Casts #4 and 13: All differences are large so this looks like large local variations, no sign of poor TSG performance.

Casts 23-48: The temperature differences are noisy for the first few casts in this section but then settle down to average values. The salinity looks odd throughout, so it looks like a problem in the TSG salinity.

Casts 49-63: The salinity now looks very bad. The temperature differences are a little noisy for the first few casts in this section, but not enough to account for such huge salinity differences. Through most of this section the temperature differences look average. It looks like there is a major problem in TSG salinity.

Casts #67-87: Salinity differences look normal. Temperature varies somewhat, but mostly close to average.

Cast #88: Both the intake temperature and the lab salinity are very different from the CTD, though the difference between lab and intake are not large. This probably indicates a poor match between what the TSG and CTD were seeing, perhaps due to poor CTD data, but does not suggest poor performance of the TSG.

When an average is based on only the casts for which the salinity differences between CTD and TSG do not look out of line with the temperature differences, the salinity is found to be low by 0.12psu.

- An interesting cast is #57. This is a very well-mixed cast to more than 15m so an exact match of depths is not critical to the comparison. The CTD temperature and TSG intake temperature are in very good agreement, within 0.003Cº. The TSG primary temperature is higher than the intake temperature by about 0.12Cº. The TSG salinity is 5.9psu lower than the CTD. So whatever is going on is affecting the salinity, but not the temperature. The TSG fluorescence is about 2.2 times that of the CTD. The ratio of CTD to TSG fluorescence looks a little out of line between casts #22 and 47, but not between casts #48 and 63 when the salinity is at its worst. It is likely that the fluorescence is fine; the variability in the ratio is not unusual.

- Calibration History

The TSG was recalibrated in December 2004. For 2005-01 and -02 there were some concerns about the performance of the Portasal when the loop samples were processed. The TSG was lower than the loop samples by about 0.02psu for those two cruises, but the results are believed to be low by at least 0.01psu due to salinometer errors, which would imply the TSG was low by at least 0.03psu. For 2005-08 there were no loop samples; the salinity was found to be low by 0.035psu. For 2005-12 there were only 6 loop samples with a lot of scatter; the salinity was found to be low by about 0.05psu. For 2005-21 there were 13 loop bottles with a lot of noise in the results. When the sections of TSG data with low variability were used, the TSG was found to be low by about 0.07psu. The comparison with CTDs showed the TSG to be low by about 0.08psu. During 2005-21 the intake temperature was found to be very close to the CTD temperature. The heating of loop water by the ship was estimated to be +0.14Cº.

- Post-cruise calibration

The TSG was recalibrated in October 2005 and the salinity was found to have an error of about 0.001psu due to conductivity drift and a further 0.001psu due to temperature drift. This is a much smaller drift than found through comparisons with CTD and loop samples from cruises in 2005. It is also a much lower drift than has been found in this instrument in the past. This suggests that either the calibration changed during or after this cruise, or that there are problems with this equipment other than calibration.

Conclusions

Doug Yelland recalls that the hose used to flush the loop was in place, though ship personnel believed the valve was turned off. When Doug noticed the salinity being way out of line, he asked that the hose be completely removed. There is no record of precisely when this happened, but most of the data in the 3rd file looks ok, so it is likely that this was the source of major trouble. Either the valve was not completely shut off, or there was a malfunction in the valve. While this accounts for the largest errors in salinity, it does not account for the 0.1psu difference between TSG and CTD near the end of the cruise. So there must be some other source of small error.

The TSG intake temperature is not in as good agreement with the CTD as during 2005-21, but this is probably due to the variability in surface conditions. At one station with well-mixed surface water the match was excellent.

The primary temperature should be lowered by 0.14Cº to account for warming in transit to the lab. This is based on the results of 2005-21 when surface temperatures were similar, and the comparison with the CTD for casts that were well-mixed to at least 8m.

In the absence of loop samples, the salinity adjustment will be based on the results of 2005-21 and the comparison of the TSG with the CTD for casts with salinity that looks reliable. Unfortunately, none of these were very well-mixed. The salinity should be adjusted by adding 0.10psu, but note should be made in the header than this should be considered ±0.1psu since there is evidence that the error may be anywhere from 0.002 to 0.020psu.

The TSG Fluorescence needs to be corrected.

f.) Editing

The time-series plots were examined and the salinity looked suspicious in many places. Based on the conclusions in section (e) suspicious salinity data were replaced with pad values. Near the end of the first file there is a sudden drop by about 0.2psu and this is close to cast #23 when the change in differences between TSG and CTD first manifest themselves. The salinity data from that point was replaced with pad values, plus all salinity from the second file and some from the beginning of the third file. One spike in salinity was cleaned in the third file. Notes were placed in the headers to indicate what editing was done.

The fluorescence data around 17:30 – 18:00 on Sept. 7 was examined because the values were mostly very high but had some low values embedded in them and the time of day does not explain this; moreover, the log states that the fluorometer was removed for repair at about that time. However the CTD casts from around the same time are in agreement with the TSG. The values were high at CTD casts #34 and 41, but low at casts #34 and 39, so the data is probably fine; no editing was done to fluorescence.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.14 Cº and +0.10psu to Temperature:Primary and Salinity:T0:C0. The fluorescence was recalibrated (offset -0.05, slope 0.913) because there was an error in the co-efficients used in the conversion.

After calibration the two temperature channels were compared and are closer.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Primary, Flag and UPloy0 (flow rate).

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, to change the location to WCVI/QCS and to add the depth of sampling to the header and to add a general comment about the sampling method and the following particular comments:

*******

WARNING

*******

Much of the salinity data from this cruise was bad, with values differing from CTD salinity by up to 10psu. Pad values were entered for large sections of salinity data, but all data must be considered suspect.   The post-cruise calibration did not match the history of the instrument or the comparisons with CTD data. This may indicate that there was a sudden change in calibration.  More likely there are errors caused by something other than calibration problems. During 2005-21 concerns were expressed that fresh water was getting mixed into the loop intermittently. During this cruise it was noticed that the hose for flushing the system was in place, though it was believed that the valve was closed. It is likely that the valve was leaking. The hose was removed late in the cruise and the data from late in the cruise is better. The temperature of the freshwater introduced appears to have been equilibrated to ambient ocean conditions.  The salinity difference between TSG and CTD is on the order of 0.1psu even after the hose was removed, so there must be a second source of small error that does not appear to be due to the calibration of the instrument.

No loop salinity samples were available for 2005-22.

The salinity should be considered +/- 0.1psu.

Temperature, salinity and fluorescence data was saved with 2 decimal places only.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

17/18. The file named as #17 on the rosette sheet was later changed to #18. The CTD log book confirms that the bottles were fired during #18.

41. Note in log: SW Loop off for repairs (fluorometer only)

51. Clock on SeaSave wrong – computer and headers ok.

69. Brief stop at 50m. Log note:  “Winds gusting to 45kt.”

70. Log note: “blowing a gale”

# Institute of Ocean Sciences    CRUISE SUMMARY

| Cruise ID#: 2005-22 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 5 September 2005 End: 13 September 2005 |  |  |  |  |  |  |
| Location: West Coast Vancouver Island / Queen Charlotte Sound |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Yelland D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-22

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19Mar05 | Factory “ |  |  |
| Conductivity | 3038 | 03Mar05 | “ |  |  |
| Secondary Temp. | 2106 | 08Jul05 | “ |  |  |
| Secondary Cond. | 1729 | 12Jul05 | “ |  |  |
| Transmissometer | 732DR | 22May05 | Factory |  |  |
| SBE 43 DO sensor | 766 | 16May05 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Altimeter | 1024 |  |  |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2005-21

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 4/12/04 | Factory | 21/10/05 | Factory |
| Conductivity | 2487 | 4/12/04 | “ | 21/10/05 | Factory |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab/wetstarFluor. | WS3S-713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? | ? |  |  |