## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2005-23

Agency: Ocean Sciences Division

Location: JdeF/WCVI/QCS

Project: La Perouse / Ecohab / PERD / Aquaculture

Party Chief: Juhasz T.

Platform: John P. Tully

Date: September 30, 2005 – October 10, 2005

Processed by: Germaine Gatien

Date of Processing: March 22, 2005 – 31 March 2005

Number of original CTD casts: 45    	Number of CTD casts processed: 45

Number of bottle casts: 26	       	Number of bottle casts processed: 25 (no sampling from #11)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#498DR), a Benthos Altimeter (#1024), a SeaBird DO sensor #0766 (on the primary pump) and a Seapoint Fluorometer (#2356 on the secondary pump) with a 10X cable. The deck unit was a model 911 (#0508) and the logging computer was an FS03. The salinometer was a Guildline model 8400B Autosal (serial number 68572).

# SUMMARY OF QUALITY AND CONCERNS

The personnel list was not entered on the first page of the CTD Daily log, though the information was entered in a comment within the log. The rosette log has one sample number listed twice. This arose because a bottle failed to fire; no note was made of that on the rosette sheet, but the CTD Daily log did have a comment about it.

The pressure sensor continues to drift as expected in an older CTD. The offset was adjusted slightly.

The dissolved oxygen data in the CTD files should be considered

- ±0.8ml/l from 0 - 100m

- ±0.4ml/l from 100 – 200m

- ±0.1ml/l from 200 - 1000m

There were two casts in Knight Inlet that coincided with observations from 2005-30. A comparison was made that offers some evidence that for upcast data there is an effective vertical offset of roughly 4db. In future we should consider applying a pressure offset to upcast data, preferably before bottle comparisons.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. There was no list of personnel on the first page, but a list was found in comments within the log.

There was no record of thermosalinograph equipment.

Dissolved oxygen data were available in individual files, with a flag channel and comments.

The bottle salinity data were obtained in spreadsheet form. There was no flag channel. There was a comment column but no comments entered. The salinity analysis sheets were also available and there is no indication there that any flags were needed.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and the only adjustments made were to the pressure calibration. There was a tiny error in one co-ordinate and based on other cruises using this instrument the pressure offset was increased to 3.5db. File 2005-23-ctd.con was prepared with those changes.

##### Conversion of Raw Data

All data were converted using file 2005-23-ctd.con.

A few casts were examined and all expected channels are present. Upcast and downcast data are similar, but the differences between the two temperature and salinity channels are significant during the upcasts.

Rosette files were converted using a start time of -2s and duration of 5s.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and BOT was used for the extensions. All BOT files were plotted and no significant outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were:

Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on two casts with fairly steady descent and ascent rates using a variety of settings for CELLTM. The best results proved to be (0.03, 7) for the primary and either (0.03, 9) or (0.0245, 9.5) for the secondary.

CELLTM was run using (0.03, 7) for the primary and (0.03, 9) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

There was one very deep cast and a few to 500 or 600m. Four casts were plotted to check for agreement between the pairs of T and C sensors. The differences were extremely noisy for the upcasts. The following values are rough estimates from downcast data:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 29 | 580 | -0.0006 | 0.00025 | 0.0025 | OK |
| 32 | 480 | -0.0005 | 0.0002 | 0.003 | OK |
| 38 | 580 | -0.0006 | .00025 | .00033 | High quiet |
| 38 | 1200 | -0.0005 | 0.0002 | 0.003 | High quiet |
| 39 | 580 | -0.0005 | 0.0002 | 0.003 | Very noisy |

The pressure dependence in the conductivity and salinity does not look significant.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and no errors were found.

The track plots (using event #s and station names) were produced and look reasonable; they were added to the end of this report.

The header summary was run and station names were checked against the log as well as a random sampling of positions and times to ensure that the GPS was working well. There were a few discrepancies in station name between the log and the headers, and the station name was changed in the headers of both the CTD and bottle files.

The average surface pressure is 1.9db. While this is a little low for the Tully, none were less than 1.3db, so this looks reasonable. A close eye should be kept on this CTD since the pressure is drifting. During 2005-32 which was about 6 weeks after this cruise Doug Anderson determined that an offset of 3.545db should be used. Over the past year values of 3.0457 was used for most cruises, and 3.4457db was used for 2005-16; 3.5 was applied to this cruise.

The altimeter readings from the headers were exported to a spreadsheet and a few casts were checked to see that the values were reasonable. No bad values were found, but as expected the header reading is higher than the minimum by an average of about 1db, because it is the median of values in the bottom 2db.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. The same sample number was entered for casts #11 and 12 on the rosette sheet. In the CTD Daily log there is a note that the bottle did not fire during cast #11, so that rosette sheet should be ignored. Sample numbers were added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

- The bottle salinity data were provided in spreadsheet form with flag and comment channels added. The file was saved as 2005-23-sal.csv. That file was then converted to individual SAL files.

- The dissolved oxygen data were provided in ADD files with flag channel included. The file names were non-standard and had to be edited.

- There was no CHL or nutrient data gathered on this cruise.

The SAL and ADD files were merged with SAMAVG in two steps. (Output: MRG1, MRG)

##### 12. COMPARE

Salinity

COMPARE was run. There was a lot of scatter in the results with the same outliers for both salinity channels, so this is assumed to be a problem in collection or analysis. Plots of transmissivity versus pressure show that none of the problems are from mud in samples. Excluding 5 outliers produced a reasonably flat trend line for the primary differences versus pressure, with primary salinity low by about 0.0016. Gradually removing more points to achieve an even flatter trend line suggests an average difference of about -0.0012. For the secondary pair the salinity is high by about 0.0012 with slight variations depending on what is included in the fit.

The following significant outliers were flagged “c” and a note put in the headers:

- Cast #3 at 150db

- Cast #12 at 149db

- Cast #15 at 114db

Dissolved Oxygen

COMPARE was run using sample number as the reference variable. The best fit was with the differences between bottles and CTD versus CTD DOX; this fit was unusually tight. When a few outliers were excluded from the fit, the fit found was:

CTD-BOT = 1.4038 DOX-CTD + 0.0692

When this sensor was used during 2005-22 the fit found for all casts was:

CTD-BOT = 1.4191 DOX-CTD + 0.0128

When this sensor was used during 2005-21 the fit found for all casts except the first three was:

CTD-BOT = 1.4255 DOX-CTD + 0.1063

The lowest value of DO sampled during the cruise was ~0.15ml/l so there are no worries about the sensor being affected by hypoxic conditions. The differences against time show little variation. (See 2005-23-dox-comp1.xls.)

None of the outliers were severe enough to warrant flagging of the data.

##### 14. SHIFT

Dissolved Oxygen

Tests were run on 2 casts using shifts of +120, 140 and 160 since +140 has been selected for this sensor in recent use. The best choice was judged by matching features in upcast and downcast and picking the shift that made the vertical offset for DO the same as for temperature. All casts were shifted by +140 records. (Output: *.SHFO)

Fluorescence

A shift of +24 records (1s) was applied as has been done for all recent cruises. A few casts were checked before and after this step and the results were satisfactory, with the resulting offset between downcast and upcast data looking much closer to that seen in the temperature signal. (Output: SHFFL)

Conductivity

During 2005-21 settings of -0.4 and +0.5 were used for the primary and secondary conductivity. Tests were run using settings around those values to ensure the same choice is appropriate AND while the primary looked best with -0.4, +0.7 looked better for the secondary.

All casts were put through SHIFT using -0.4 records for the primary conductivity and +0.7 records for the secondary. (Output *.SHFC0 and SHFC1).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00 	Minimum Surface Salinity: 5

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None.

##### DETAILED EDITING

The primary sensors were chosen for the archive. Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensors were used during 2005-21, 2005-22 and 2005-25 when they were found to produce salinity that was high by 0.0006, 0.0005 and 0.0006, respectively. During 2005-11 and 2005-12 the salinity was found to be low by 0.0001 and 0.0027, but there was a lot of scatter in both those comparisons. The secondary sensors were used during 2005-21, 2005-22 and 2005-25 when the salinity was found to be high by 0.0027, 0.0021 and 0.0032.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The only excursions from the local historic climatology were from the northern half of Juan de Fuca Strait, where the deepest salinity data were occasionally above the maxima. This is not considered indicative of problems with calibration since excursions in that region are common, probably due to limitations in the climatology.

Other data for comparison

Cruise 2005-30 visited the same area within a few hours of 2005-23. Casts 2005-30-0074 and 2005-30-0076 are at the same locations as casts 2005-23-0046 and 2005-23-0047 from this cruise. An SBE25 was used during 2005-30. An initial comparison shows salinity to be reasonably consistent; variations in temperature are more significant. Further comparison will be done after recalibration and bin averaging. See section 24 for further discussion.

##### Initial Recalibration

The comparison with bottles indicates that the primary salinity is low by about 0.0012 which is similar to the results of other cruises that used this equipment. No recalibration will be applied since the difference is small and the scatter large.

An initial recalibration was run on the MRG and SAM files using file 2005-23-recal1.ccf to apply the following correction to the dissolved oxygen:

CTD-DOX (Corrected) = 1.4038 DOX-CTD + 0.0692

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the calibration was applied to the EDT files as well. (See 2005-23-dox-comp2.xls.)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

SHIFT addresses errors due to transit time and the comparison with titrated samples addresses the drift in the instrument’s calibration. But there remains an error due to poor time response of the sensor. To at least partly address this issue we compare downcast data to the titrated samples. COMPARE was run using thinned downcast CTD data (after initial recalibration and bin-averaging) and the titrated DO values. When outliers were excluded, the differences were fairly flat when plotted against pressure or bottle DO, with the CTD reading high by 0.09ml/l. For 2005-21 and -22 with the same instrument, this offset was 0.12 and 0.13ml/l. When plotted against DOX-BOT the fit is even flatter and not strongly dependent on what points were excluded. (See 2005-23-dox-comp3.xls)

File 2005-23-recal2.ccf was applied to subtract 0.09ml/l from DO channel in the thinned files and COMPARE was rerun to check that the calibration was run correctly. (See 2005-23-comp4.xls.) The AVG files were then recalibrated. This correction was not applied to the bottle files since it addresses an error that does not affect the CTD data while it is stopped. (Output: COR2)

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE (volts/kg), Status:Pump, Descent_Rate and Flag. CHANGE UNITS was used to add a second DO channel in umol/kg and REORDER was run to get the two DO channels together. (Output: REO)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.8ml/l from 0 - 100m

- ±0.4ml/l from 100 – 200m

- ±0.1ml/l from 200 - 1000m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data. Profile plots were made displaying Temperature, DO, fluorescence and transmissivity profiles.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE (volts), Status:Pump, Descent_Rate and Flag. (Output: *.MRGREM)

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. At this point the file for cast #11 was removed since there was no bottle sampling. (Output: CHE)

##### 24. Comparison of Knight Inlet casts from 2005-30 and 2005-23

The final data from the two casts in Knight Inlet were compared. The data from 2005-30 were collected using an SBE25 and upcast data were archived. The SBE25 has been plagued by problems during downcasts, but Dario Stucchi made a discovery during 2005-30 that there was an inappropriate setting for the pump turn-on, so that the pump came on late in casts with low surface salinity. Another cause of pump trouble in downcasts has been found to be clogged valves which clear themselves around 50db.

When the pairs of casts are examined on T-S surfaces the salinity varies little along constant density surfaces, but there is more variation in temperature. When profiles are examined the reasons becomes clear. The salinity is monotonically increasing and has a low gradient below the surface, but the temperature has a lot of structure even at 100m for these casts. For the upcast data we know that the CTD will drag deep water up with it. There is a clear vertical offset in all variables, on the order of 4 or 5m. The deep values are quite close in salinity and dissolved oxygen, a result of low gradients and, to some extent, because bottle calibrations have forced the match. The resulting errors in temperature are on the order of 0.01Cº at depth and 0.1Cº near the surface. The salinity errors are only significant very close to the surface. Complicating the comparison is the fact the upcast data are also marked by some mixing by the CTD package and the SBE25 has a lower sampling rate, so there is not the level of detail that is seen by the 911+.

The question arises as to how to use this result to improve our calibration of upcast data. Is it possible that we may be creating a mismatch of T and S by calibrating salinity? Probably not, since the calibration of salinity generally uses only low-gradient parts of the water column. The DO calibration is a little more problematic, since we have samples from all depths, but unless T vs DO is used by researchers this is not a problem. Should we apply a pressure offset, if we can determine what it is? If so, it makes sense to apply a pressure offset first, then do the bottle comparisons, especially to get better DO calibration. If the occasion arises again, some time should be spent determining if there is a way to calculate the effective offset.

##### 25. Thermosalinograph Data

a.) Checking calibrations

There was 1 file containing TSG data. The con file had an error in the fluorescence calibration; that was corrected and the file saved as 20023-TSG.con.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

A time-series plot was produced and spikes were noted in the primary temperature and salinity in the first 28 hours of the record.

CTDEDIT was used to clean primary temperature and salinity at about 16 points between records #233 and 3149. When the TSG track was plotted it became obvious there was an error in one position. CTDEDIT was used again to clean latitude and longitude for record #4023. The track then looked fine and was added to the end of this report. A text editor was used to remove the first 10 records because the flow rate was very low.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within 1db of 4db and exported to a spreadsheet. The TSG files were opened in EXCEL and reduced to the times when CTDs were run. Those files were combined in a spreadsheet matching times. The positions were compared and were very close, with average differences for both latitude and longitude of 0.0001º and no difference was greater than 0.0006º, so the clock would appear to be accurate. This spreadsheet will also be used in the next step to compare temperature, salinity and fluorescence. (See 2005-23-TSG-CTD-comp.xls)

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 When the differences between intake temperature and lab temperature were plotted there was a lot of noise. The average difference was -0.214 Cº, with a median value of -0.209 Cº.  In a quiet section of 400 scans the average was -0.177Cº. Looking at 2005-23-TSG-CTD-comp.xls, we find that during stops for CTDs, the temperature from the lab is higher than the remote temperature by an average of 0.225Cº, and the median value is 0.219Cº.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. A subset of those casts was chosen as well-mixed based on a mixed-layer depth calculation. (See 2005-23-TSG-CTD-comp.xls and 2005-23-TSG-CTD-comp-well-mixed.xls.)

The TSG fluorescence is higher than that of the CTD by 25 to 30%.

The TSG salinity is low by between 0.025 and 0.075, with a difference of 0.037 in the patch of data with the lowest variability.

The TSG intake temperature is lower than the CTD temperature by <0.007Cº. This is as close as one could expect.

The lab temperature is higher than the CTD temperature by an average of 0.21 to 0.22Cº  and the average difference between lab and intake is 0.22Cº.

e.) Loop Salinity comparison

There were 23 salinity samples taken from the loop. The salinity data in the TSG file were averaged and the standard deviation found over 2 minutes (5 samples); the average value and standard deviation were recorded for the times of the loop samples. These were compared and the TSG salinity is low by 0.036 when all samples are included. When outliers are removed it is low by 0.046 and when only the 10 samples with the lowest standard deviations are included it is low by 0.045.(See file 2005-23-loop-comp.xls)

f.) Calibration history

- Previous experience with this equipment

The TSG was recalibrated in December 2004. For 2005-01 and -02 there were some concerns about the performance of the Portasal when the loop samples were processed. The TSG was lower than the loop samples by about 0.02 for those two cruises, but the results are believed to be low by at least 0.01 due to salinometer errors, which would imply the TSG was low by at least 0.03. For 2005-08 there were no loop samples; the salinity was found to be low by 0.035. For 2005-12 there were only 6 loop samples with a lot of scatter; the salinity was found to be low by about 0.05. For 2005-21 there were 13 loop bottles with a lot of noise in the results. When the sections of TSG data with low variability were used, the TSG was found to be low by about 0.07. The comparison with CTDs showed the TSG to be low by about 0.08. During 2005-21 the intake temperature was found to be very close to the CTD temperature. The heating of loop water by the ship was estimated to be +0.14Cº. During 2005-22 it was discovered that a hose was connected to the loop system in such a way that fresh water might have been introduced to the loop. The salinity was found to be low by 0.10.

- Post-cruise calibration

The TSG was recalibrated in October 2005 and the salinity was found to have an error of about 0.001 due to conductivity drift and a further 0.001 due to temperature drift. This is a much smaller drift than found in comparisons between CTD and loop samples from cruises in 2005. It is also a much lower drift than has been found in this instrument in the past. This suggests that either the calibration changed during or after this cruise, or that there are errors in this equipment other than those due to calibration drift.

Conclusions

The intake temperature is in good agreement with the CTD.

The primary temperature should be lowered by 0.22Cº.

The salinity should be raised by 0.045. This is a smaller error than has been seen in other cruises between July and September of 2005, but problems were found in the loop system during 2005-22 that may have led to fresh water getting into the system.

The fluorescence is reasonably close to the CTD.

f.) Recalibration

File 2005-23-TSG-recal.ccf was used to apply offsets of -0.22 Cº and +0.045 to Temperature:Primary and Salinity:T0:C0. After calibration the two temperature channels were compared and are much closer.

h.) Preparing Final Files

CLEAN was rerun to fix the header start and stop positions. A text editor was used to fix the geographic area in the header.

REMOVE was used to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Primary, Flag and UPloy0 (flow rate).

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and to add the depth of sampling to the header and to add a general comment about the sampling method.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated for the TSG and CTD sensors.

Particulars:

15. Stop at 90m to adjust wire angle

21. Odd spikes in salinity at 97m and 70m upcast

22. Stop at 186

30. Salinity spike at 248m on upcast

31. Bottle 17 empty.

35. Bottle 7 empty

39. Cast cancelled due to weather

44. Seasave crashed while getting initial data

# Institute of Ocean Sciences    CRUISE SUMMARY

| Cruise ID#: 2005-23 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 30 September 2005 End: 10 October 2005 |  |  |  |  |  |  |
| Location: West Coast Vancouver Island / Knight Inlet |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Juhasz T. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-23

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19Mar05 | Factory “ |  |  |
| Conductivity | 3038 | 03Mar05 | “ |  |  |
| Secondary Temp. | 2106 | 08Jul05 | “ |  |  |
| Secondary Cond. | 1729 | 12Jul05 | “ |  |  |
| Transmissometer | 498DR | 22May05 | Factory |  |  |
| SBE 43 DO sensor | 766 | 16May05 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Altimeter | 1024 |  |  |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2005-23

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 4/12/04 | Factory | 21/10/2005 |  |
| Conductivity | 2487 | 4/12/04 | “ | 21/10/2005 |  |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab WetstarFluorometer | WS3S-713P | 10/01/01 | ? |  |  |
| Flow Meter | ? | ? | ? |  |  |