## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2005-24

Agency: PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.

Project: Sea Lice Sampling

Area: Broughton Archipelago

Chief Scientist: Hargreaves B.

Platform: Clupea

Date: 24 March 2005– 29 April 2005

Processed by: Germaine Gatien

Date of Processing: 13 February 2006 –17 February 2006

Number of original CTD casts: 16 hex files

Number of casts processed: 13 (3 of the files were duplicates)

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#1294) was used.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard. The consecutive numbers in the log book started with “0”, so 1 should be added to each to find the event number that was used to name the files. So file 2005-24-0011 is Event #11, but in the log is Consecutive #10.

There were problems in judging which files corresponded to which log entries. The log sheets were critical to being able to use these data. There were a few positions missing, but in general, there was good detail. It was also very useful that notes had been made in the header files about timing problems. Data from 2 casts are missing – the files that appear to contain those data are actually duplicates of other casts. And there is one data file that does not appear to correspond to any log entry and one log entry with no corresponding data file.

Some doubts remain about the times and locations of these CTD casts. Casts #6, #13 and #15 are of particular concern. These data should be used with care. There were many discrepancies among available information, so this is a “best guess”.

There was no salinity calibration sampling available and no history for this instrument since its last calibration, so no estimate can be made of those errors.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The Daily Log sheets were obtained.

The original file names were non-standard and contain a description of the cast. Together with the information in the log sheets, standard names were devised. There were two files that were not clearly associated with the log entries, but a guess was made for one as being event #15. The other was called #99 and looks like a duplicate of cast #15.

File 2005-24-cast-id.xls is a spreadsheet containing the original file names, the event number assigned to that file and the consecutive number in the log book for the entry believed to correspond to the file. In most cases the file is clearly identified by details in the comments of the log sheets. Cast time helps in the identification, but in several cases there is a note in the header that the time was incorrect, and the correct time is entered as a comment. At one point a note is entered that the CTD clock was not working properly. Errors are possible in this process, so care should be taken later to examine the track plot and ensure the positions correspond to the comments in the log sheets.

There were two HEX files, named LICE504 and LICE505, which appear to be from 2003. They will not be processed.

The consecutive numbers in the log book start with #0, so 1 should be added to each of them to correspond to the event number upon which the file names were based.

There is mention of surface salinity bottles on the log sheets, but the data were not found.

Configuration file 1294.con contains the most up-to-date pre-cruise calibration coefficients for this instrument, which were from November 2002. The temperature and conductivity coefficients were checked and are correct. The pressure coefficients were not available. There is a post-cruise calibration available from November 2005, before repairs were done to the instrument. At that time it was reported that the drifts in calibration were -0.00061 Cº/year and -0.0001psu/month. If we assume drift linear with time we would expect that the temperature error is about 0.0005 Cº and the salinity is low by about 0.003psu. However, given the instrument needed repair it is possible that significant changes in calibration occurred after this cruise.

- Conversion of Raw Data

- The raw data were converted using conversion file 1294.con.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature. The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

5. DERIVE

Program DERIVE was run to calculate salinity.

6.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. There are notes in the headers of several files indicating that the dates and times are wrong in the header. All the times seem slightly different, so all times were changed to match the log sheet entries.

Cast #99 was examined at this point and it is clear that it is a copy of cast #15, so it will not be processed further.

All casts were put through HEADEDIT to add lines for latitude, longitude and water depth entries to the Location section of the headers. The project area was changed as well. Then a text editor was used to enter the positions and bottom depths and a station name based on the entries on the log sheets.

Times were made to match the log sheets. It is notable that all the times are between 08:00 and 17:00 hours, so these are most likely PST. The headers say the times are UTC, but most of those times were wrong.

Cast #13 had no positions entered, but the description for #13 says Thompson Sound, so the same position was used as for cast # 9.

Cast #14 had no position entered, but the station description is the same as for cast #8 so that location was entered.

Cast #15 had a position entered in the log book, but it does not fit the description given in the headers. The log entry was used but will be investigated.

7. Checking Headers

Initial checks turned up many errors in times and positions. After these were fixed, a Cruise Track Plot was produced. While the track looks a little odd, this mission was done in distinct groups, and when plotted that way the only odd thing is the placement of either cast #13 or 14. Those positions were inferred, not written in the log book. The Header Check indicates no problems with speeds, and no further errors were found when the cross-reference listing was compared with the log book.

The Header Summary was prepared. From this the maximum depth in each file was compared with the log entries for bottom depth and maximum depth. There are 3 cases (casts #5, 6 and 15) where the maximum depth sampled is significantly different from the bottom depth recorded in the log, but this is an area of rapid change in bottom depth, so this may just reflect a few minutes between the log recording and the CTD cast. There are notes on the log sheets about large drift during casts #5, 6 and 14. There are many more discrepancies between the maximum sampling depth recorded in the log and what actually is in the files. This occurs even for casts where there is little doubt about the identification of station site. It is not known how the maximum pressure was determined for the log. Bottom depths for the repeat casts were compared and there is good agreement for casts 7&12 and 10&11. For casts 8&14 and 9&13 there are significant differences.

T-S plots for pairs of casts are not very useful because there were huge changes in the month between measurements, even for pairs that are almost certainly identified correctly.

There remains considerable doubt about the identification of some of the files, especially #5, 6, 13, 14 and 15. The greatest doubt is for cast #14 given that the bottom depth is about half that for cast #8.

Cast #15 is described in comments in the header as being from the junction of Kingcome Inlet and Wakeman Sound. The last entry in the log is for a cast in Tribune Channel referred to as Site #5, but no file seems to correspond to that spot. The data in cast #15 looks more like the nearby sites from Kingcome and Wakeman than Tribune Channel and is too deep for site #5. A estimated position was assigned based on station Kin1.4 from 2005-30 which is near the juncture of Kingcome and Wakeman.

The surface check was run. The average surface pressure is –0.14db but the CTD was turned on in air. Examining files closely suggests that the conductivity went to non-zero values at about 0db. No adjustment will be made.

8. Test Plots

Profiles were plotted for all casts and up and down traces compared. There were no differences so great as to suggest errors in processing.

9. SHIFT

The conductivity was shifted by -1.4 records for all casts as was done for 2003-17 and 2003-23 using the same equipment.

11. DELETE

During 2003-17 and 2003-23 the upcasts were found to be more useful than the downcasts. For this cruise the descent rate was fairly quiet and the downcast data looks ok.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min     	Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	 	Pressure filtered over width: 5

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

The only warnings were from the upcast section for cast #1.

11. CTDEDIT

At this point T-S plots were made and it became obvious that cast #5 is a duplicate of #6 and #14 is a duplicate of cast #13.

- In the latter case it is obvious that #14 is not at the site given on the log sheets, so it will not be processed further. For now it is assumed that cast #13 is correctly identified, but it could be wrong too.

- It is not clear which of #5 and #6 is misidentified and they could both be wrong. The files have maximum sampling pressures that seem too low for either of the sites named in the original file names according to the log sheets. The temperature and salinity look more like cast #8 than anything else, so it will be presumed that the data are from Tribune Channel at Watson Cove. There obviously were data gathered at Fife Channel, but one file or other appears to be lost. For now it will be assumed that the file named 2005-24-0006 is the correctly named file and cast #5 will not be processed further. Later, a decision can be made on whether to keep #6 or rename it.

Downcast page plots were produced and were examined for spikes and instabilities to guide the use of CTDEDIT.  A few surface records were removed from all casts and salinity was cleaned for most. A few other records were removed for a few casts. Note was made of the editing details in the relevant files.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value = .000

Average value will be used.

Interpolated values are NOT used for empty bins

Page plots were examined and no problems noted. No further editing will be done.

- Intercomparisons

COMPARE - There were no data available for salinity calibration.

Previous Use of CTD – The only data processed using this instrument since its latest calibration were from 2003-17 and 2003-23 for which there was no salinity calibration sampling.

Historic ranges – There was local climatology available for only one cast; the data fit within the historic ranges.

# Post-cruise calibration – There is a report on drift in calibrations between November 2002 and November 2005. If we assume that the drift is linear with time, then the conductivity error would lead to salinity that is low by 0.003psu and temperature would be low by 0.0014Cº which would partly offset the salinity error. However, the drift is more likely linear with use and since the instrument needed repair in November 2005 it is possible that a large change occurred after this cruise.

14. Final Plots

THIN and DERIVED QUANTITIES routines were used to produce information for tables. Page plots were prepared using the edited data and tables.

15. REMOVE, FIX TIME and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The station names were changed from geographic descriptions to site # where the latter was available.

The position for cast #15 was changed to one that fits the description in the headers.

The ADD TIME CHANNEL routine was used to add 8 hours to the times in the headers.

The HEADEDIT routine was used to fix formats, the chief scientist’s name and channel names and to add the following comment to the headers:

There are some doubts about the times and locations of these CTD casts. See

the processing report for details.

There was no salinity calibration sampling and no history for this instrument

since its last calibration, so no estimate can be made of those errors.

Salinity for this cruise has been edited but large errors likely remain.

SeaCat salinity is prone to error due to mismatch of conductivity and temperature

response times, especially when the descent rate is non-uniform. These errors are

likely to be much larger than calibration errors. SeaCat salinity errors are

expected to be as high as 0.05units in areas of high gradients.

The standards check routine was run and HEADEDIT rerun until all problems were resolved.

16. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Notes were entered in the headers of casts 6, 13 and 15 warning that there remain particular doubts about the locations of those casts.

Particulars

5/6. Duplicate files; not clear which is correctly identified – could both be wrong.

13. No position given in log but probably same site as cast #9 according to description.

14. No position given in log; duplicate of #13; log sheet says same site as cast #8.

13-14, No bottom depths listed.

15. Header says April 30 and log says April 29. Upload date is the 30th.  Time in header does not agree with log either. Location in log disagrees with header. Maximum sampling pressure too high for log location. Changed station name to Site #14. Position taken from station KIN1.5 of Dario Stucchi.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2005-24

Dates: 	Start:	29 March 2005	    End:	29 April 2005

Location: Broughton Archipelago

Vessel:	Clupea

Party Chief: Hargreaves B.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/19 SEACAT / 1294

Cruise ID#:		2005-24

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 22 Nov. 2002 | Factory |  |  |
| Conductivity | 1294 | 22 Nov. 2002 | Factory |  |  |
| Pressure Sensor | 163223 | 17 June 1999 | Factory |  |  |