## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected Salinity:Bottle lost during addition of HPLC. S.H. |
| 12 January 2021 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2005-25

Agency: OSD

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG - JdeF

Party Chief: Peña A.

Platform: Vector

Date: September 12, 2005 – September 18, 2005

Processed by: Germaine Gatien

Date of Processing: 11 January 2006 – 7 February 2006

Number of original CTD casts: 73

Number of CTD casts processed: 73

Number of rosette casts: 27

Number of rosette casts processed: 21 (no samples from 6 casts)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#498DR), a PAR sensor (#4565), a surface PAR sensor (#16504) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#0766) was mounted on the primary pump and a Seapoint Fluorometer (#2356) with a 10X cable was mounted on the secondary pump. The deck unit was a model 911 (#0508). The salinometer was an 8400B model Autosal (S/N 68572).

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book was in good order. The new section on the CTD test cast was completed, but the entry is out of line with other observations of the surface pressure. There were some errors in cast number on the rosette sheets, which led to some errors on the analysis spreadsheets.

It is recommended that salinity samples not be taken in turbid waters; such samples were taken at 7 casts and were outliers in the comparison between CTD and bottles.

There were notes in the rosette log that indicated a need for flags in the dissolved oxygen files, but only one was entered; there were a few comments entered, but they were in the wrong place.

The dissolved oxygen data in the CTD files should be considered:

- ±0.4ml/l in the top 150m

- ±0.1ml/l below 150m.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Note was made that there were some problems with bottle firing.

The test cast found a pressure of 1.36 ±0.5db on deck.

The rosette sheet for station 14 gives the cast number as 80; it should be 81.

Chlorophyll data were not available when processing began because of problems that needed further investigation.

The nutrient and salinity data were each obtained in spreadsheet format with flag channel and comments entered.

The dissolved oxygen data were obtained; a flag channel was present.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no errors were found.

All configuration files were the same and no errors were found, so 2005-25-0001.con was selected for conversion.

##### Conversion of Raw Data

The files were converted to CNV files

A few casts were examined and all expected channels were found and looked reasonable.

As usual for this area the descent rate was highly variable; it was high and almost constant for some casts, very noisy and fairly low for others. Bottle stops typically lasted about 30s.

The two temperature channels sometimes show somewhat different response to sharp features. The down and upcast temperature and conductivity channels are reasonably close, though the two channels vary more on the upcast than the downcast.

Transmissivity looked fine with good agreement between downcast and upcast.

The dissolved oxygen had the usual sort of offset.

The fluorescence looks ok.

The altimetry is often extremely noisy but generally looks reliable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s. The files for casts #43 and 59 were deleted since bottles were fired only for test purposes and no sampling was done.

All files were then converted to IOS HEADER format and renamed as *. BOT.

All BOT files were plotted and significant outliers were found only in cast #77; CTDEDIT was used to remove a small spike in the primary salinity around 50db, and the output was renamed as BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Three deep casts with a steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01,7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7) and (0.03, 9) were tried and there was little difference between several of them. A choice of (0.02, 7) looks reasonable for both pairs and was also selected for 2005-08 and 2005-10 when the same equipment was used.

CELLTM was run using (0.02, 7) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 280 | -0.0007 | +0.00015 | +0.002 | High, steady |
| 59 | 280 | -0.0005 | +0.00010 | +0.0017 | High, fairly steady |
| 59 | 405 | -0.0005 | +0.00013 | +0.002 | High, fairly steady |
| 66 | 280 | -0.0007 | +0.00015 | +0.002 | High, steady |
| 90 | 280 | -0.0007 | +0.00015 | +0.0017 | High, steady |

There were a number of casts with very steady descent rate, making the comparison easy. There was no noticeable variation in salinity differences with pressure. The primary salinity looked a little spiky in places. The two temperature channels differ a little more in detail than usual, but the average differences are small. There is no evidence of temporal drift in the differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run. No errors were found.

The average surface pressure is 1.6db which is a little low for the Vector. For cast #65 there are in-water values at a pressure of -0.4db at the end of the upcast. The pumps were turned off at 1db so the pumped channels may not be reliable, but the transmissivity does indicate that the CTD was below the surface. There are some records with zero transmissivity and fluorescence for pressure around -0.42db. The test at the jetty indicated a surface value of 1.4 ±0.5db. It seems likely that the sign was wrong in that reading, since the pressures are, if anything, low. As the pressure sensors age they are expected to drift to lower values. The con file includes an offset of 3.046db. A further offset of +.4db will be applied to this data.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. For several casts there was a rosette file, but the log indicates there was no sampling, so these files were not processed further. Cast #81 was originally named as #80 on the rosette sheet. It was later corrected but the oxygen and nutrient files were named as #80 and needed to be corrected in order to merge the data. Cast #73 was named as #72 on the rosette sheet which also resulted in the misnaming of samples.

Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data was provided in EXCEL format with a flag channel and comments. Channel names were changed to standard format and a column was added to which event numbers were added. The file was saved as 2005-25 Sal.csv. The files were converted to individual SAL files.

The nutrient data was provided in spreadsheet format. Channel names were changed to standard format and unneeded lines were removed; the file was saved as 2005-25-nuts.csv. The data was reordered on sample number and then converted to individual files.

The dissolved oxygen data was received in individual files with flag channel. Only one flag had been entered (cast #2) but was in the wrong column; a corresponding comment was moved as it was in the wrong place. There were some comments at the bottom of files explaining that some bottles did not close; the comments were removed. There are other comments on the rosette sheets that suggest that more flags may be necessary. Based on those notes the following samples were flagged “c” and comments entered:

- Samples 31 and 32 (cast 6) were flagged “c” and a note added that there is some confusion about sample numbers.

- Sample 45 (cast 9) is noted as having a “strange oxy curve”

- Samples 125 and 126 (cast #48): There were 4 bottles fired at 171db and the sample numbers are confused. On the chart there are single entries for dissolved oxygen samples by Niskin 3 and 4, so no indication that there were duplicate samples taken. But a note does suggest that these were both from Niskin 4. A “c” flag was entered by sample #125 and a note entered that it might be a replicate of 126.

- Samples 144 (cast #50) was flagged “c” because of notes on the rosette sheet about the stir bar being added late and sample #148 was flagged “d” because the stir bar was not added at all. The latter value was way out of line and the former a little out of line.

Cast #6 had a blank line and sample numbers were out of order and one sample number was wrong. Those were fixed with an editor.

The extracted chlorophyll data was in spreadsheet format and included a flag channel and comments. Many data were removed due to large variability between duplicates that could not be explained by natural variability, sampling or analysis methods, so the records were removed by the analyst. A comment concerning this will be added to the headers of all bottle files. The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The resulting file is 2005-25-chl.csv. That spreadsheet was converted to individual files. (Output: CHL)

The SAL, CHL, ADD and NUTS files were merged with SAMAVG in four steps. (Output: MRG1, MRG2, MRG3 and MRG)

##### 11. COMPARE

Salinity

There is a lot of scatter in the comparison of bottles and CTD. The sampling was relatively shallow so this is not surprising. Many of the samples are from the bottom. In the past problems have been noted in the quality of salinity bottle data in areas of high turbidity. This area is prone to such problems, so plots were made of all casts with salinity calibration sampling and low transmissivity (<20%) was noted at the bottom of casts #50 and 54, and fairly low (<40%) for casts #1, 35, 48, 57 and 75. The COMPARE plot shows that 6 of those 7 bottom samples are anomalously low. The bottom sample from cast #1 looks only slightly low.

When 4 other outliers were removed from the comparison a flat trendline was achieved with the primary high by 0.0006psu and the secondary high by 0.0032psu. There are, however, only 11 data points left. There is no obvious temporal drift.

The analyst had already put “c” flags on the two samples from casts #57 and 75 because 3 readings were needed. The bottom samples for casts #35, 48, 50, 54 and 57 were flagged “d”, the flag was changed from “c” to “d” for the bottom sample from cast #75. A comment was put in about sample #1 but it was not flagged.

Dissolved Oxygen

COMPARE was run. Outliers were identified by removing residuals <-.25 and >0.25 which produced a fairly tight relationship:

CTD-BOT = 1.3941 * DOX-CTD + 0.1253

The only outlier severe enough to be flagged was from cast #50 which had already been flagged.

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. The only significant outliers had already been flagged.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. There was a large scatter for fluorescence values higher than 5 mg/m3. When those were excluded the extracted chlorophyll values were found to be about 80% of those of the CTD fluorescence.  Removing all data with “c” flags did not make much difference. None of the outliers suggest problems with the fluorometer or extracted chlorophyll analysis.

##### Other Comparisons

Previous experience with these sensors –

- The primary conductivity sensors were used for 2005-11, -12, -21 and -22. From the latter two cruises there were a lot of bottles and the salinity was found to high by 0.0006 and 0.0005psu.  The secondary sensors were used for 2005-21 and -22 when they produced salinity high by 0.0027 and 0.0021psu.

- During the most recent uses of this DO sensor the results were:

CTD-BOT = 1.3705 DOX-CTD + 0.049   (2005-10) SoG/JdeF

CTD-BOT = 1.4103 DOX-CTD + 0.0742 (2005-14) SoG/JdeF

CTD-BOT = 1.4255 DOX-CTD + 0.1063 (2005-21) offshore

CTD-BOT = 1.4191 DOX-CTD + 0.0128 (2005-22) offshore

- The pressure sensor was believed to be reading a little low during 2005-22, but no recalibration was applied.

Historic ranges – There were many minor excursions from the local climatology limits. The only group of casts that fell entirely within the historic ranges were those from the western half of Juan de Fuca Strait. In most areas there were temperatures slightly above the maxima , but in the northern Strait of Georgia west of Texada Island the excursions looked more significant. The only salinity excursion was near-shore for one cast in Juan de Fuca Strait. Give the high variability in this region and the fact that the climatology does not contain recent data, these excursions are not considered to be evidence of calibration problems.

Post-cruise calibration – There was a post-cruise calibration for the DO sensor before repairs were made to a loose membrane. The drifts were read off the graph on the calibration sheet and compared to the effect of using the calibration from section 12. The graph readings are rough and somewhat scattered. The following table shows how the two compare:

| Original DO CTD | DO after calibration | effect of calibration | drift seen in Nov. 2005 at factory |
| --- | --- | --- | --- |
| 1.6 | 2.36 | 0.76 | 0.7 |
| 2 | 2.91 | 0.91 | 0.7 |
| 3.6 | 5.14 | 1.54 | 1.0 |
| 3.8 | 5.42 | 1.62 | 1.6 |
| 6.2 | 8.77 | 2.57 | 1.7 |
| 7.4 | 10.44 | 3.04 | 3.0 |

##### The agreement is reasonable given that the graph readings are very rough and the calibration reliability is known to be limited by the time response of the sensor.

##### 13. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were on the order of 1s. A shift of +24 records (1s) is the shift that has been used in most other cruises. All data was shifted by +24records. (Output: SHFFL)

Dissolved Oxygen

Tests were run using +120, +130, +140, +150 and +160 records on a few casts. As well as making the trace offset similar to those of temperature we also judge by how the DO looks during a bottle stop. We expect the DO to settle to the correct value by the end of the bottle stop. That value should be roughly half way between the up and down values in motion. The results vary with DO gradient, so there is no clear RIGHT answer. However, the best choice overall appears to be to advance the DO channel by +130 to +150 records. For most of the 2005 cruises during which this sensor was used, a setting of +140 was chosen. So +140 records was applied to the DO channel for all casts.

Conductivity

Tests were run on a 4 casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. There was very little change but the best settings proved to be -0.4 for the primary and +0.4 for the secondary. Those parameters were applied to all casts. (Output: *.SHFC0 and SHFC1).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salinity

Maximum Surface Pressure (relative): 10.00

Minimum Salinity: 5

Surface Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

##### DETAILED EDITING

The primary sensors were slightly closer to the bottles and the calibration was consistent with other recent cruises when the data from those sensors were archived.

Page plots were produced using (T0, S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used. The descent rate was generally very steady with the exception of some casts in Juan de Fuca Strait.

All casts required some editing. The following casts required heavy editing: 5-9, 21-22.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

The primary salinity is in good agreement with the bottles, and was so for 2005-21 and 2005-22 as well. No recalibration will be applied to the salinity channel.

From section 11 we have the following equation for recalibration of DOX:

CTD-BOT = 1.3941 * DOX-CTD + 0.1253

File 2005-25-recal1.ccf was prepared to apply the above equation to the DO channel and to add 0.4db to the pressure channel to the SAM and MRG files. COMPARE was rerun for dissolved oxygen and when the same points were left out of the comparison as in the first run, the average difference is -0.00009ml/l indicating that the recalibration worked properly. (See 2005-25-dox-comp2.xls.)

The edited downcast files, EDT, were recalibrated using 2005-25-recal1.ccf. (Output:COR)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure		Averaging interval = 1.000

Minimum bin value =   .000	Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

The averaged downcast files were thinned and the CTD DO values were compared with the upcast bottle DO values. When plotted against pressure and removing outliers identified by residuals, the trendline implied the following correction: (See 2005-25-dox-comp3.xls.)

CTD DOX (Corrected) = CTD DOX + 0.0005*Pressure - 0.1395

That correction was applied to the thinned files using calibration control file 2005-25-recal2.ccf. COMPARE was rerun and the results were satisfactory, so the correction was applied to the bin-averaged downcast files. (See 2005-25-dox-comp4.xls.)

The second recalibration will not be applied to the bottle files because it only applies to data collected while the CTD was moving.

The dissolved oxygen data in the CTD files should be considered:

- ±0.4ml/l in the top 150m

- ±0.1ml/l below 150m.

##### Special Fluorometer Processing

All COR files were put through CLIP to produce files with data to 100db only for the use of A. Peña. Those files were then bin-averaged (1/4db bins), put through REMOVE to remove extraneous channels, recalibrated using 2005-25-recal2.ccf and HEADEDIT to fix formats and channel names. The final files (FCTD) were saved in a separate directory. The recalibrated CTD bottle files (SAMCOR) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.4ml/l from   0 -  150m

•	±0.1ml/l below 150m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

##### Final Bottle Files

The MRGCOR files were put through CLEAN to remove the SeaBird headers.

SORT was used to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods, plus the analyst’s remark about the rejection of some data. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared using the edited data and displaying T and S. Profile plots were produced showing PAR, DO, fluorescence and transmissivity.

##### Cross-reference lists were produced for the final CTD and CHE file.

The sensor histories were updated.

Particulars from logs

6. Some confusion over DO samples at 5db.

30. Pump turned off at bottom.

35. Bottle #1 did not trip. Bottle #15 empty. Problem with computer.

43. Bottle #1 fired only to test, no sampling.

44-45. Bottle #1 fired, but no sampling.

48. No DO and Nuts from bottom 3 bottles.

50. Note on rosette sheet for sample #144: “Oxy stirbar added late” and for sample #148 “stir bar not added”. Sample 145 has note “called 154” but there is a sample #154 as well, and no sign of an error in the DO file, so perhaps this was fixed at sea.

57. Bottle #15 did not fire.

59. Bridge depth 423, lab depth 365. Bottle #15 fired at 410db just for test.

60. Bottle #1 fired, but no sampling.

65. Bottle #1 fired, but no sampling.

77. Bottle #15 did not fire.

81. Bottle #15 did not fire.

88. Pump not on, restarted.

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-25 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 12 Sept. 2005 End: 18 Sept. 2005 |  |  |  |  |  |  |
| Location: SoG/JdeF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Peña A. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-25

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19/03/05 | Factory “ |  |  |
| Conductivity | 3038 | 03/05/05 | “ |  |  |
| Secondary Temp. | 2106 | 08/07/05 | “ |  |  |
| Secondary Cond. | 1729 | 12/07/05 | “ |  |  |
| Transmissometer | 498DR | 22/05/05 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 16/11/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| PAR | 4565 | 23/12/04 |  |  |  |
| Surface PAR | 16504 | 02/01/04 |  |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |