## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2005-27

Agency: IOS, Ocean Sciences Division, Sidney, B.C.

Project: Sea Lice

Chief Scientist: Sweeting R.

Platform: VECTOR

Date: 9 July 2005 – 12 July 2005

Processed by: Germaine Gatien

Date of Processing: 18 January 2006 –  24 January 2006

Number of original CTD casts: 17

Number of casts processed: 17

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (#0334) was run with pressure sensor 290544 and transmissometer #197. Salinity samples were collected using a Niskin bottles

# SUMMARY OF QUALITY AND CONCERNS

The CTD Daily Log Book gives a list of personnel but only first names.

The file names were non-standard.

There was no salinity calibration sampling available for this cruise at the time of processing. According to the log samples were taken. Other cruises in 2005 using the same CTD found salinity to be low by 0.004psu (2005-17) and by about 0.001psu (2005-30). The latter comparison had many bottles.

Upcast data was selected because the top 15 to 30m of downcasts data were unusable. It is likely that the pump did not come on because of low conductivity.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

The Daily Log was obtained. Note was made of spikes in the data due to transmission errors. This was also noticed during 2005-30.

While there were salinity samples taken during this mission, the data could not be found. There is a note in the log that they were labelled as 2005-17, but that file was checked without success. There was salinity sampling during 2005-17 and 2005-30, so those results will be applied to this mission.

The cruise summary sheet was completed.

The configuration file from 2005-17 was obtained and the calibration constants were checked and found to be correct.

## 3. Conversion of Raw Data

The file names were changed to standard format.

The data was converted using the 0334CTD.con.

A few casts were plotted. The downcasts and upcasts were reasonably similar for most casts, but for some there were significant differences in the top 50m to 100m. This is likely due to the pump being set to come on at a conductivity not reached in surface waters in this area. Similar problems occurred during 2005-30 until a change was made to the pump turn-on parameter.

The pressure signal has a spike in at least one cast; during 2005-30 there were many such spikes which were removed using WILDEDIT. The pressure has many steps of about 1db; the manufacturer states the resolution is 1db.

All casts checked had a constant transmissivity value. This should be rechecked later, and that channel removed if there is no signal. There is no transmissometer listed in the log book, but only the CTD serial number is given with no sensors specified.

Conductivity becomes non-zero at P~1db.

The descent rate is generally moderate and steady.

The headers do not contain positions.

4. WILDEDIT

WILDEDIT was run on all casts on pressure, temperature and conductivity channels using 2,20,25,0 for “Standard deviations for pass 1” and “Standard deviations for pass 2”, scans per block and “Keep data within this distance of the mean”.

5. ALIGN

This step was skipped. The conductivity will be aligned later using SHIFT.

6. WFILTER

As was found during the 2001 Arctic cruises and during 2003-04, 2005-17 and 2005-30 the SBE 25 pressure is noisy.

In processing the data from 2005-30 tests were done on filters and a cosine, size 5 produced results that left a few unstable features, but removed most. The same parameters were selected for this data.

The SeaSoft routine WFILTER was run for all casts to apply a cosine filter, size 5, to the pressure, temperature and conductivity.

7. CELLTM

Based on results for the 2001 Arctic data, 2003-04, 2005-17 and 2005-30, CELLTM was run on all casts using (α = 0.03, 1/β = 9.0).

8. DERIVE

Program DERIVE was run to calculate salinity.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 25 data to IOS Headers.

The headers lack positions and station names. The time in the headers is correct and the water depth is in the headers for about half the casts.

A text editor was used to add longitude, latitude and station name (and water depth where necessary) to the headers based on the Daily Log Book entries.

Plots were made to check on the transmissivity and the only values were 0 or 35%/m. This channel should be removed later.

CLEAN was run to replace pad values in the pressure channel using interpolation based on scan number.

10. Checking Headers

The header check and header summary were run and no errors were found. The times in the log are in PDT, but the times in the CTD files are in UTC. The speeds from casts #11 to 12 and #13 to 14 seem a little high, but the positions agree with the log book, and on a track plot the station positions look correct, so it is assumed that the Frosti is capable of higher speeds in tight quarters than the Vector.

Track plots were produced and appended to this report.

A surface report was produced and all pressures were close to zero, but positive. The initial conductivity values are very low. Because the CTD pump was set to turn on after a minimum conductivity was reached, it is assumed that all values are “in-water”. The average surface value was 0.28db.

11. SHIFT

Based on the results of 2005-17 and 2005-30 SHIFT was run on all casts advancing the conductivity by +0.7 records. (Output: SHF)

- REVERSE

Because the data from the top 50m is bad for some casts, all casts were put through REVERSE, so upcast data is available. (Output: SHFREV)

- DELETE

The SHF and SHFREV files were put through DELETE using the following parameters:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.2m/s (calculated over 7 points) was deleted *

Sample interval taken from the header.

Pressure was not filtered.

* After an initial run using 0.3m/s it was decided that was too severe for the upcast data. There were large sections of data with low but very steady descent rate, so using 0.3 removed data that looked ok.

COMMENTS ON WARNINGS: There were no warnings.

- DETAILED EDITING

Each cast was examined to see if upcast and downcast differ significantly in the top 30m. The downcasts look unusable in the top 15m for all casts, and to more than 30m for many. Upcast data will be selected.

Page plots were produced using the DELREV files; they were examined for spikes and instabilities and used to guide the use of CTDEDIT.

CTDEDIT was used to clean noise in S. Near-surface records were removed from most casts. Salinity was cleaned where there were spikes or “overshoots” suspected to be due to misalignment of T and C. Where data corrupted by shed wakes was identifiable, it was removed.

All casts required some editing. Only cast #13 required heavy editing (from 175-225db).

Note was made of the editing details in the files.

- Intercomparisons

Sensor History –For 2005-17 in June the salinity was found to be low by about 0.004psu.

Historic ranges – There is no useful climatology for this area.

Comparison of nearby casts – Temperature and salinity for groups of nearby casts were plotted together to check for unbelievable values; no problems were noted. Casts #3 and 19 were about 15km apart and at the bottom they were within 0.002psu of each other at constant density.

- Recalibration

Based on the comparisons of 2005-30 the salinity was found to be low by less than 0.001psu when a few outliers were excluded. There were many bottles but they were not done with a rosette, so the comparison is not as certain as usual. No recalibration will be applied to this data set.

- BIN AVERAGE

The following Bin Average values were applied to the EDT files:

Bin channel = pressure

Averaging interval = 1db

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins.

- Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Scan_number; Conductivity:Primary, Transmissivity, Descent Rate, Bottle Number and Flag.

- HEADER EDIT

Header Edit was used to change the agency name and to fix formats and channel.

- Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

c.) The conductivity and pressure sensor history files were updated.

Particulars

Times in log book are all PDT.

3. CTD to bottom, touched weight

5. Salinity calibration sample 1m above CTD

8, 13, 15, 17. Salinity calibration samples.

16. CTD touched bottom.

## Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2005-27

Dates: Start: 9 July 2005 End: 12 July 2005

Location: Kingcome Inlet and Broughton Archipelago

Vessel: Frosti

Party Chief: Sweeting R.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334		Cruise ID#:		2005-27

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 290544 | 13May05 | Factory |  |  |
| Temperature | 4054 | 14Oct04 | Factory |  |  |
| Conductivity | 2754 | 15Oct04 | Factory |  |  |
| Transmissometer | 197 | 20June2005 | Factory |  |  |

i