## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 6Feb. 2019 | Bottle spreadsheet converted to searchable BOT files. |

## PROCESSING NOTES

Cruise: 2005-28

Agency: PBS, Ocean Sciences Division, Sidney, B.C.

Project: High Seas Salmon

Chief Scientist: Morris J.

Platform: Frosti

Date: 26 October 2005 – 23 November 2005

Processed by: Germaine Gatien

Date of Processing: 18 February 2006 – 2 March 2006

Number of original CTD casts: 129

Number of casts processed:  127 (battery failure at surface for 2 casts)

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (#0334) was run with pressure sensor 290544 and SBE43 dissolved oxygen sensor (#0047). A Niskin bottle was mounted on the wire.

# SUMMARY OF QUALITY AND CONCERNS

The file names were non-standard.

There were no positions in the headers and many dates and times were wrong in the headers; that information was obtained from the CTD Log Book. Corrections were done using a text editor and were very time-consuming.

The CTD Log Daily Log book was in good shape and included remarks about equipment failures.

There were only 2 deep salinity samples available for calibration although there were many surface samples, some of which were in well-mixed water suitable for calibration.

Because there was no DO sampling the CTD DO channel will not be archived. However, a set of files with DO was prepared for the use of Marc Trudel. Calibration was based on the results of 2005-30.

The salinity calibration is considered to be ±0.002psu, but there are other errors due to variations in alignment as the descent rate varies. Errors are expected to be on the order of 0.01psu in high gradient areas and much smaller at depth.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX. The file names are non-standard, missing a 0.

2. Preliminary Steps

The Daily Log was obtained. Note was made of problems with the CTD battery and computer clock errors. The time changed from PDT to PST during the cruise. Two deep salinity calibration samples are mentioned and there were many surface samples.

The cruise summary sheet was completed.

A configuration file was obtained and the calibration constants were checked and found to be correct.

The salinity bottle data was obtained in spreadsheet 2005-28-sal.xls; it was simplified and saved as 2005-28-sal.csv.

The file names were changed to standard format.

## 3. Conversion of Raw Data

The data was converted using the 0334CTD.con. The first two casts could not be converted until a separate header file was used. The header of those 2 HEX files had errors in the sensor numbers.

A few casts were plotted. During 2005-30 there was a severe problem with spiking due to transmission problems. During this mission the pressure has only tiny spikes and steps ~0.2db for most casts; the manufacturer states the resolution is 1db. There were large spikes in a few casts, one of which had only 6db of data, but according to notes in the log this appears to be due to battery failure. Conductivity becomes non-zero at P~0.2db during upcasts. The downcasts generally started at about 1 to 2db. The temperature and conductivity look ok, and the dissolved oxygen shows the usual offset between downcast and upcast. The descent rate is generally high and varies from extremely quiet to extremely noisy.

The headers do not contain positions.

4. WILDEDIT

WILDEDIT was run on all casts on pressure, temperature and conductivity channels using 2,20,25,0 for “Standard deviations for pass 1” and “Standard deviations for pass 2”, scans per block and “Keep data within this distance of the mean”.  Plots show that WILDEDIT got rid of some pressure spikes.

5. ALIGN

To check the alignment of the dissolved oxygen channel, tests were run on 3 casts to see which advancement made the differences between the downcast and upcast look most alike. A setting of +6s or +7s relative to pressure worked best. Since there was no DO calibration sampling for this cruise, the results of 2005-30 will have to be used for calibration. So it is best to use the same alignment as well, to ensure the calibration is appropriate. So +6s was applied to all casts.

The conductivity alignment was also checked. The temperature and conductivity channels seem fairly well aligned. This will be fine-tuned later using SHIFT.

ALIGNCTD was run with a setting of +6s applied to the dissolved oxygen channel.

6. WFILTER

As is generally the case with an SBE 25 CTD, there is small-scale noise in the pressure, conductivity and temperature channels. For 2003-04, 2005-17 and 2005-30 a cosine filter of size 5 was found to do a good job in removing reversals without oversmoothing.

The SeaSoft routine WFILTER was run for all casts to apply a cosine filter, size 5, to the pressure, temperature and conductivity. A few casts were examined to ensure that the choice was appropriate and it was found that it had removed most noise without oversmoothing.

7. CELLTM

Tests were run using CELLTM on 3 casts with choices for (α,1/β) being (0,0), (2,9), (3,9) (2,7), (3,7) and (2.45,9.5). The best results were with α = 0.03, 1/β = 9 in each case.

For the 2001 Arctic data, 2003-04 and 2005-17 the same results were found.

CELLTM was run on all casts using (α = 0.03, 1/β = 9.0).

8. DERIVE

Program DERIVE was run to calculate salinity.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 25 data to IOS Headers.

The headers lack positions and station names. The time is correct.

Edit Headers was used to add lines for longitude, latitude and water depth in the Location section of the headers. A text editor was then used to enter the longitude, latitude and water depths based on the Daily Log Book entries. Test track plots were produced to check that no gross errors were made in entering the data, and …..

CLEAN was run to add event numbers and to replace pad values in the pressure channel using interpolation based on scan number and to ensure the pad values for the DO channel are correct.

10. Checking Headers

The header check was run and a few errors in headers were found and corrected.

The header summary was printed and checked against the log entries. A few errors in position and station name were found, and errors were found in date and time from cast #92 to 177, when the computer clock was fixed. These errors were corrected.

The log gives the file name and event number for station T05 as 385, but the data file was actually named as 384 which was a bongo cast according to the log. That was changed to match the log.

The header check was rerun and no errors found.

The header summary was rerun and checked against the log for the casts that had been found incorrect on the first run; one more error was found and corrected.

Track plots were produced and appended to this report.

A surface report was produced and the average surface value was 1.2db.

The mixed layer is deep enough that the following casts should be considered for a comparison of CTD salinity and surface salinity bottles for calibration purposes: 29-41, 50, 56, 141, 195, 237, 243, 327, 330, 333, 339, 357-378.

11. SHIFT

A few casts were selected to study alignment. Tests were run advancing the conductivity by between 0 and +0.9 records relative to temperature and the results examined in T-S space to see if the salinity became more or less stable. The best results were found when the shift was +0.7 or +0.9 records, with the latter perhaps a little over-smoothed. For the 2001 Arctic data +0.3 records was used, for 2003-04 the setting used was +0.4 records and for 2005-17 and 2005-30 it was +0.7 records.

SHIFT was run on all casts advancing the conductivity by +0.7 records.

Tests were also run on the dissolved oxygen channel to see if the alignment could be improved further, but all settings made it worse in the top 50m and any improvements below that were slight. The ALIGNCTD routine appears to have the best possible setting.

- DELETE

The files were put through DELETE using the following parameters:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 7 points) was deleted

Sample interval taken from the header.

Pressure was not filtered.

COMMENTS ON WARNINGS: The only warnings pertained to two very shallow casts, #83 and #86. These were casts during which the battery failed and only a few records were acquired. They will not be processed further.

All DEL files were copied to EDT files.

- DETAILED EDITING

Page plots were produced and examined for spikes and instabilities and used to guide the use of CTDEDIT. All casts were examined in CTDEDIT. A little surface data was removed from most casts. Salinity was cleaned where there were spikes or “overshoots” suspected to be due to misalignment of T and C or where there was evidence of corruption due to shed wakes, especially at the bottom of casts.

Cast #71 had gaps in the data; an examination of the SHF file showed that there were spikes near the surface that led to DELETE making inappropriate choices. The first 17 records were removed from the SHF file and DELETE was rerun.

For cast #80 a patch of spiky data was removed by making records #982 to 987 match record #981 and DELETE was rerun.

The following casts were edited heavily: 29, 50.

The following casts required no editing: 11, 44, 71, 261, 312, 348, 351.

All other casts were edited lightly. Note was made of the editing details in the files.

- BIN AVERAGE

The following Bin Average values were applied to the EDT files:

Bin channel = pressure

Averaging interval = 1db

Minimum bin value =   .000	Average value will be used

Interpolated values are NOT used for empty bins.

The output files were named AVG.

- Comparison of bottles to CTD salinity

There were two deep salinity calibration samples and many shallow ones. Based on a mixed-layer depth calculation done during the surface check, 14 casts were found with surface samples and low salinity gradients near the surface. Plotting descent rate against pressure indicated that the CTD stopped between 9.5 and 12.5db, with 11db being the average. The bottle is believed to have been mounted 5db above the CTD, so CTD data from 6db was used for the surface comparison. The firing levels for the 2 deep bottles were determined in similar fashion. The edited files were averaged using 0.5db bins and thinned to values from 1db above firing to 1db below. The differences were found between CTD salinity 1db above and below the bottle firing level, and where the difference was >0.005psu they were deemed to have a moderate gradient. Others were described as having a low gradient.

When all 16 bottles were included the CTD was found to be reading high by 0.002psu. When the casts with moderate gradients were excluded the CTD was high by 0.0026psu. When 1 outlier was excluded from that group the CTD was found to be low by 0.0017psu. Of the two deep bottles the one with low gradient gave CTD salinity high by 0.0007psu and the one with moderate gradient showed it to be low by 0.0066psu. There is a lot of scatter in the results, but it is probably a reasonable conclusion that the CTD is within 0.002psu of the bottles. (See 2005-228-sal-comp.xls.)

- Other Intercomparisons

Sensor History

- SALINITY: This CTD was used during 3 other cruises in 2005. During 2005-17 there was only surface sampling and the results had a lot of scatter; the CTD appeared to be reading low by about 0.004psu. During 2005-27 there was no calibration sampling. There were many samples during 2005-30 using multiple Niskin bottles on the wire with the CTD. There was a lot of scatter, but when a few outliers were excluded the CTD was found to be low by about 0.001psu.

- DISSOLVED OXYGEN: The DO sensor was used for 3 cruises in May, June and October 2005. Because of anoxic conditions 2005-11 required different corrections for different parts of the cruise. For 2005-12 the following correction was applied:

DO Corrected = 1.342 * DO-CTD – 0.0143

And for 2005-30, with many hydro bottles but no rosette, the following correction was used, but the comparison was done with upcast data:

DO Corrected = 1.574 * DO-CTD – 0.154

Historic ranges –Most of the data fell within the historic ranges, but there were excursions. Many of the casts were in inlets that are probably not well represented in the climatology. The surface salinity was often lower than the historic minima and the temperature sometimes fell higher than the historic maxima at the bottom of the mixed layer, which may mean that it was deeper than during other cruises represented in the climatology. Since there were excursions to both higher and lower salinity and temperature there does not appear to be a systematic instrumental error.

Comparison of nearby casts – T-S plots for groups of nearby casts were checked for unbelievable values and none were found. The differences of nearby casts were small at depth. For example, the differences between casts #165 and 168, about 20km apart, at about 200m were on the order of 0.002psu and less than 0.01Cº.

- Recalibration

The bottle comparison and that during 2005-30 suggest that the CTD salinity is within 0.002psu. No recalibration will be applied.

There was no calibration sample for DO. A method was found to recalibrate this data for the use of Marc Trudel, but it was judged unwise to put the data in the DATA_LIB archive. Document “2005-28-DO-calibration-method.doc” describes how the calibration method was derived and “DO calibration study.xls” contains a comparison of different calibration methods. (In brief, the calibration used for upcast data from 2005-30 was applied to a subset of the upcast data from 2005-28. Then subsets of the downcast and upcast data were compared to derive a calibration equation applicable to downcast data.)

File 2005-28-recale.ccf was prepared to apply the following correction to the dissolved oxygen channel:

DO Corrected = 1.5925 * DO-CTD – 0.2665

The output files were given extension COR.

- REMOVE

The following channels were removed from the AVG (bin-averaged but uncalibrated) files for all casts: Scan_number; Conductivity:Primary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Descent Rate and Flag.

NOTE: These files had not been recalibrated as described above as they are intended to go in to the

archive.

- HEADER EDIT

Header Edit was used to fix the agency name (PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.) and to fix formats and channel names and to add the following note to the headers:

The salinity calibration at rest is considered ±0.002psu.

Other instrumental errors, including variations in alignment due to descent rate variations, are expected to cause errors in salinity of up to 0.01psu in high gradient areas and much smaller errors at depth.

The dissolved oxygen channel was removed from the file because there was no calibration sampling during this cruise.

The files produced were given extension CTD.

- Special files for Marc Trudel

Special files were prepared for Marc Trudel using the COR files described in section 17.

Those files contained some negative or unbelievably low values of DO near the surface; CLEAN was used to replace with pad values all DO readings at pressures <2db. Plots were then examined and another12 casts were identified with unbelievable surface values. Those values were replaced with pad values using a text editor.

Then all files were put through CHANGE UNITS to add a second DO channel with umol/kg units. REORDER was used to get the two DO channels together.

REMOVE was run with the only the DO voltage channel selected, and text file 2005-28-hdr3.txt was used when HEADER EDIT was run. These files were given the extension CTDX.

- Final steps

a.) THIN and DERIVE were run to obtain values for tables and page plots were prepared using the final files.

b.) A cross-reference listing was produced.

c.) The conductivity, DO and pressure sensor history files were updated.

Particulars

31. Off-course from designated transect. Too far north.

50. Time change to PST.

71. CTD battery failure - 40,000 data points.

80. CTD battery failure – 30,000 data points.

83. CTD battery failure – only 29 records. DELETE

85. CTD battery failure – only 13 records. DELETE

178. Computer data and time changed after cast due to error in both.

216. Deep salinity sample at 400m.

228. Stopped at 150m because of shallowing depth.

279. Niskin bottle failure.

309. Deep salinity sample at 350m.

## Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2005-28

Dates: Start: 26 October 2005 End: 23 November 2005

Location: NE Pacific

Vessel: Frosti

Party Chief: Morris J.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334		Cruise ID#:		2005-28

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 290544 | 13May05 | Factory |  |  |
| Temperature | 4054 | 14Oct04 | Factory |  |  |
| Conductivity | 2754 | 15Oct04 | Factory |  |  |
| Dissolved Oxygen | 0047 | 10Feb2005 | Factory |  |  |