## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 18 April 2006 | CHE files fixed for casts 20,23,29,39,41,45,54,56,64,67 See note at end |
| 14 May 2007 | CTD Fluorescence restored to CHE files to first 8 of files listed above – channel got lost in previous correction. |

## PROCESSING NOTES

Cruise: 2005-32

Agency: Ocean Sciences Division

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG - JdeF

Party Chief: Masson D.

Platform: Vector

Date: November 28, 2005 – December 2, 2005

Processed by: Germaine Gatien

Date of Processing: 21 March 2006 – 11 April 2006

Number of original CTD casts: 72

Number of CTD casts processed: 72

Number of rosette casts: 20

Number of rosette casts processed: 20

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#498DR), a PAR sensor (#4656), a surface PAR sensor (#16504) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#0047) was mounted on the primary pump and a Seapoint Fluorometer (#2356) with a 10X cable was mounted on the secondary pump. The deck unit was a model 911 (#0508). The salinometer was an 8400B model Autosal (S/N 68572).

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book and rosette sheets were generally in good order, but there were some inconsistencies in cast numbers. The intended first cast was abandoned and the next cast was called #1 in the CTD log and files and #2 in the rosette log and analysis results. The CTD file names were changed to 2. Cast #23 was misidentified in the rosette log so that analysis files did not have the same names as the CTD files.

Stops for rosette sampling were mostly very short.

The Reference PAR data may have been affected by the presence of dots of paint on the sensor.

The salinity analysis results are considered unreliable; this is probably due an error made by the analyst who was using the new Autosal for the first time. As a result the recalibration of salinity was based on an earlier cruise. This can be revisited in future when post-cruise calibrations or the results of other cruises become available.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 50m

- ±0.3ml/l from 50m to 150m

- ±0.1ml/l below 150m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. It is noted in the log that the pressure offset was changed from 3.045 to 3.545 in file 0443ctd.con after cast #2. There is no cast #2, but it is presumed that the note refers to cast #1. There is no difference in configuration files from casts 1, 3, 4 and 5 and no file 0443ctd.con was found. The calibrations were checked and no errors were found. It is assumed that the comment is based on a test at sea, and it makes sense based on the history of the instrument.

When the Reference PAR sensor (SPAR) was examined after this cruise it was found that there were small red dots of paint on the sensor and it is likely that this occurred during this cruise; this might have a small effect on the data.

The chlorophyll, nutrient and salinity data were each obtained in spreadsheet format with flag channel and comments entered.

The dissolved oxygen data were obtained; a flag channel and comments were included.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

The files were converted to CNV files

The confusion over cast numbers continues in that the bottle analysis files all refer to the first event as #2, whereas the CTD file is named as #1. It seems easiest to change the CTD files since there is no event #2 in the log book.  The files names were changed from 2005-32-0001 to 2005-32-0002 after conversion.

A few casts were examined and all expected channels were found and looked reasonable.

As usual for this area the descent rate was highly variable; it was high and almost constant for some casts, very noisy and fairly low for others.

Bottle stops typically lasted much less than 30s. Only the first three rosette casts had stops that were mostly 30s or longer. Some stops were less than 10s.

The down and upcast temperature and conductivity channels are reasonably close, though the two channels vary more on the upcast than the downcast.

Transmissivity looked fine with good agreement between downcast and upcast.

The dissolved oxygen had the usual sort of offset.

The fluorescence looks ok.

The altimetry is often extremely noisy but generally looks reliable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format. They were put through CLEAN to add event numbers and named as *. BOT.

All BOT files were plotted and no significant outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.  (It was discovered later that WILDEDIT was applied to PAR and PAR:Reference as well which was a mistake. A correction was made later in the processing. See section 16 for details.)

Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Three deep casts with a steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01,7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7) and (0.03, 9) were tried and there was little difference between several of them. There are significant differences between the upcast and downcast data so that the judgment of the optimal setting is difficult. A choice of (0.03, 9) looks best for the primary and (0.02, 7) for the secondary. CELLTM was run using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 29 | 150 | -0.0012 | +0.00016 | +0.003 | Fairly high, fairly steady |
| 44 | 150 | -0.0013 | +0.00009 | +0.002 | Fairly high, very steady |
| 44 | 275 | -0.0013 | +0.00012 | +0.0025 | Fairly high, very steady |
| 72 | 150 | -0.0012 | +0.00007 | +0.002 | Fairly high, very steady |
| 72 | 275 | -0.0010 | +0.00007 | +0.0015 | Fairly high, very steady |
| 72 | 330 | -0.0012 | +0.00010 | +0.002 | Fairly high, very steady |

The temperature differences are much larger than usual even when the descent rate was very steady. Where those differences are noisiest there appear to be differences in alignment between the two temperature traces rather than spikes in one or the other. There is no evidence of temporal drift in the differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check showed that station names were missing from the first few casts. Those were added to the headers. The header summary was run and checks made for accuracy. The only error found was that while the file name of the first cast had been changed from 1 to 2, the event number in the header had not been changed. That was fixed. One discrepancy was found between the headers and CTD log, but the error is clearly in the log book.

The average surface pressure is 1.4db which is a little low. A few casts were examined to see if the records with pressures <0.3db have reasonable near-surface values. The salinity was low in each case, but in this region it is not obvious what is near surface and what is not. Since the pressure offset was based on a test at sea, pressure will not be recalibrated further.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data was provided in EXCEL format with flag and comment channels. No data had been flagged. Channel names were changed to standard format and the file was saved as 2005-32 Sal.csv. The files were converted to individual SAL files. The file identified as from cast #22 was changed to #23.

The nutrient data was provided in spreadsheet format with flag channel and comments. Channel names were changed to standard format; the file was saved as 2005-32-nuts.csv. The data was reordered on sample number and then converted to individual files. There were a few errors in pad values; these were fixed with a text editor. The file identified as from cast #22 was changed to #23.

The dissolved oxygen data was received in individual files with flag channel, but the flags were in the wrong column. These were fixed using a text editor. The file identified as from cast #22 was changed to #23.

The extracted chlorophyll data was in spreadsheet format and included a flag channel and comments.. The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The resulting file is 2005-32-chl.csv. That spreadsheet was converted to individual files. The following comment was added to the headers of the CHL files: “Average of two samples is reported. Variability is assessed as the % (std dev/mean*100).”

The file identified as from cast #22 was changed to #23. (Output: CHL)

The SAL, CHL, ADD and NUTS files were merged with SAMAVG in four steps. (Output: MRG1, MRG2, MRG3 and MRG)

##### 11. COMPARE

Salinity

The CTD primary and secondary channels produced salinity that was lower than that of the bottles by about 0.050 and 0.047. There are hints of time and pressure dependence that disappear if two records are dropped. There is a lot of scatter in the results. It is suspicious that both sensors look so much alike. The differences between the two CTD channels are consistent with the observations of section 7. It is hard to imagine what problem would cause both pairs to be off by so much and by almost exactly the same amount. This looks like a problem with the salinometer. The analyst reports that this analysis was his first using a new Autosal salinometer and he thinks it is possible he made an error in the setting of the potentiometer.

All salinity samples were be flagged “d” with the comment “Samples consistently differ from both CTD channels by about 0.05; the analyst reports a likely error in use of the salinometer.”

Dissolved Oxygen

COMPARE was run. Plotting the differences versus CTD DO excluding any samples flagged “d”, 6 outliers were identified by removing differences >10 or <-10 and residuals <-.2 and >0.2. This produced a fairly tight relationship. When differences were plotted versus pressure one other outlier was identified, and the fit was then found to be:

CTD-BOT = 1.4614 * DOX-CTD - 0.032

The fit versus time shows no hint of temporal drift in calibration.

Two outliers were flagged “c” with notes of explanation in the headers (cast #17, sample #55 and cast #51, sample #172). Comments were added to the headers for a few samples that already had been flagged, confirming the validity of the flags.

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. A few slight outliers were investigated, but there was nothing to suggest that further flagging of bottle DO data was justified.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. The fluorescence is almost twice as high as the titrated chlorophyll which is unusual when the values are low.

##### Other Comparisons

Previous experience with these sensors –

- The primary conductivity sensors were used for 2005-16 when the bottle comparison had a lot of scatter and showed the CTD salinity to be high by about 0.0047.  The secondary sensors were used for 2005-21, -22 and -25 when they produced salinity high by 0.0027, 0.0021 and 0.0032.

- During the most recent uses of this DO sensor the results were:

CTD-BOT = 1.3423 DOX-CTD - 0.143   (2005-12) WCVI

CTD-BOT = 1.574  DOX-CTD - 0.154     (2005-30) Broughton

Historic ranges – The temperatures below 120m in the northern part of the Strait of Georgia were frequently above the historic range, but that is likely due to the limits of the data in the climatology; there have been frequent observations of high temperatures in that region in recent years. There were a few high salinity values in the same area, but only slightly above the historic maxima. These excursions are not considered indicative of instrument calibration problems.

##### 13. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were on the order of 1s. A shift of +24 records (1s) is the shift that has been used in most other cruises. All data was shifted by +24records. (Output: SHFFL)

Dissolved Oxygen

Tests were run using +120, +130 and +140 records on a few casts. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by +130 records (about 5.5s). So +130 records was applied to the DO channel for all casts.

Conductivity

Tests were run on a 2 casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.4 for both sensors. That parameter was applied to all casts. (Output: *.SHFC0 and SHFC1).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salinity

Maximum Surface Pressure (relative): 10.00

Minimum Salinity: 5

Surface Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

##### DETAILED EDITING

The primary sensors show fewer spikes and were chosen for 2005-16, but the salinity produced was off by about 0.005 during that cruise. There is more history for the secondary sensors and they have produced salinity that is consistently high, but the differences varied from 0.0012 to 0.0032. Overall the primary sensors look better.

Page plots were produced and used to guide the editing. On-screen plots of descent rate and pump status were also used. The descent rate was often very noisy early in the cruise, but was very steady after the Juan de Fuca section.

All casts required some editing. Note was made of the editing details in the headers of the relevant files.

##### PAR and SPAR reprocessing

At this point it was discovered that there were many pad values in the PAR and SPAR data because WILDEDIT had been accidentally applied to those channels. So the original CNV files were stripped of most channels, converted to IOS HEADER format and put through DELETE and CLEAN; these will be treated as the SECONDARY FILES. The original EDT files were stripped of PAR and SPAR with output EDTREM and will be the PRIMARY FILES. Those two files were then merged with the PAR and SPAR only taken from the secondary files. The output of this operation was called EDT2.

##### Initial Recalibration and Fluorescence Filter

The results of 2005-16 will be applied to the primary salinity channel, subtracting 0.0047.

From section 11 we have the following equation for recalibration of DOX:

CTD-BOT = 1.4614 * DOX-CTD - 0.032

File 2005-32-recal1.ccf was prepared to apply the above corrections to the DO and primary salinity channels in the SAM and MRG files. COMPARE was rerun for dissolved oxygen and when the same points were left out of the comparison as in the first run, the average difference is -0.0005ml/l indicating that the recalibration worked properly. (See 2005-32-dox-comp2.xls.)

The edited downcast files, EDT, were recalibrated using 2005-32-recal1.ccf. (Output:COR1)

The COR1 files were clipped to 100db and set aside to be processed later for Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

The downcast files were bin-averaged using 0.5db and then thinned to usual bottle levels and the CTD DO values were compared with the upcast bottle DO values. When plotted against pressure and removing outliers identified by residuals, the following correction was derived from the trendline:

CTD DOX (Corrected) = CTD DOX + 0.0004*Pressure - 0.1019 (See 2005-32-dox-comp3.xls.)

That is similar to the fit found for 2005-25.

That correction was applied to the thinned files using calibration control file 2005-32-recal2.ccf. COMPARE was rerun and the results were satisfactory, so the correction was applied to the bin-averaged downcast files. (See 2005-32-dox-comp4.xls.)

The second recalibration will not be applied to the bottle files because it only applies to data collected while the CTD was moving.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5ml/l in the top 50m

- ±0.3ml/l from 50m to 150m

- ±0.1ml/l below 150m

##### Special Fluorometer Processing

All COR1 files were put through CLIP to produce files with data to 100db only for the use of A. Peña.

Those files were then bin-averaged (1/4db bins), put through REMOVE to remove extraneous channels, HEADEDIT to fix formats and channel names and recalibrated using 2005-32-recal2.ccf. The final files (FCTD) were saved in a separate directory. A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data was put through a median filter with fixed width 11, before bin-averaging. The recalibrated CTD bottle files (SAMCOR) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The PAR:reference data may be affected by the presence of small dots

of paint on the sensor. Painting was going on aboard the Vector during

this cruise and paint was found on the sensor afterwards, but it is

unknown when it first appeared.

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.5ml/l from   0 -  50m

•	±0.3ml/l from  50 - 150m

•	±0.1ml/l below 150m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

##### Final Bottle Files

The MRGCOR files were put through CLEAN to remove the SeaBird headers.

SORT was used to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### Final Plots

THIN and DERIVE were run to obtain values for tables, and page plots were prepared using the edited data and displaying T, S and DO. Profile plots were produced showing PAR, DO, fluorescence and transmissivity.

##### Cross-reference lists were produced for the final CTD and CHE file.

The sensor histories were updated.

REVISION April 17, 2006.

Errors were found in the CHE files – some were missing CHL data and others DO data. The DO data for casts #39, 46, 64 and 67 were missing. The OXY file for cast #39 contained data from both casts #39 and #46. An ADD file was created for each missing file and then a text editor was used to change the sample numbers and DO values to those on the rosette log sheet for casts #39 and 46. The other columns were not changed. A note of explanation was entered in the ADD file headers. For cast #67 there was an OXY file, so an ADD file was created for that. The ADD file for cast #62 contained the data for that cast plus the data for cast #64, so two separate files were created. The new ADD files plus CHL files were merged into the CHE files.

Particulars from logs

2. Log note that Pressure offset changed after cast #2. There is no cast #2 and there is no difference between the con files from the first five casts.

5. Reversal of direction during upcast to get missed bottle.

23. Bottle #9 didn’t fire.

32. Deep water inversion noted.

57. File named 58. Name changed to 57.

58. File named 58b. Name changed to 58.

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-32 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 28 November 2005 End: 2 December 2005 |  |  |  |  |  |  |
| Location: SoG/JdeF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Masson D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-32

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 01/02/05 | Factory “ |  |  |
| Conductivity | 1766 | 07/06/05 | “ |  |  |
| Secondary Temp. | 2106 | 08/07/05 | “ |  |  |
| Secondary Cond. | 1729 | 12/07/05 | “ |  |  |
| Transmissometer | 498DR | 22/05/05 | Factory |  |  |
| SBE 43 DO sensor | 0047 | 10/02/05 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| PAR | 4656 | 15/02/03 |  |  |  |
| Surface PAR | 16504 | 02/01/04 |  |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |