## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2005-33

Agency: Ocean Sciences Division

Location: Strait of Georgia, Juan de Fuca Strait

Project: Strait of Georgia

Party Chief: Ross P.

Platform: Vector

Date: November 23, 2005 – November 26, 2005

Processed by: Germaine Gatien

Date of Processing: 19 May 2006 – 24 May 2006

Number of original CTD casts: 6

Number of CTD casts processed: 6

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#498DR), an Altimeter OA-916D (#1024), a SeaBird dissolved oxygen sensor (#0047) and a Seapoint Fluorometer (#2356) with a 10X cable was mounted on the secondary pump.

# SUMMARY OF QUALITY AND CONCERNS

There was no calibration sampling during this cruise, but it occurred shortly before cruise 2005-32 in the same region, and carried most of the same equipment, so processing was based on the later cruise.

The salinity analysis for 2005-33 was considered unreliable. The results for 2005-16 and 2006-08 using the same equipment are very different, so no recalibration has been applied. This should be revisited after factory recalibration results become available.

Based on the results of cruise 2005-32 the SBE dissolved oxygen data in the CTD files should be considered

•	±0.5ml/l from   0 -  50m

•	±0.3ml/l from 50 - 150m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT. The file names were non-standard, so were changed to the format 2005-33-00**.*.

##### Preliminary Steps

The Log Book was obtained. There is no equipment list but the configuration file indicates that the equipment was the same as that carried for 2005-32 except that there was no PAR or Surface PAR on this cruise. The calibrations were checked and no errors were found. During 2005-32 there was a test that indicates that a pressure offset of 3.545 is appropriate so con file 2005-33-ctd.con was prepared with that change.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

The files were converted to CNV files

A test conversion of rosette files produced files with 1 surface bottle for each of 4 casts. The pumps were not operating when this occurred and there is no indication of any sample analysis, so no processing of bottle files will be done. It is assumed that this was done to gather water only.

A few casts were examined and all expected channels were found and looked reasonable.

The descent rate was high and quite steady.

The down and upcast temperature and conductivity channels are reasonably close, though the two channels vary more on the upcast than the downcast.

Transmissivity looked fine with good agreement between downcast and upcast.

The dissolved oxygen had the usual sort of offset.

The fluorescence looks ok with the usual spikiness.

The altimetry is often extremely noisy but looks reliable near the bottom.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

The CELLTM choice for 2005-32 will be applied: (0.03, 9) for the primary conductivity and (0.02, 7) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

One cast was examined to see if the differences are reasonably close to those seen during 2005-32 and they are. The temperature differences are larger than usual and during 2005-32 this was thought to be due to variations in alignment, rather than calibration problems or spikes in one or the other.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check showed that the project name was wrong and errors were found in two station names in the header summary. Those errors were corrected in the CLN files and the checks rerun without further errors being found. Some of the header times differ by up to 40 minutes from those of the log book, but that probably reflects the fact that other activities took place at those sites besides the CTD casts. One time is missing from the log book. The times in the log book are in PST; in the headers they are in UTC. There were no problems with the GPS time during 2005-32

The average surface pressure is 0.33db which is a little low and there are some records with negative pressures. However, the conductivity is ~0 for the low values and the pumps were not running. The pressure appears to be correct

The altimeter values in the header were checked against plots of the altimetry and the algorithm was found to have worked well.

##### Other Comparisons

Previous experience with these sensors –

- The primary conductivity sensors were used for 2005-16 when the bottle comparison had a lot of scatter and showed the CTD salinity to be low by about 0.0047.  The secondary sensors were used for 2005-21, -22 and -25 when they produced salinity high by 0.0027, 0.0021 and 0.0032.  During 2005-32 the salinity analysis was considered unreliable.

- During the most recent uses of this DO sensor the results were:

CTD-BOT = 1.3423 DOX-CTD - 0.143   (2005-12) WCVI

CTD-BOT = 1.574   DOX-CTD - 0.154   (2005-30) Broughton

CTD-BOT = 1.4614 DOX-CTD - 0.032   (2005-32) SoG/JdeF

- The fluorescence was found to be almost twice as high as the titrated chlorophyll, which was unusual given the CHL values were low and the fit is usually fairly close for low CHL.

Historic ranges – All data fell within the local climatology limits.

##### 13. SHIFT

Fluorescence

As usual all fluorescence data were shifted by +24records. One cast was plotted before and after this step and the alignment is much improved by the shift. (Output: SHFFL)

Dissolved Oxygen

During 2005-32 a setting of +130 records was found best for the alignment of DO with temperature. One cast was tested for this cruise and the same setting proved best. So SHIFT was running applying +130 records to the DO channel for all casts.

Conductivity

For 2005-32 the best setting proved to be -0.4 for both sensors. Since the primary sensors were selected for 2005-32, SHIFT was run only on that channel using -0.4. (Output: *.SHFC0).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salinity

Maximum Surface Pressure (relative): 10.00	  Minimum Salinity: 5

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None

##### DETAILED EDITING

Page plots were produced and used to guide the editing. On-screen plots of descent rate and pump status were also used. There were many unstable features, but these are likely to reflect actual conditions as there is active mixing in this region. All casts required some editing, mostly at the surface and bottom of the casts. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

For 2005-16 (July) the salinity was found to be low by 0.0047 but the comparison was considered particularly noisy. For 2006-08 (January) the salinity was found to be close enough to bottles to not bother recalibrating. This data will not be recalibrated, but when the sensors have been recalibrated that decision should be reviewed.

File 2005-33-recal1.ccf was prepared to apply the following correction to the DO channels in the EDT files. (Output:COR1)

CTD-BOT = 1.4614 * DOX-CTD - 0.032

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

Based on the results of 2005-32 file 2005-33-recal2.ccf was prepared to apply the following correction to the AVG files:

CTD DOX (Corrected) = CTD DOX + 0.0004*Pressure - 0.1019

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

Based on the results of cruise 2005-32 the SBE dissolved oxygen data

in the CTD files should be considered

•	±0.5ml/l from   0 -  50m

•	±0.3ml/l from  50 - 150m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

Header Check was rerun on the CTD files to check for errors; none were found.

##### Cross-reference lists were produced for the final CTD files.

The sensor histories were updated.

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2005-33 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 23 November 2005 End: 26 November 2005 |  |  |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Ross P. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2005-33

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 01/02/05 | Factory “ |  |  |
| Conductivity | 1766 | 07/06/05 | “ |  |  |
| Secondary Temp. | 2106 | 08/07/05 | “ |  |  |
| Secondary Cond. | 1729 | 12/07/05 | “ |  |  |
| Transmissometer | 498DR | 22/05/05 | Factory |  |  |
| SBE 43 DO sensor | 0047 | 10/02/05 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |