## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2005-35

Agency: IOS, Ocean Sciences Division, Sidney, B.C.

Project: Strait of Georgia Pollution Transport

Chief Scientist: Ross P.

Platform: VECTOR

Date:  9 September 2005 – 10 September 2005

Processed by: Germaine Gatien

Date of Processing: 19 May 2006 – 23 May 2006

Number of original CTD casts: 2

Number of casts processed:  2

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (#0334) was run with pressure sensor 290544 and SBE43 dissolved oxygen sensor (#0047).

# SUMMARY OF QUALITY AND CONCERNS

There was no calibration sampling.

There was no configuration file included with the data, but it is presumed that the same file as used for 2005-30 in October 2005 is appropriate.

There have been problems with the pump not coming on at the surface for this CTD during other cruises. For 2005-35 there does not appear to be any problem, probably because the surface conductivity was high enough to turn on the pump.

During 2005-30 when there were many bottles on a hydro wire, the salinity was found to be low by 0.001. No recalibration was applied to this data.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

The Daily Log was obtained. There is no list of equipment in the log and no configuration file. There are only 2 hex files and 2 cnv files, but the latter have odd channels, not the ones we usually archive.

During 2005-30 it was noted that the SBE internal clock is fast by about 2 minutes. This is not a significant error.

Configuration files used for 2005-27 and 2005-30 (July and October) were obtained and trial conversions done to test which, if either, is an appropriate configuration for this data. Plots were made of the CNV files produced.

Using the 2005-27 file which includes a transmissometer, the data looks reasonable except that the transmissivity signal is constant, so that instrument was presumably not mounted.

Using the 2005-30 con file, the dissolved oxygen voltage is constant, so it is assumed the DO sensor was not mounted.

The configuration file from 2005-27 was adapted by removing the transmissivity channel and the two files were then converted. It was saved as 0334CTD.con.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

The data were converted using the 0334CTD.con.

Both casts were plotted. The pressure signal has steps of about 1db; the manufacturer states the resolution is 1db.

The temperature and conductivity look normal.

The descent rate is quite steady and mostly above 0.5m/s.

The headers do not contain positions, but they are available from the log book.

4. WILDEDIT

WILDEDIT was run on all casts on pressure, temperature and conductivity channels using 2,20,25,0 for “Standard deviations for pass 1” and “Standard deviations for pass 2”, scans per block and “Keep data within this distance of the mean”.

5. ALIGN

This step was skipped. SHIFT will be used to align T and C.

6. WFILTER

Based on the results of many other cruises using this equipment, the SeaSoft routine WFILTER was run for all casts to apply a cosine filter, size 5, to the pressure, temperature and conductivity.

7. CELLTM

Based on the results of other recent uses of this equipment CELLTM was run on all casts using α = 0.03, 1/β = 9.0.

8. DERIVE

Program DERIVE was run to calculate salinity.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 25 data to IOS Headers.

The headers lack positions, station names and water depth. The time is reasonably close to the time in the log book. A text editor was used to add longitude, latitude, water depth and station name to the headers based on the Daily Log Book entries.

10. Checking Headers

Header Check turned up no further errors. The average surface value was 0.3db with very low salinity values that look reasonable for that level.

The track plot looks fine and was added to the bottom of this report.

11. SHIFT

Based on the results of other cruises using this equipment SHIFT was run on all casts advancing the conductivity by +0.7 records. The data was plotted before and after on a T-S surface and the shifted data looked worse than the original data. A few different values were tested, and none looked better than the original, so this step was skipped.

- DELETE

The files were put through DELETE using the following parameters:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00    Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 7 points) was deleted

Sample interval taken from the header.

Pressure was not filtered.

COMMENTS ON WARNINGS: There were no warnings.

- DETAILED EDITING

Page plots were produced and examined for spikes and instabilities and used to guide the use of CTDEDIT.

Both casts were examined in CTDEDIT and both required some editing. Note was made of the editing details in the files.

- BIN AVERAGE

The following Bin Average values were applied to the ED1 files:

Bin channel = pressure 		Averaging interval = 1db

Minimum bin value =   .000	Average value will be used

Interpolated values are NOT used for empty bins.

- Intercomparisons

Sensor History

The pressure sensor has been used for 5 other cruises and no offset was applied to any of them.

The conductivity sensor has been used on 4 other cruises since last calibration, 3 of which had some calibration sampling. It was found to be low by 0.004 during 2005-17 but there was only shallow sampling. During 2005-30 there were many bottles on a hydro wire and the salinity was found to be low by 0.001. During 2005-38, with limited bottles, salinity was considered to be within 0.002.

Historic ranges –There is no local climatology available for this area.

- Recalibration

No calibration was applied.

- REMOVE

The following channels were removed from all casts: Scan_number; Conductivity:Primary, Descent Rate and Flag.

- HEADER EDIT

Header Edit was used to fix formats and channel names. Standards Check was run and Header Edit rerun until all problems were resolved.

- Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

c.) The conductivity and pressure sensor history files were updated.

## Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2005-35

Dates: Start: 9 September 2005 End: 9 September 2005

Location: Strait of Georgia

Vessel: Vector

Party Chief: Ross P.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334		Cruise ID#:		2005-35

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 290544 | 13May05 | Factory |  |  |
| Temperature | 4054 | 14Oct04 | Factory |  |  |
| Conductivity | 2754 | 15Oct04 | Factory |  |  |