## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 29 Nov 2018 | Corrected DMS data in file 2006-08-0037.che and in DMS summary. |
| 22 Aug 2016 | Found missing CHE file (event 23) and updated flags to numerical format |
| 24-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. |
| 25-Sep-2012 | File 2006-08-0056.CHE: Jim Christian requested that quality flag of d be used for Nitrate_plus_Nitrite, Silicate, and Phosphate at 199.6 dbars (sample #337). Suspect that bottle was misfired. J.L. |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 11-Apr-2010 | Added Lisa Miller’s Dissolved Inorganic Carbon and Alkalinity data to the rosette files. J.L. |

## PROCESSING NOTES

Cruise: 2006-08

Agency: Ocean Sciences Division

Location: North-East Pacific / Rivers Inlet

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: January 30, 2006 – February 19, 2006

Processed by: Germaine Gatien

Date of Processing: April 11, 2006 – May 19, 2006

Number of original CTD casts: 50   	Number of CTD casts processed: 50

Number of bottle casts: 43       		Number of bottle casts processed: 43 (40 for archive)

Number of TSG files: 2			Number of TSG files processed: 2

# INSTRUMENT SUMMARY

Two CTDs were used with a mid-ship winch:

- SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#498DR), a Benthos Altimeter (#1024), a SeaBird DO sensor (#0766) and a Seapoint Fluorometer (#2356) with a 10X cable. The deck unit was a model 911 (#0508) and the logging computer was PAC02328.

- SeaBird Model SBE 911+ CTD (#0558) was mounted with a Chelsea/Seatech transmissometer (#333DR or #498DR), a SeaBird DO sensor (#0047) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0508) and the logging computer was PAC02328.

The Rosette had 24 10-litre bottles.

Two salinometers were used: model 8410 (serial number 58879) and model 8400B Autosal (serial number 68572).

# SUMMARY OF QUALITY AND CONCERNS

This was an extremely difficult cruise to process. Line P is always complex so that errors are bound to occur, but the problems were more severe than usual. There were a few errors in sample numbers in the rosette log sheets, CTD Daily log book and the analysis spreadsheets, and format problems with the dissolved oxygen and salinity analyses, and many salinity samples were missing. The missing salinity samples were randomly distributed so they were not discovered until the processing job was almost completed, with the result that the bottle preparation and calibration comparison had to be repeated.

The most serious error is that the log book and all con files indicated that Sea-Bird Dissolved Oxygen sensor #0047 was used for all casts, but it was in fact only used for 6 casts, none of which included bottles. All the casts using CTD #0443 were calibrated incorrectly initially. Fortunately there were many dissolved oxygen samples from the rosette, so recalibration will probably have done an adequate job of correcting the initial faulty assumption. The 6 non-rosette casts were calibrated initially using the correct co-efficients, but an inappropriate recalibration was applied. There are no bottles available for those casts, so final correction was based on the comparison of 2 casts at the same site but using the 2 different sensors.

Another fact that was unknown to the processor until the end of the job was that a protocol was used for the SBE dissolved oxygen sensor that had never been followed before at IOS; a plug was occasionally attached to “starve” the sensor of oxygen. This may explain why the offset in the fit of bottles to sensor varied with time, though the slope did not. Different re-calibrations were used for different parts of the cruise to allow for this variability. Considerable time was spent investigating the unusual calibration; knowing about the switch in sensors and the “starving” technique might have saved some time.

The titrated oxygen files had format errors and incorrect entry of pad values.

The salinity bottle data was provided in 11 spreadsheet files in order of analysis without event numbers.

There was an error in the pressure sensor configuration used at sea for CTD #0585; an offset of -0.6db was entered whereas +0.4db would be appropriate. The operator would think the CTD is a metre shallower than it really is. If it is important to get close to the surface, then con files need to have the correct pressure co-efficients. As they get older there is drift in the offset of the pressure sensors, so this will become a more serious issue with time.

Two salinometers were used and 19 samples were run on both. The Portasal gave values higher than the Autosal by from 0 to 0.004, with an average difference of 0.0021. For one cast 24 bottles were fired at 2000db; when the salinity samples were analyzed with the Autosal the salinity varied by 0.0045 (excluding one outlier). The CTD values from the rosette file varied by 0.0023, suggesting that the Autosal scatter is ~ ±0.001. There was no indication of problems with Niskin bottles.

The dissolved oxygen data in the CTD files using CTD #0443 is considered:

•	±0.4ml/l from 0 - 200m

•	±0.2ml/l from 200 – 1000m

•	±0.1ml/l from 1000 - 1500m

•	This instrument is considered unreliable below 1300db; for this mission bottle comparisons suggest that the Oxygen:Dissolved:SBE channel is low by from 0.1 to 0.2ml/l between 1500 and 4000db.

There are no calibration samples for the DO sensor mounted on CTD #0585, so it was calibrated by comparing 2 casts at the same site using the 2 different sensors. No error estimate was attempted.

The transmissometer data was removed from several casts because the data was bad or the instrument had been removed.

The CHE files for casts 47, 55 and 63 were prepared for the use of UBC researchers. They contain the usual CTD data from the time of rosette firing but no bottle sampling data; it would make processing easier if sample #s were assigned to all bottles even if they are not sampled. There are also some lines in the CHE files for cast #1 that had no corresponding IOS sampling, hence no sample #s.

The Thermosalinograph fluorescence is much higher than the titrated chlorophyll samples and about twice as high as the CTD fluorescence. No recalibration was done to that channel.

There was no record in the log book of which fluorometer used on the TSG, it was presumed to be the one usually mounted on the Tully TSG.

The report from the Chief Scientist was very helpful in determining where there were equipment changes and header or file name errors.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained along with rosette log sheets and a report by the chief scientist about activities and problems encountered during the cruise. This included a request to create rosette files for some casts run without any IOS sampling, since UBC researchers did do sampling.

There was no record in the log book of the fluorometer used on the TSG or of the salinometer used. The fluorometer ID was not in the configuration file either. The calibration given corresponds to that of fluorometer WS3S-713P, so that was assumed to be the serial number of the sensor. The salinometer information was available from the salinity analysis sheets.

Dissolved oxygen data were available with flag channel and comments, but the format of the comments was wrong. It was necessary to go through each ADD file to ensure the format was correct.

The Autosal bottle salinity data were obtained in spreadsheet form with a flag channel and comments.

The Portasal bottle salinity data were in 10 spreadsheets. Most have a quality channel, but since most have no headers it was not completely obvious if there are comments or not. There is at least one flag with no comment. This will be investigated later.

The nutrients were available in spreadsheet format with flags and comments.

The cruise summary sheet was completed. The histories of the conductivity, dissolved oxygen and pressure sensors were obtained as well as that of the thermosalinograph conductivity and temperature.

Two CTDs were used during this cruise; the report from the chief scientist makes it clear what was used when. Three CON files were prepared as follows:

2006-08-ctd1.con – CTD #0443 with transmissometer #498DR (casts 1-39, 43-47, 50-52, 54-63**)

2006-08-ctd2.con – CTD #0585 with transmissometer #333DR (casts 40-42*)

2006-08-ctd3.con – CTD #0585 with transmissometer #498DR (casts 50, 52, 53*)

*For casts #42 and 53 no transmissometer was actually mounted, but it was included in the CON file.

**The DO sensor was left off for the first and last casts (to avoid the effects of anoxia in Saanich Inlet), but the sensor was included in the con file.

The calibration constants were checked for all instruments and the only adjustments made were to the pressure calibration of CTD 0585. The pressure offset was set to -0.6db whereas +0.4db has been used in all recent uses of that sensor, so +0.4db was entered for this cruise as well.

A comparison was made between the two CTDs by doing a repeat cast at #42 and #43.

##### Conversion of Raw Data

All data were converted using the con files given above. Files named 2006-08-0004.* were renamed 2006-08-0005.*. A few casts were examined and all expected channels are present. Both T and C channels show fine-scale noise even when the descent rate is steady and the upcast differs significantly from the downcast. Rosette files were converted using a start time of -2s and duration of 5s.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and BOT was used for the extensions. All BOT files were plotted and noise was noted in casts 2, 7, 28, 43, 46 and 58. CTDEDIT was used to remove some records from cast #2 and to clean both salinity channels mostly at the surface. The edited files were copied to *.BOT. Comments describing the editing were entered in the relevant headers. Based on notes from the Chief Scientist, a text editor was used to correct or add station names to a few casts.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### The station names were corrected for a few casts.

##### CELLTM

There were two CTDs used. Tests were run on two casts for each CTD. It is best to do the tests on casts with fairly steady descent and ascent rates using a variety of settings for CELLTM. However, the descent rate is extremely noisy for all the casts using CTD 0443. The best results proved to be (0.03,7) for all casts though there was little difference between a number of the choices.

CELLTM was run using (0.03, 7) for the primary and secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

There were many deep casts but the descent rate was very noisy for those. Two casts were plotted for CTD #0585 and three for CTD #0443 to check for agreement between the pairs of T and C sensors. The following values are rough estimates from downcast data:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| CTD0443 in rosette |  |  |  |  |  |
| 29 | 2000 | -0.017 | +0.00004 | +0.0024 | Very noisy |
| 47 | 2000 | -0.018 | +0.00005 | +0.0025 | Very noisy |
| 52 | 1000 | -0.015 | +0.00002 | +0.0017 | Extremely noisy |
| 52 | 2000 | -0.018 | +0.00004 | +0.0024 | Noisy but high |
| 52 | 3000 | -0.017 | +0.00005 | +0.0023 | Very noisy |
| 52 | 4000 | -0.017 | +0.00005 | +0.0024 | Extremely noisy |
| CTD0585 stand-alone |  |  |  |  |  |
| 49 | 2000 | +0.0007 | -0.0003 | -0.0045 | Extremely noisy |
| 53 | 2000 | +0.0003 | -0.00035 | -0.0043 | Extremely noisy |

The temperature difference for CTD 0443 is unusually large. There is no significant temporal or pressure dependence. The casts using CTD 0585 were so noisy that it is difficult to pick out a value.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and a few errors were found. The Sea-Bird headers were scrambled for cast #3. That was corrected and conversion and CLEAN rerun. The speed check showed an average of greater than 10 knots between casts 58 and 59; this looks reasonable because there was a long steam with no intervening activities that required a stop.

At this point in the processing job, it was noted that the last cast was named #63 in the log book, but the file name had #62, so the file name and event number was fixed in both the CTD and bottle files.

The track plots (using event #s and station names) were produced and look reasonable; they were added to the end of this report.

The header summary was run and station names were checked against the log as well as a random sampling of positions and times to ensure that the GPS was working well. No further errors were found.

For CTD #0443 the average surface pressure was 2.4db which is about what we expect for the Tully. One cast had a surface value of 0.7db so that was examined closely and the salinity values support the idea that the CTD was very close to the surface; that cast was in protected waters. The pressure sensor for this CTD is drifting. During 2005-32 Doug Anderson determined that an offset of 3.5457db should be used. Over the past year values of 3.0457 was used for most cruises, and 3.4457db was used for 2005-16; 3.5457 was applied to this cruise. The average surface pressure is 3.8db for CTD #0585 which is a little deeper than expected, but it was used during the roughest weather and two of the values are around 2.2db which looks reasonable.

The mixed layer depth calculation shows that near-surface samples will be useful for salinity calibration for most of the offshore casts, but not for most of the early casts.

The altimeter readings from the headers were exported to a spreadsheet and a few casts were checked to see that the values were reasonable. No values were found that were clearly bad, but the CTD was sometimes bobbing about at the bottom so that it is impossible to make a very reliable estimate. The altimeter has done as well as possible under the circumstances.

##### BOTTLE FILE PREPARATION

NOTE: More salinity samples analyzed on the Portasal were obtained after the bottle file preparation was completed, so it had to be redone. This did not affect the comparison with the Autosal samples. A run of COMPARE showed no significant change in results.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. There were three casts with bottles done for UBC; there are no sample numbers for those, so pad values were entered. There is a discrepancy between the log notes and rosette notes as to how many bottles were fired for a few casts, and in at least one case the same sample number is recorded twice on the rosette sheets. The Daily Log book was taken as correct in most cases. Sample numbers were added to the BOT files (output: SAM) and bin-averaged (SAMAVG.)

- The bottle salinity data were provided in 11 spreadsheets with flag and comment channels added. The preparation was complex because there were two different salinometers and both loop and rosette samples. The spreadsheets were edited to correct channel names and remove some information that is not needed and to separate CTD samples from loop samples. The CTD and loop files were then combined and saved as 2006-08-ctd-all-sal.csv and 2006-08-loop-all-sal.csv. There are some duplicates with analysis run on both the Autosal and Portasal. The question arises as to which is better, and what, if any, offset should be applied to the other.

- There were 19 samples that were analyzed on both the Autosal and the Portasal. A study was made of the differences between them. On average, the Portasal was high by 0.0015 and when one obvious outlier was excluded, by 0.0021. There was a lot of scatter with values from 0 to 0.0044 (excluding the outlier.) A plot of differences against salinity value looks quite flat. The Portasal analyses were done on 5 different days between Feb. 11 and Feb 15 while the Autosal analyses all were done on Feb. 22. The differences were plotted against the time between the two analyses and there is some suggestion that it increased as the elapsed time increased. I would expect that there might be some evaporation occurring which would lead to higher salinity from the Autosal, but what we are seeing is the opposite; varying ambient temperature might account for some of this trend. (The ship went quite far north late in the cruise and the analyst noted a change in temperature in the lab where the Portasal analysis was done.) The differences versus sample number show a similar relationship, but the sample number is closely related to the elapsed time, so this is not surprise. Differences versus nominal pressure show no hint of a trend.

- A spreadsheet was prepared in which the Autosal values were chosen when there was a choice and the Portasal values were adjusted by subtracting 0.0021 from each value. These were then converted to SAL files.

- The dissolved oxygen data were provided in ADD files with flag channel included, but the flags were in the wrong column and pad values were entered incorrectly. Errors were corrected.

- The nutrient data was received in spreadsheet format QF2006-08nuts.xls. That file includes a detailed discussion of errors. The spreadsheet was edited (changing headers and removing unnecessary columns) and saved as 2006-08-nuts.csv and converted to NUT files which were then re-ordered on sample-number. Sample #360 from cast #58 was renamed #9360 since #360 was used for cast #59 as well.

- The chlorophyll data was received in spreadsheet format. The CTD data was on a single sheet and was edited and saved as 2006-08-CTD-chl.csv. That file was converted to CHL files. Two errors were found in the data.

Sample #360 was used twice, for casts #58 and 59. The rosette sheet shows no sampling from bottle #14, cast #58 (sample #360), but Janet Barwell-Clarke says that chlorophyll was taken because they could not get the intended sample from bottle #13 (sample #359). There was salinity sampling from cast #59 which is identified as sample #360. So the sample # for bottle #14 from cast #58 will be renamed as #9360 to avoid confusion.

Sample #228 is said to be from bottle #14 of cast #37, but there was no such bottle # or sample # for that cast. There was a sample #228 for bottle #23 of cast #39. However, there is no indication on the rosette sheet that there was any CHL sampling during that cast. Janet Barwell-Clarke confirms that the value comes from cast #39.

The ADDSAMP file was converted to CST files to form a framework to which the various chemistry files will be merged. The SAL, CHL, ADD and NUTS files were merged with the CST files in four steps.

CLEAN was run to remove all remarks, but leave comments. (Output: MRGCLN)

Those files were then merged with the SAMAVG files. (Output: MRG)

##### 11. COMPARE

Salinity

NOTE: Some salinity values were found after this work was done. A quick comparison was made with the full suite of bottle data, but no significant difference was noted. (See 2006-08-sal-comp-full.xls.)

COMPARE was run. The analysis is expected to be complex for two reasons. The early part of the cruise sampled shallow calm waters and the latter half sampled offshore casts in extremely difficult conditions. So the area from which we usually expect the best comparisons because of the deep sampling may not be as reliable as usual. The other complication is that two salinometers were used, one of them only for data from the offshore casts. As explained in the previous section the Autosal values were assumed to be correct and the Portasal values adjusted to match them.

There were 7 severe outliers (samples 10, 69, 70, 71, 83, 174 and 374) but only 2 of those were below 10m. Most of the scatter is confined to the top 200m. Many different plots were produced to study the calibration and the reliability of the analyses and the Niskin bottles:

- When differences >0.02 were excluded the primary differences showed very little pressure dependence and the secondary was fairly flat. Both channels were high on average, the primary by 0.0008 and the secondary by 0.0015.

- When differences were plotted against file pair number the slope was similar for both pairs, suggesting that the temporal variations reflect either the depth of sampling or the state of the sea.

- Using only Autosal samples for bottles below 200m both channels had salinity that was fairly flat with pressure and the primary was high by an average of 0.0006 and the secondary high by 0.0046.

- Normally we would trust the offshore results more, but in this case it is not at all clear. To reduce the effect of CTD motion, points below 200m were chosen, excluding the severe outliers, and then standard deviation in the pressure was used to exclude more data; only 13 points were left. The primary salinity was then found to be high by an average of 0.0013 and the secondary high by 0.0035.

- While there was a significant salinity gradient near the surface early in the cruise, later there were many casts with very well-mixed surface waters. When only shallow samples from well-mixed surface waters were used (and many outliers removed) the primary salinity was low by an average of 0.0004.

- 24 samples from 2000m for cast #59 were analyzed in order of sample number on the Autosal. This provides a measure of the variability to be expected from this salinometer. CTD pimrary salinity in the rosette files varied from 34.5768 to 34.5791. The Autosal results vary from 34.5751 to 34.5796 when 1 outlier is excluded and the trendline suggests a drift downwards, though there is a lot of scatter. Looking at the CTD data the results are flat when averaged over the 5s window around firing time, but the minimum values for each bottle firing do appear to drift downwards towards the end. The pressure is varying notably towards the end with an 8db range during the last 4 bottles. The Autosal variability is twice that of the CTD using unaveraged CTD data.

- During cast #59 the value from Niskin #15 was an extreme outlier with the bottle value low by 0.26. For Niskin bottles #3, 11 and 17 the bottles were also low by about 0.0025 to 0.003. The CTD salinity shows little variability for all four. This may indicate that there was leakage into the Niskin bottles, or there may have been problems in sample handling or analysis. If those values are excluded the average difference indicates that the CTD is high by 0.0002.

- COMPARE was rerun using Niskin bottle # as the reference to see if there is any indication of systematic problems with bottles. There are many outliers but no evidence of problems with specific bottles, except that, as expected, there is more scatter in the first few bottles since they are sampled more often.

Most of the outliers in COMPARE were from near the surface where high gradients can probably account for the differences between the bottles and the CTD. The deeper outliers were investigated, and while the standard deviation in the pressure was larger than usual, the variations in CTD salinity were much too small to account for the bottle values, especially for the deepest outlier. The following significant outliers were flagged and a note of explanation put in the headers:

- Cast #32 at 800db, sample #174: flagged “d”. Nutrients odd as well. Looks like bottle didn’t close until 600db.

- Cast #59 at 2000db, sample #374: flagged “d”.

A comparison also had to be made between the two CTDs used. There was no bottle sampling for system CTD 0585, but there was a repeat cast using the two systems. Unfortunately, both were in heavy seas but we would expect the average differences to be useful for calibration. COMPARE can be used for this task only after thinning of the data and this is probably best left until after editing and metre-averaging, so this will be done later.

Summary of salinity comparisons:

The Portasal produced salinity that is higher than the Autosal by about 0.002.

The Autosal shows variability that after allowing for real changes is on the order of ±0.001.

The primary CTD salinity is closest to the bottles.

If the Autosal is reliable then the primary is somewhere between being low by 0.0004 to being high by 0.0013 depending on which data is selected and the secondary high by from 0.0015 to 0.0046.

If the Autosal is not the more reliable salinometer then the primary is high by ~0.003 and the secondary by at least 0.004 and possibly as much as 0.007.

Dissolved Oxygen (Note: This part of the report was written before discovery of an error in sensor serial number. See section 18 for details.)

COMPARE was run using pressure as the reference variable. (A second run was done later using sample number but there was no significant difference.) The best fit was with the differences between bottles and CTD versus CTD DOX. When 2 outliers and records for pressure >1500db were excluded the fit was:

CTD-BOT = 0.7983 DOX-CTD + 0.268

This result is very different from that during all previous uses of this sensor. For example during 2005-32 the fit was

CTD-BOT = 1.4614 DOX-CTD - 0.032

and during 2005-30 the fit was:

CTD-BOT = 1.5740 DOX-CTD - 0.154

The CTD DO values are higher than those of the bottles, most notably at the surface. The value used for shifting the DO channel was also different from previous uses.

The ratio between DO from the CTD and BOTTLE was calculated for the data used in the fit above. That ratio was plotted against a variety of parameters. The ratio of the two DO values is quite steady when DO values are high, but when DO<1 the ratio becomes low, but this does not look significant because the offset of 0.268 is much more significant when DO values are low. The trendline for a plot of DOX_BOT versus DOX_CTD changed very little when only DO_BOT > 1.5ml/l was used or DO_BOT < 1.5ml/l. While it looks like there might be some time-dependence this is probably due to the fact that the lowest DO values were late in the cruise.

Possible reasons for this change in calibration include:

- Wrong calibration used – a check was made and there was no change since the last usage and no error in the configuration file. The only other possibility is that the wrong sensor # was recorded.

- Problems with titrations – The bottle files look reasonable and there is no known reason to doubt them. The titrated DO was plotted against salinity and the only problems are with a few samples that were outliers in COMPARE.

- Problems in rosette file creation – The rosette files agree with the CTD files.

- Damage to sensor or large drift in calibration – The sensor will be sent for recalibration soon, but no information is available at this time.

- Unusual conditions such as very deep mixed layer – The mixed layer was very deep which might conceivably affect how the sensor behaves.

- Change in usage to avoid sampling anoxic waters - For this cruise the DO sensor was removed for casts in Saanich Inlet to avoid the problems that have occurred in the past with calibration drift after anoxic sampling. This could be a factor although there was no anoxic sampling during 2005-32 to the best of my knowledge.

- Sampling of low DO (but not anoxic) waters - the DO was very low at depth in the offshore waters. Plotting of a selection of individual casts in COMPARE suggests that the offset may start drifting from cast #26 (Station P4) onwards though the slope does not. Since each cast has a different range of values this may not be significant. Moreover, up to P4 the casts are all coastal & similar – after that the casts are well offshore and there are large separations.

Whatever the cause, the fit against bottles is reasonably tight and we can do a recalibration.

Flags were entered for the following titrated DO samples:

Sample #174 Cast 32 at 800m; flagged “d”, all bottle samples indicate misfire.

Sample #337 Cast 56 at 200m; flagged “c” by analyst, changed to “d” based on COMPARE.

Nutrients

The nutrients for sample #174 were flagged “d” since they also show evidence of the bottle firing at 600db instead of at 800db as was found for DO and salinity.

Fluorescence versus titrated chlorophyll

COMPARE was run using sample number as the reference variable. CHL was plotted versus FL. The following trendline was found when it was forced to go through the origin:

Titrated Chlorophyll = 0.3321 * Fluorescence.

##### 12. SHIFT

Dissolved Oxygen

Tests were run on 2 casts using shifts of +100 to 150. The best choice was judged by matching features in upcast and downcast and picking the shift that made the vertical offset for DO the same as for temperature. For cast #30 +100 looked best, but +120 looked slightly better for cast #53. This is a lower setting than has been used in recent uses of this instrument. A third cast was selected that had a steady descent rate to see if that made a difference. For that cast +100 looked best in the top 100m and +120 at greater depths. An intermediate value, +110 records, was used for all casts. (Output: *.SHFO)

Fluorescence

A shift of +24 records (1s) was applied as has been done for all recent cruises. A few casts were checked before and after this step and the results were satisfactory, with the resulting offset between downcast and upcast data looking closer to that seen in the temperature signal. (Output: SHFFL)

Conductivity

Tests were run using a variety of settings and the best results vary from cast to cast and depth to depth. The best overall results were with the following choices:

CTD #0443   Primary -0.2	Secondary	-0.2

CTD #0585   Primary -0.5	Secondary 	-0.2

All casts were put through SHIFT using those settings. (Output *.SHFC0 and SHFC1).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              		Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for 2 casts. For cast #44 they pertained to the upcast only so are not significant. For cast #41 the warnings are in areas of very noisy descent rate. DELETE appears to have worked well, with no large jumps in pressure in the output file.

At this point a preliminary comparison was made between casts #42 and 43 to decide which sensors to use for CTD #0585. The DEL files were bin-averaged and plotted together on a T-S surface. The two primary channels look close while the secondary channels differ considerably along σt-surfaces. So the primary will be selected for both CTDs. After editing a more thorough comparison should be done to determine whether recalibration is needed for CTD #0585.

##### DETAILED EDITING

The primary sensors were chosen for the archive. Page plots were produced using (T0, S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing. Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Comparison of two CTDs – Casts #42 and 43 at P15 were compared, the first was run using CTD #0585 and the second with #0443. The EDT files were bin-averaged and then thinned. COMPARE was run. The differences showed a lot of scatter, but since the scatter looks the same for primary and secondary sensors this is not likely a measure of sensor problems, but rather noise in the descent rate. Outliers were removed until the differences were flat with pressure. The shallower of the two casts went to 537db. For both sensor pairs the data from cast #42 produced salinity lower than that of cast #43. The primary was low by an average of 0.0008 whereas the secondary was low by ~0.0065. The CTD used for cast #42 was system #0585. If the primary salinity from CTD 0443 is high by ~0.001 as is suggested in section 11, then that of CTD 0585 is very close to being correct. The secondary salinity for cast #42 appears to be low by something like 0.003. This comparison is very rough. (See 2006-08-sal-comp-2CTDs.xls.)

Previous experience with these sensors –

CTD #0443

- The primary sensors were used during 2005-16 when the salinity was low by about 0.005, but the comparison was very noisy.

- The secondary sensors were used during 2005-21, -22, -25 & -23 when salinity was found to be high by 0.003, 0.002, 0.003 and 0.0001, respectively. The latter two comparisons had few bottles or a lot of scatter.

CTD #0585

- The primary sensors were used for many cruises during 2005 and salinity was found to be within 0.0012 except for one cruise when it was low by about 0.003. There was a lot of scatter for that comparison. The most recent comparison showed the salinity to be high by 0.0012.

- The secondary sensors were used during 2006-08 when the salinity was found to be high by 0.002 but there was a lot of scatter.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The only excursion from the local historic climatology was for cast #5 in Queen Charlotte Sound, where the temperature data were above the maxima between 110 and 150db. All the data from the Line P section were comfortably within the historic ranges.

##### Initial Recalibration

The comparison with bottles indicates that the primary salinity from CTD #0443 is high by ~0.001 if we take the Autosal analysis as being more reliable than the Portasal. If the Portasal is correct, then the salinity is high by at least 0.003 and possibly as much as 0.007. No recalibration will be applied. However, when these sensors are recalibrated this should be revisited. For CTD #0585 primary salinity we have only the one intercalibration which suggests that it is very close to the bottles. No recalibration is justified, but a check of the post-cruise calibration is advised when it becomes available.

An initial recalibration was run on the MRG and SAM files using file 2006-08-recal1.ccf to apply the following correction to the dissolved oxygen: (MRGCOR1 and SAMCOR1)

CTD-BOT (Corrected) = 0.7903 DOX-CTD + 0.268

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the calibration was applied to the EDT files as well. (See 2006-08-dox-comp2.xls.)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Cast #19 had some unstable features, but an examination of the data and the location suggests that the instability could be real, and even if it isn’t it is not at all clear which would be the data that should be removed. No further editing was not to T-S.

DO, Transmissivity and Fluorescence profiles were plotted on-screen. The only problems noted were in transmissivity. Casts #41, 42 and 53 had no signal in TR, so that channel will be removed later. Casts 20 and 39 had a few spikes, but no editing was done as TR is usually left unedited unless there is obvious instrumental failure. TR for cast #40 looked unbelievable and that channel should be removed later.

##### Final Calibration of DO

SHIFT addresses errors due to transit time and the comparison with titrated samples addresses the drift in the instrument’s calibration. But there remains an error due to poor time response of the sensor. To at least partly address this issue we compare downcast data to the titrated samples. COMPARE was run using thinned downcast CTD data (after initial recalibration and bin-averaging) and the titrated DO values. When outliers and deep values were excluded, the plot of differences against CTD DO looks very noisy with the CTD reading high by an average of 0.04ml/l. The source of the variability was sought in a series of plots.

When DO was plotted against pressure the trendline is flat until about 600m and below that it drops off to lower values. The plot against file pair number shows a big change after cast #32. Casts #39 and 44 look similar and then the differences drop to negative values. As mentioned in section 11, the slope of the calibration did not appear to change through the cruise but the offset did. The reason may be that the CTD was exposed to low DO values only late in the cruise and that this is causing a change in the offset.

Plots were made of the DO values at a number of casts. Cast #2 has DO values ~1ml/l but then all DO values are >2 until cast #23. From that cast onwards there was frequent sampling with DO<1. The lowest value was at cast #26. Cast #29 was the first to sample below 2000m and cast #39 the first to go below 3000m. So the changing offset could be associated with accumulated effects of sampling low DO water or exposure to high pressures. (See 2006-08-dox-comp3.xls)

File 2006-08-recal2.ccf was prepared to apply an offset as follows:

Casts  1-35:	+0.075

Casts 37-47:	+0.041

Casts 48-63:	-0.071

Note that the sensor was not mounted for casts #1 and #63, but the casts were put through the calibration step anyway to simplify the processing stream.

COMPARE was rerun to check that the calibration was run correctly and the average difference between CTD and bottles was then 0.003ml/l. (See 2006-08-comp4.xls.) The AVG files were then recalibrated. This correction was not applied to the bottle files since it addresses an error that does not affect the CTD data while it is stopped. (Output: COR2)

As a final check of the DO calibration, the DO saturation was calculated and plotted near the surface. It was in the range of 90 to 100% except for the casts using CTD #0585. Those values are typical for winter. Casts #41-42, 48-49 and 53 have low saturations, and when the results for casts #42 and 43 were plotted together it is clear that the results for CTD #0585 are not correct. There are no bottle data available for that sensor. The calibration used in the configuration file was correct.

A comparison was made between casts #42 and 43 since they used the two different CTDs at the same site. The data on a plot of DO from one cast against the other fell into a straight line with the following relationship:

DOX_CTD cast 42 =  0.5336 * DOX_CTD cast 43 + 0.2257

DOX_CTD cast 43 =  1.8737 * DOX_CTD cast 42 - 0.4219

The relationship found in section 11 can be expressed as:

CTD_BOT = 0.7983 DOX_CTD_Cast 43 + 0.268

From that we can deduce that

CTD_BOT = 1.4958 DOX_CTD_Cast 42 + 0.0696

This looks much more like the results of other cruises such as the following from 2005-32:

CTD-BOT = 1.4614 DOX-CTD - 0.032

This shows that the problems noted earlier are likely not due to sensor trouble, but rather the misidentification of sensors. Despite what is said in the log book and in all configuration files, the DO sensor must have been different on the two different CTD systems. Doug Anderson confirms that this was likely the case. He believes that sensor #0047 was only used on the 6 non-rosette casts (CTD #0585), and that sensor #0766 was used for all rosette casts (CTD#0443). To check this, cast #44 was reconverted using the calibrations for DO #0766 and the results compared with bottle DO. The fit looks much more like what we have found in the past. (See new-DO-config-check.xls.)

All the rosette casts were all converted incorrectly. However, there were a lot of bottles and the comparison with the CTD DO provided a correction that should remove that error, at least in as much as the error is linear.

The 6 non-rosette casts were converted correctly but the correction based on the bottle comparison was inappropriate. However, there is no point in going back and doing it again since we need to apply a further correction anyway. Those 6 files should be recalibrated based on the cast 42-43 comparison shown above. So file 2006-08-recal3.ccf was prepared to apply the following equation to casts 40, 41, 42, 48, 49 and 53:

DOX_CTD (corrected) = 1.8737 * DOX_CTD - 0.4219

A check was made that casts #43 and 42 are close after the correction and they were.

Finally DO saturation was recalculated for all the non-rosette casts and surface values are between 97 and 103%.

Another discovery made late in processing is that there was sporadic application of a different protocol for handling the dissolved oxygen sensor. Doug Anderson reports that he applied a plug that “starves” the DO sensor of oxygen. At first he did it regularly, but later changed to doing it only during long steams and removing it several hours before use. This change was made because some negative DO values were observed, but it is now obvious that those were due to the faulty configuration file.

Using this technique consistently looks like a promising way to get steadier calibrations and to enable sampling in anoxic waters without calibration drift.

The irregular application of the plug may be the cause of the drifting calibration described above.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE (volts/kg), Status:Pump, Descent_Rate and Flag.

The transmissivity channel was removed only from casts #40, 41, 42 and 53 where the data was either absent or bad. The Oxygen:Dissolved:SBE channel was removed from cast #1 and #63 since the sensor was not mounted for those casts.

CHANGE UNITS was used to add a second DO channel in umol/kg.

REORDER was run to get the two DO channels together. (Output: REO)

HEADER EDIT was used to fix formats and channel names, to fix the DO sensor serial number and to add the following comments to casts using CTD #0443:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The DO sensor was originally identified incorrectly and the wrong

calibration was used. There were many bottles, so the calibration which

was based on comparison to titrated values is believed to have corrected

this error.

The dissolved oxygen data in the CTD files should be considered

•	±0.4ml/l from 0 - 200m

•	±0.2ml/l from 200 – 1000m

•	±0.1ml/l from 1000 - 1500m

•	This instrument is considered unreliable below 1500db; for this

mission bottle comparisons suggest that the Oxygen:Dissolved:SBE channel

is low by from 0.1 to 0.2ml/l between 1500 and 4000db.

HEADER EDIT was used to fix formats and channel names and to add the following comments to casts using CTD #0585:

DO sensor #0047 was used without a rosette, so there are no bottles

for comparison. Recalibration was based on a site that was sampled twice,

once with this sensor and immediately afterwards with sensor #0766 for

which there were many bottle calibration samples. It is impossible to

estimate the errors for sensor #0047 during this cruise but the results

when used during other cruises suggests errors would be on the order of:

•	±0.5ml/l from 0 - 200m

•	±0.2ml/l from 200 – 400m

•	±0.1ml/l from 400 - 1500m

•   The sensor is considered unreliable below 1500db producing values

that are generally low, typically by 0.1 to 0.3ml/l.

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The comments were adjusted for individual files to reflect the absence of transmissivity or dissolved oxygen channels.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to rearrange files with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE (volts), Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.MRGREM)

The Oxygen:Dissolved:SBE channel was removed from cast #1 and #63 since the sensor was not mounted for those casts.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods, and to fix the serial number for the DO sensor.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

CLEAN was run to remove Sea-Bird headers and “Comments from secondary file” from the header with the output extension.

The CHE files for casts 47, 55 and 63 were set aside for the use of UBC researchers; there was no IOS sampling so they are not to be put in DATA_LIB. Cast #1 had some IOS sampling and some bottles fired only for the use of UBC researchers. Two versions of that file were created, a full form for the UBC group, and one edited to leave only those bottles sampled by IOS. The latter was run through CLEAN to reset the headers and it is intended that it be placed in the DATA_LIB archive. (Output: CHE)

##### 21. Thermosalinograph Data

a.) Checking calibrations

There were 2 files containing TSG data. The con files had an error in the temperature calibration and the fluorometer serial number was missing; the corrected file was saved as 2006-08-TSG.con.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPoly0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

A time-series plot was produced and spikes were noted in all channels, but especially in the salinity channel. The flow rate was fine and the temperature differences reflect the temperature variability.

An initial track plot was produced and looks ok; there is an odd jog in the middle of the Pacific that was probably due to weather conditions. The positions fit the log records before and after the jog.

CTDEDIT was used to remove many large spikes in salinity and a few large spikes in primary temperature and fluorescence; salinity was also cleaned lightly around records 40975 and 42375 where unbelievable salinity data was seen that is likely due to the mismatch of conductivity and temperature in an area of large temperature gradients. Two runs of CTDEDIT were used. Only the 2nd file needed editing.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within 1db of 4db and exported to a spreadsheet. The TSG files were opened in EXCEL and reduced to the times when CTDs were run. Those files were combined in a spreadsheet matching times. The positions were compared and were very close, with average differences for both latitude and longitude of 0.0001º and no difference was greater than 0.0004º, so the clock would appear to be accurate. This spreadsheet will also be used in the next step to compare temperature, salinity and fluorescence. (See 2006-08ctd-tsg-comp.xls)

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 When the differences between intake temperature and lab temperature were plotted there was a lot of noise. The average difference was -0.231Cº and the median value was the same. When only 19 casts with a deep mixed-layer were included, the average was -0.225Cº.   Looking at 2006-08-CTD-TSG-comp.xls, we find that during stops for CTDs, the temperature from the lab is higher than the remote temperature by an average of 0.225Cº (median value is 0.23Cº).

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. A subset of those casts was chosen as well-mixed based on a mixed-layer depth calculation. (See 2006-08-ctd-tag-comp.xls and 2006-08-ctd-tsg-wellmixed.xls.)

The TSG fluorescence is higher than that of the CTD by a factor of from 1.2 to 2.8.

The TSG salinity is lower than that of the CTD by an average of 0.105 but when 4 extreme outliers are excluded the average is 0.028. When only the well-mixed casts are used the TSG is low by 0.0237 and that comes down to 0.0218 when 1 outlier is excluded.

The TSG intake temperature is lower than the CTD temperature by <0.007Cº. This is as close as one could expect.

The lab temperature is higher than the CTD temperature by an average of 0.225 using all casts and 0.215Cº using just the well-mixed casts.

No time-dependence is noted but might well be masked by the scatter in the results.

e.) Loop Salinity comparison

There were 52 loop samples analyzed; a Portasal was used for most but a few were from the Autosal. The study made in section 10 indicates that the Portasal salinities are higher than those of the Autosal by 0.0021. Hence, all Portasal values were reduced by 0.0021. A spreadsheet (2006-08-TSG-sal-comp.xls) was prepared with the loop sample ID, the salinity from the analysis spreadsheet and the date and time of collection as noted in the log book. There was no time given for one sample.

The salinity data in the TSG file were averaged and the standard deviation found over 2 minutes (5 samples); the average value and standard deviation were recorded for the times of the loop samples. These were compared and the TSG salinity is low by 0.035 when all samples are included. When 10 outliers are removed (8 because of high standard deviation and 2 with very large salinity differences) it is low by an average of 0.025 (median of 0.018) and when only the 12 samples with the lowest standard deviations are included it is low by 0.019 (median 0.017). (See file 2006-08-TSG-loop-comp.xls)

NOTE: 3 more loop samples were found after this analysis; they were from the first file. The near-surface salinity gradient was quite high in that area and there was a lot of variability in the TSG data, but they fall roughly in line with other data. For the one sample with the lowest standard deviation the TSG was low by 0.015.

f.) Loop chlorophyll comparison

Chlorophyll titration data from the loop was combined with CTD data in a spreadsheet; the loop data was extracted, edited and saved as 2006-08-loop-chl.csv. A few values were compared with the TSG data. The TSG fluorescence was about 6 times higher than the titrated chlorophyll.

g.) Calibration history

The TSG was recalibrated in October 2005. The drift found at that time was much less than the errors found by comparison with loop samples and CTD for previous cruises; salinity was found to be low by from 0.03 to 0.10 during 2005. This could mean that the drift was irregular during the year, or that there is some cause of error beyond calibration drift. Some freshwater may be getting into the system. During 2005-22 it was discovered that a hose was connected to the loop system in such a way that fresh water might have been introduced to the loop. The salinity at that time was found to be low by 0.10. Another possible cause of error is bubbles in the system as was reported in the log for Loop sample 19.

The temperature rise in the loop due to heating by the ship was found to be 0.2Cº during February 2005.

Conclusions

The intake temperature is lower than the CTD temperature by about 0.007. The lab temperature is higher than the CTD by 0.215Cº and higher than the intake temperature by 0.225Cº, so an estimate of 0.2Cº is reasonable.

The salinity was found to be low by between 0.017 and 0.028. An estimate of 0.02 is a compromise between the comparison with the well-mixed CTD casts and the loop samples with low standard deviations.

The fluorometer on the TSG is almost twice that of the CTD and 6 times the loop chlorophyll. No calibration will be applied to the fluorescence.

f.) Recalibration

File 2006-08-TSG-recal.ccf was used to apply offsets of -0.22 Cº and +0.020 to Temperature:Primary and Salinity:T0:C0. After calibration the two temperature channels were compared and are much closer.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Primary, Flag and UPoly0 (flow rate).

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and to add the depth of sampling to the header and to add a general comment about the sampling method as well as the following remark:

Fluorescence is nominal. Temperature and salinity have been recalibrated

based on comparison with CTD data, bottle salinity and intake temperature.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun. One CHE file was found to have salinity data missing; there was an error in the file that produced the SAL files. That was corrected and the CHE file re-derived; no further errors were detected.

The sensor history was updated for the TSG and CTD sensors.

Particulars:

1. No DO sensor on this cast.

5. Saved as #4; file name changed after SeaBird conversion.

26. Alarm went off at about 1000m. No spikes noted

29. Alarm went off at 1320 & 1790 on way down, 1300 on way up.

40. Problem with transmissivity from 1300m down onwards.

41. Transmissometer cable changed – still no signal.

42. No transmissometer mounted

43. Repeat of cast #42 using different CTD for comparison.

44. Few beeps and error light at 125db of upcast.

48. CTD 0585 with transmissometer 498DR.

52. Reversed direction during upcast to fix spool on winch.

Loop 19 – lots of bubbles in TSG 0003 Feb. 13.

63. No DO sensor. #13 did not fire.

# Institute of Ocean Sciences    CRUISE SUMMARY

| Cruise ID#: 2006-08 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 30 January 2006 End: 19 February 2006 |  |  |  |  |  |  |
| Location: Line P |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Robert M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |
| 2 | SEABIRD | 911+ | 0585 | No | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2006-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 11Feb05 | Factory |  |  |
| Conductivity | 1766 | 07June05 | “ |  |  |
| Secondary Temp. | 2106 | 08Jul05 | “ |  |  |
| Secondary Cond. | 1729 | 12Jul05 | “ |  |  |
| Transmissometer | 498DR | 22May05 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 24Nov05 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Altimeter | 1024 |  |  |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

Make/Model/Serial#:	SEABIRD/911+/0585         Cruise ID#:		2006-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19/03/05 | Factory |  |  |
| Conductivity | 3038 | 03/03/05 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 2102 | 07/06/05 | “ |  |  |
| Transmissometer | 333DR | 29/02/05 |  |  |  |
| SBE 43 DO sensor | 0047 | 10/02/05 | Factory |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 77511 | 30/09/1999 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2006-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 21/10/05 | Factory | 21/10/2005 |  |
| Conductivity | 2487 | 21/10/05 | “ | 21/10/2005 |  |
| Secondary Temp. | 2416 | 5/08/04 | “ |  |  |
| Wetlab WetstarFluorometer | WS3S-713P | 10/01/01 | ? |  |  |
| Flow Meter | ? | ? | ? |  |  |