## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 6 Feb. 2019 | Bottle spreadsheet converted to searchable BOT files. |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2006-09

Agency: Ocean Sciences Division

Location: North-East Pacific

Project: High Seas Salmon

Party Chief: Morris J.

Platform: W.E. Ricker

Date: February 28, 2006 – March 26, 2006

Processed by: Germaine Gatien

Date of Processing: June 15, 2006 – June 30, 2006

Number of original CTD casts: 129   	Number of CTD casts processed: 128

Number of TSG files: 2			Number of TSG files processed: 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0558) was mounted with a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0471) and the logging computer was the PAC CTD on the Ricker. The salinometer used was a model 8400B Autosal (serial number 68572).

# SUMMARY OF QUALITY AND CONCERNS

There was an error in the pressure sensor configuration used at sea; an offset of -0.6db was entered whereas +0.4db would be appropriate. The operator would think the CTD is a metre shallower than it really is. If it is important to get close to the surface, then con files need to have the correct pressure co-efficients. As they get older there is drift in the offset of the pressure sensors, so this will become a more serious issue with time.

The pumps were generally not turned on until the CTD was well below 5db.

Salinity calibration sampling is limited to 10db and 1 deep sample. The history for these conductivity cells is also based on very little deep sampling. While there is no indication of large errors in the salinity the decision to not recalibrate should be revisited when post-cruise calibration is done.

There were severe problems in the salinity for 12 casts between #244 and #325. There is a note in the CTD log that the CTD often appeared to be “frozen” when deployed and ice in the cell is mentioned once. It is not uncommon to see bad conductivity data that clears at about 50m – such problems may be caused by clogged bleed valves, but during this cruise the problems usually affect both upcast and downcast primary data in areas of large temperature gradient and occasionally both pairs of sensors were affected. Secondary downcast data were selected for casts 244, 247, 250, 256, 262, 268, 283 and 289. For cast 256 the upper 23db was removed from the downcast secondary data. Upcast primary data was selected for casts 259, 265 and 379. For cast 277 the downcast primary salinity was selected, but edited heavily in the top 10db.

The thermosalinograph data is suspicious in Muchalat Inlet. The CTD data were not shallow enough to rule out the highly variable temperature and salinity values observed, so the data was left unedited.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained along with salinity analysis sheets.

The file names were changed to standard format.

The bottle salinity data were obtained in spreadsheet form with flag and comment channels; there were no flags applied.

The nutrients were obtained in spreadsheet format with flags and comments.

The cruise summary sheet was completed. The histories of the conductivity and pressure sensors were obtained as well as that of the thermosalinograph conductivity and temperature. Calibration constants were checked. One error was found and corrected in the conductivity calibration. As has been found many times for this CTD, the pressure offset was set to -0.6db whereas +0.4db has been found appropriate for the past 2 years, so +0.4db was entered for this cruise as well. The new con file was named 2006-09-ctd.con.

##### Conversion of Raw Data

All data were converted using the con files given above. A few casts were examined and all expected channels are present. Both T and C channels show fine-scale noise even when the descent rate is steady and the upcast differs significantly from the downcast. The differences between sensors are larger on the upcast than downcast. The secondary temperature looks smoother than the primary but the secondary conductivity has much more fine-scale noise than the primary.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on three casts using different settings for CELLTM to determine the best choice of parameters. CELLTM was run using (0.03, 7) for the primary and (0.03, 9) for the secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Three casts were plotted to check for agreement between the pairs of T and C sensors. The following values are rough estimates from downcast data:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 65 | 250 | -0.0006 | -0.0003 | -0.0025 | Steady |
| 133 | 250 | -0.001 | -0.0004 | -0.003 | Steady |
| 253 | 500 | -0.0007 | -0.0003 | -0.0032 | Steady |

The temperature difference is fairly large. There is a slight indication of drift in the salinity differences, but no significant pressure dependence. There is a lot of noise in the conductivity differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and an error was found in one file. The CNV file had scrambled Sea-Bird headers. Those were fixed and conversion and CLEAN rerun. No further errors were found.

The header summary was run and station names were checked against the log as well as a random sampling of positions and times to ensure that the GPS was working well. The only problem noted was that station names were missing from the first 30 files. A text editor was used to add station names based on the log entries.

The track plots (using event #s and station names) were produced and look reasonable; they were added to the end of this report.

The average surface pressure is 6.1db which is deeper than expected. Hugh Maclean says that he does not actually see the deployment while he is operating the computer, and that the crew probably just had a different approach from the usual.

There is a note in the log that the data from the top 10m was considered suspect and on at least one occasion the CTD was said to be FROZEN when first deployed.

On cast #4 the pressure was very odd at the beginning of the record, with readings of -9db which gradually changed to reasonable values with no significant change in T and S. Doug Anderson believes that this is probably due to a bit of salt in the system and since it was a one-time only occurrence no further action is needed.

##### 10. COMPARISON WITH BOTTLES

Salinity

The mixed layer was shallow for most of these casts. Given that it is impossible to determine the exact depth from which the water in the Niskin bottles derive, salinity calibration should only be based on areas where salinity gradient is low around the depth of sampling which was usually around 10-11db. So a calculation of mixed-layer depth was done based on a salinity change of 0.002 and a reference depth of 9db. The casts that were best mixed based on that calculation were examined; some were rejected because the local values were very noisy even though the average gradient was ok. The secondary salinity is noisy so there is an added risk of error in that channel as this comparison was done by picking off values from an on-screen plot.

The spreadsheet of salinity values was simplified and reduced to those casts with low gradient around 10db and the deep bottle from cast #253. The best 11 shallow casts suggest that the primary salinity is low by 0.0054 if we use downcast data and 0.0058 using upcast data. The secondary salinity was found to be low by 0.0098 and 0.008 using downcast and upcast data respectively. The downcast data is considered cleaner, since the CTD is ahead of the wire and bottle so the ambient water is less disturbed. However, the CTD was in motion at 10m of the downcast, but stopped for the upcast and the samples were taken during the upcast and surface water can experience significant differences in the time taken to run the cast.

There were two bottles that were fired at 1 or 5m (this is not clear) and in each case the shallowest CTD downcast primary salinity is close but slightly less than the bottles (0.004 and 0.019 difference).

Comparison with the one deep bottle (500m) shows that the primary salinity is high by 0.0014 and the secondary low by 0.0019. Unfortunately this cast is one for which the primary salinity was poor in the top 100m. It was probably ok at the level of the bottle.

Another source of error is that the Niskin might not have been exactly 5m above the CTD as assumed in the comparison. Hugh Maclean says that the usual method was used in mounting the bottle, so there should not be large variations, but for the shallow comparison even small variations could result in significant and possibly systemic errors.

Fluorescence versus titrated chlorophyll

Plots were made of CHL from the upcast bottle and Fluorescence from averaged downcast CTD files. We don’t expect a great comparison. The ratio of Fluorescence to CHL ranges from 0.15 to 4.3 with the highest values occurring for low chlorophyll values. The average ratio is 0.84 when values are included for CHL<3. There is considerable variability with time probably due to the frequent excursions into inlets followed by offshore sections.

##### 11. SHIFT

Fluorescence

The fluorescence is frequently very noisy and often low, so it is difficult to judge what setting is best for aligning the signal with temperature. A shift of +24 records (1s) was applied as that has been used for all recent cruises. (Output: SHFFL)

Conductivity

Tests were run on the primary conductivity using a variety of settings and the best results overall were with a choice of -0.5s, the same setting used for this sensor when it was last used.

All casts were put through SHIFT to advance the primary conductivity by -0.5s. (Output *.SHFC).

SHIFT was not run for the secondary conductivity for most of the casts since the primary was a better choice for archiving, but it was later discovered that for some casts there were significant problems in the primary channel so the secondary was considered as an alternative. SHIFT was run on those casts only using a setting of -0.2 based on the results of 2006-08.

##### 12. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The pumps were not turned on for cast #1 so it will not be processed beyond this point.

##### Other Comparisons

Previous experience with these sensors –

- The primary sensors were used for many cruises during 2005 and salinity was found to be within 0.0012 except for one cruise when it was low by about 0.003. There was a lot of scatter for that comparison. During 2006-08 there were no bottle comparisons for this CTD but a repeat cast using another CTD suggested that there was no significant error in salinity.

- The secondary sensors were used during 2006-09 when the salinity was found to be high by 0.002 but there was a lot of scatter. During 2006-08 there were no bottle comparisons for this CTD but a repeat cast using another CTD indicated that salinity was low by ~0.003.

Historic ranges – Profile plots were made of T and S with local climatology superimposed. The only excursions from the climatology were for salinity at a few shallow casts, with salinity often lower than the historic minima. In one case values were found that were above and below the historic limits at different depths. The ranges themselves look odd for many of these cases as though there was very little sampling. These excursions probably reflect real variability that is not represented well in the historic ranges. There is no suggestion of instrumental error.

##### DETAILED EDITING

The primary salinity appears to be closer to the bottles and the secondary conductivity signal was very noisy. There is some indication of problems with the downcast data near the surface but it occurs in both sensor pairs. The secondary temperature seems particularly noisy near the surface. The data used for the bottle comparison was also examined to see if the downcast data at 10db was significantly different from that of the upcast data in areas of low-gradient. The average difference for the primary salinity was 0.0002 and for the secondary it was 0.0026. The primary looks like the best choice overall.

Because there are some doubts about the surface data, consideration was given to using upcast data. One cast from which the downcast and upcast data were quite different was studied and the downcast data looked much better. The upcast data had large areas of instability (especially at the surface) even after putting the data through DELETE. The upcast looks like a choice of last resort..

Page plots were produced using (T0, S0). These were used to guide the editing.

In the course of editing, 12 casts were found (in the second half of the cruise) for which the downcast primary salinity data was found to be unusable especially between the bottom of the mixed layer and about 60db. There were large instabilities and upcasts and downcasts were notably different. In some cases both channels were bad, but for most, either the secondary or upcast data looked better. The low-gradient casts discussed in section 10 did not include any of the problem casts. The malfunction appears to be related to the time-response rather than the calibration since it is worst in areas of large temperature change. There is mention in the log book that there might have been ice in the cell, which would certainly explain problems in the downcast, but would that explain upcast problems? The following choices were made:

- For casts 244, 247, 250, 262, 268, 283 and 289 the primary conductivity was bad in the top 100m during both downcast and upcast. The secondary conductivity was ok for those casts so the secondary channels were selected. A return was made to step 11 to run SHIFT on the secondary conductivity and DELETE was rerun.

- For cast 256 both conductivity channels were bad to 20db during both downcast and upcast. The secondary data was better below that. Downcast secondary data was used but the data above 20m was removed from the cast using CTDEDIT.

- For casts 259, 265 and 379 both conductivity channels were bad (379 was suspicious to 20db) during the downcast, so upcast primary data was selected. The shifted data was put through REVERSE and then DELETE.

- For casts 277, both channels had problems; the primary salinity was selected and edited heavily in the top 10db.

On-screen plots of descent rate and pump status were also used to guide editing.

All casts required some editing except: 53, 89, 104, 107, 111.

Note was made of the editing details in the headers of the relevant files.

After editing a comparison was made of all shallow casts for which there was bottle sampling. Earlier only the casts with well-mixed surface waters were used. A comparison was made with both unedited and edited data (the latter include some upcast data) to see if there was any significant difference around the time when the CTD data looked poor, in the hope that any other bad casts could be identified. The differences show tremendous variability but there is no obvious pattern suggesting which casts had the problems noted above. When casts were rejected that had differences >0.1 for either salinity channel the average showed the primary being low by 0.0015 and the secondary low by 0.004. The deep bottle indicated that the primary was high by about 0.0014 and the secondary low by 0.019. The shallow results each show the CTD being lower by 0.002 to 0.003 than found by the deep bottle. Keeping in mind that the CTD data comes mostly from the downcast without a stop and the bottles from the upcast after a stop, this is reasonable. We expect the CTD to read slightly lower than the bottles under those conditions.

##### Initial Recalibration

The salinity comparisons suggest that the primary salinity is either high by 0.0014 based on the one deep bottle or low by 0.0015 based on the shallow bottles when outliers are excluded. The history of the sensor is consistent with the results from the one deep bottle. For the secondary sensors previous comparisons suggest salinity is high by 0.002 or low by 0.003, but both are based on very little information. The one deep sample from this cruise suggests it is low by 0.002.

Recalibration is not justified.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Some plots do show some unstable features and the files were checked again in CTDEDIT. Some were in areas of active mixing where the features are believable. For the others there was no basis for choosing which data is believable and which not. No further editing was done to T-S.

Temperature, Salinity and Fluorescence profiles were plotted on-screen and no problems noted.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts except 244, 247, 250, 256, 262, 268, 283, 289: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

The following channels were removed from casts 244, 247, 250, 256, 262, 268, 283, 289: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

HEADER EDIT was used to fix formats and channel names, to fix the DO sensor serial number and to add the following comments:

Fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

There were problems in the salinity data for casts #244–250, 256-268, 277, 283,

289 and 379. Large unstable features and large differences between downcast and upcast

were noted. The trouble appears to be due to problems in the pump or conductivity cell,

possibly involving ice in the cell. Secondary sensors or upcast data were chosen

where appropriate (noted in the individual headers), but all data  from those casts

should be  considered of lower quality since there is no calibration  sampling from

the  depths where the problems were noted.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### 18. Thermosalinograph Data

a.) Checking calibrations

There were 2 identical configuration files for the TSG. One was renamed 2006-09-tsg.con. The calibrations were correct. There is only 1 temperature channel and no fluorometer or flow channel.

The 2 data files had non-standard names. They were renamed as 2006-09-0001.cnv and 2006-09-0002.cnv.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Latitude, Longitude, Salinity:T0:C0, Time Julian and Pressure. There was no channel for Flow Rate and no fluorometer. Pressure is not usually converted but in order to study alignment it was done this time. These files were then converted to IOS HEADER format. The name of the pressure channel had to be changed from prM: Pressure to prSM: Pressure in order to be converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

A time-series plot was produced. There is a lot of variability in T and S and most is believable given that the ship moved in and out of inlets. However, there are some spikes in salinity that need investigating and a large section where the temperature drifts upwards and suddenly comes back to expected values.

An initial track plot was produced and shows that there are some bad data points in positions in both files. Using a text editor the bad positions were replaced with values determined by interpolation. (File 1: scans #6609, 29435 and 48327; File 2: scan #2100.) The plots then looked fine.

An initial look at the data show a lot of spikes that are not huge, but clearly not real. A study was made to see if shifting the conductivity relative to temperature would remove the spikes in salinity. An advance of about +1 records did reduce some spikes slightly but not enough to avoid editing, and it increased at least a few other spikes. The mismatch appears to be non-constant, so SHIFT is probably not appropriate.

CTDEDIT was used to edit the first file to smooth salinity spikes and to remove a section in which temperature steadily increased to 14ºC and then suddenly (over 3 records) reverted to values of <8ºC. The second file had no data that was clearly bad. There is a section at the end of the first file and the beginning of the second that is suspicious. There is very cold fresh water with occasional warmer, saltier values. A study was made of the temperature and salinity in the Muchalat Inlet area. As the ship went eastwards the temperature and salinity fell to very low values. At the head of the inlet there is a spike to high, salty water. On the return trip westwards the values are generally cool and fresh, but there are intermittent bursts of higher values. It is not clear if this could be real or not. The higher values are associated with changes of direction in the ship track, so might be due to the ship heeling over enough so that the intake moves from one layer to another. The depth of the TSG intake is about 3m. It seems unlikely that such a cold, fresh layer would exist at that depth, but the CTD and calibration samples are not shallow enough to rule out such low values. Is it possible that water from a fresh surface layer got sucked down by some ship effect? (See tsg_study_Muchalat.xls for plots.)

c.)  Checking Time Channel

The CTD data from the downcast is not useful since the pumps were rarely turned on at the level of the TSG. Upcast data that was put through REVERSE and DELETE was used to get estimates for any cast that sampled above 5m. The data was bin-averaged using 0.5m bins and thinned. The record between 4 and 5db that was closest to 4.5db was chosen. Not all casts had any sampling at that depth.

The TSG files were opened in EXCEL and reduced to the times when CTDs were run. Those files were combined in a spreadsheet matching times. The positions were compared and the average differences in latitude and longitude were 0.0012º and 0.0025º. There were 5 casts with fairly large differences, but all were close to shore and most in narrow passes where drift was likely to be large. When those 5 casts are excluded no differences were larger than 0.002º and the average differences were both 0.0001º, so the clock would appear to be accurate.

This spreadsheet will also be used in the next step to compare temperature and salinity. (See 2006-09-tsg-ctd-comp.xls)

d.) Comparison of T and S from TSG and CTD data

Note: There was no intake temperature channel for comparison with the TSG. There were no loop samples, and only 2 bottles from shallower than 10m.

The spreadsheets comparing CTD and TSG files were examined to find the differences between the salinity and temperature channels for the CTD and the TSG. This comparison is limited by the fact that we had to use upcast CTD data, but have TSG data from the beginning of the cast. When all data is included the TSG temperature was lower than the CTD by 0.0193Cº which is unexpected since it is usually higher due to heating from the ship. The TSG salinity was found to be lower than the CTD by 0.4525 which is a much greater difference than noted in other recent uses.  There were some very large differences. This is likely due to the presence of a large gradient between the TSG intake and the shallowest CTD data. This is particularly likely for the end of the first file and the beginning of the second where there were huge variations in T and S. When casts with salinity differences >0.1 were excluded, the temperature of the TSG was found to be high by 0.071Cº and the salinity high by 0.004. The temperature difference is reasonable for the RICKER. The salinity difference is smaller than noted in previous uses. There is too much scatter to see any time dependence in the differences. (See TSG-CTD-comp-outliers_excluded.xls.)

g.) Calibration history

The TSG has been used for 8 other cruises since it was last recalibrated in January 2005 but all other uses were on the TULLY. During 2005-22 it was discovered that a hose was connected to the loop system in such a way that fresh water might have been introduced to the loop. The salinity at that time was found to be low by 0.10. Since then, the temperature was found to be high by 0.22Cº in September 2005 and February 2006 and the salinity was low by 0.045 and 0.02. The heating due to the ship is greater for the TULLY, so the temperature adjustment should be lower for this cruise. An estimate of 0.05C due to ship heating was made for 2004-32 when this equipment was used on the RICKER. That cruise was in autumn when the temperature difference between ocean and ship would be less, so we might anticipate a slightly higher ship effect on temperature for this cruise.

Conclusions

When outliers are ignored, the TSG temperature is higher than the CTD temperature by about 0.07Cº. This is consistent with the history of the instrument. Salinity was found to be high by 0.004 which is not consistent with the history of the instrument. Given all the uncertainties in using upcast CTD data in the comparison and not being able to match times accurately with the TSG, it seems best to use the salinity adjustment used for the February 2006 data. When the instrument is next recalibrated this should be revisited.

f.) Recalibration

File 2006-09-TSG-recal.ccf was used to apply offsets of -0.07 Cº and +0.020 to Temperature:Primary and Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary, Flag and Pressure.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and to add the depth of sampling to the header and the following comment about the data.

The temperature and salinity data from Muchalat Inlet (March 25) look

suspicious with values varying between warm, salty and cold, fresh. It

is not expected that water as cold and fresh as seen in the record

would occur as deep as 3m. The data was not edited since the CTD and

calibration sampling were not shallow enough to confirm that it is wrong.

As a final check plots were made of the cruise track and data; no problems were noted.

##### 19. Producing final files

A cross-reference listing was produced for CTD files.

HEADER CHECK was rerun and a cruise track plotted and no errors found.

The sensor history was updated for the TSG and CTD sensors.

Particulars: (Most bottles at 10m, suspicious surface data to 10m)

- Pumps off

4.	Temp diff large at 4m. Pressures negative at beginning of file.

151. Bottles at 1 and 5m

154. Bottles at 1 and 5m

226. Cast rerun because of problems with primary temp on first attempt

244. Primary salinity looks bad to about 60db; chose secondary channels.

247. Problems again with data, went to 20m, back to surface and then full cast – still bad, used secondary

250, 262, 268, 283 and 289 – used secondary channels downcast

253. Bottles at 10 and 500m

259, 265, 389 – Used UPCAST primary channels

262. Stopped at 155 of downcast for 1 minute

277. Used downcast primary, but edited heavily in top 10m

# Institute of Ocean Sciences    CRUISE SUMMARY

| Cruise ID#: 2006-09 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 28 February 2006 End: 26 March 2006 |  |  |  |  |  |  |
| Location: NE Pacific |  |  |  |  |  |  |
| Vessel: W.E.Ricker Party Chief: Morris J. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0585 | No | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585         Cruise ID#:		2006-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19/03/05 | Factory |  |  |
| Conductivity | 3038 | 03/03/05 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 2102 | 07/06/05 | “ |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 77511 | 30/09/1999 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2006-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 06/01/05 | Factory |  |  |
| Conductivity | 2488 | 06/01/05 | “ |  |  |