## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2006-10

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: WCVI / Broughton Inlet

Project: ECOHAB / LaPerouse / Juan de Fuca / Aquaculture

Party Chief: Juhasz T. / Stuchhi D.

Platform: Vector

Date: May 29, 2006 – June 12, 2006

Processed by: Germaine Gatien

Date of Processing: 4 August 2006 – 15 August 2006

Number of original CTD casts: 13 HEX files and 45 DAT files

Number of SBE25 CTD casts processed: 13

Number of SBE911+ CTD casts processed: 45

Number of rosette casts: 26

# INSTRUMENT SUMMARY

There were two legs and different equipment was used on each.

- LEG 1

A SeaBird Model SBE-25 CTD (#0334) was run with pressure sensor 290544; an SBE43 dissolved oxygen sensor (#0047), a transmissometer (S/N 498DR) and a fluorometer (SeaPoint S/N 2228) were also mounted on the CTD.

- LEG 2

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#723DR) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#0766) was mounted on the primary pump and a Seapoint Fluorometer (#2356) with a 10X cable was mounted on the secondary pump. The deck unit was a model 911 (#0425).

The salinometer was an 8400B model Autosal (S/N 68572).

# SUMMARY OF QUALITY AND CONCERNS

There were duplications in cast numbers. This was noted in the log. In order to resolve the problem the event numbers for Leg 2 were advanced by 3, so that cast #12 became cast #15 etc.

- LEG 1

Some data was lost because the pump did not operate in low-salinity water. This problem has occurred on other cruises - the cut-off conductivity has to be set much lower for work in water with such low surface salinity. For some casts upcast data was processed as well as or instead of downcast data. All files with a 9 in the first digit of the event number contain upcast data. So file 2006-10-0003 is from the downcast, 2006-10-9003 is from the upcast. For cast #1 the only useful data is from the upcast and only to 10m. For casts #3, 5, 6 and 7 both downcast and upcast files were prepared because the pumps did not come on until the CTD was down to 6, 9, 8 and 12m respectively. For all other casts only downcast data is to be archived.

The information on the rosette log sheet about how the Niskin sampling was done was misleading. The correct information was obtained from Lucius Perreault.

There was insufficient dissolved oxygen sampling to provide recalibration information. The sensor was used in anoxic waters which may cause calibration drift. The results of Leg 2 were applied to this data. Based on the few bottles available, it is believed that the data is probably high by ~0.15 in anoxic waters. Based on Leg 2 results errors of about ±0.6ml/l are expected above 20m and ±0.2ml/l below 20m.

The stops for bottle sampling lasted 30s or more.

- Leg 2

The first three casts of Leg 2 had event numbers assigned that had already been used during Leg 1, so all numbers were changed in Leg 2 by adding 3 to the event numbers. Thus, casts #11 through 55 became #14 through 58. A general note was made in the log book and on each of the rosette sheets.

Salinity calibration sampling near the bottom continues to look less reliable than a little higher in the water column. The transmissivity was not very low so mud in samples seems unlikely; perhaps the larger differences between CTD and bottles are due to the higher gradients seen right at the bottom so that the vertical separation between the CTD and the rosette is more significant.

The SBE dissolved oxygen data in the CTD files should be considered:

•	±0.6ml/l from    0 -  20m

•	±0.2ml/l from   20 - 175m

•	±0.1ml/l below 175m

The stops for bottle sampling lasted 30s or more.

Dissolved oxygen saturation was very high near the surface for many casts, but bottle samples confirm that such values are correct.

Fluorescence went off-scale near the surface for most casts during both downcasts and upcasts.

************************************************************************************

FROM THIS POINT ON THE PROCESSING IS DIVIDED INTO SEPARATE SECTIONS FOR THE TWO LEGS. ALL MAPS ARE AT THE END OF THE REPORT.

************************************************************************************

LEG 1 : May 29, 2006 – June 4, 2006

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

The Daily Log was obtained. The configuration file 0334CTD.con was obtained and all calibration co-efficients were found to be appropriate.

The times in the log book are in PDT.

The file names were almost in standard format, but required the addition of an extra “0” in the event numbers.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

All hex files were converted using 0334CTD.con. A few casts were plotted. The pressure signal has reversals on the order of 0.2db which is typical of this instrument.

All variables look normal except during cast #1. There is a note in the log that the pumps may not have turned on because the water was too fresh. It does look as though they came on during the upcast.

The descent rate is low for the early casts, but quite steady; later in the cruise the descent rate is noisier but also higher, on average.

4. WILDEDIT

WILDEDIT was run on all casts on pressure, temperature and conductivity channels using 2,20,25,0 for “Standard deviations for pass 1” and “Standard deviations for pass 2”, scans per block and “Keep data within this distance of the mean”.

5. ALIGN

This step was skipped. SHIFT will be used to align T and C.

6. WFILTER

Based on the results of many other cruises using this equipment, the SeaSoft routine WFILTER was run for all casts to apply a cosine filter, size 5, to the pressure, temperature and conductivity. After this step the pressure is quite smooth.

7. CELLTM

Based on the results of 2006-20 and other recent uses of this equipment CELLTM was run on all casts using α = 0.03, 1/β = 9.0. Tests on a few casts show small improvements from applying this routine.

8. DERIVE

Program DERIVE was run to calculate salinity.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert Sea-Bird ASCII data to IOS Headers.

The times are in PDT according to the log book. As noted during 2006-20 and 2006-27, the SBE25 date was ahead by 1 day. To correct for these two errors, ADD TIME CHANNEL was used to subtract 17 hours from the times (-24 for the date error +7 to convert to UTC).

At this point a text editor was used to add entries for Latitude, Longitude and Water Depth based on log book entries.

10. Checking Headers

Header Check was run and turned up no obvious errors.

A Header Summary was produced and checked against the log.

The Surface Check produces an average of 0.11db.

Track plots were produced and added to the end of this report.

Plots were examined on-screen of all variables. The first cast looks bad, presumably the pumps came on late, if at all. The upcast will be examined later to see if it is useful. Fluorescence is frequently off-scale in the top 20m. Other variables look as expected.

11. SHIFT

Conductivity

On previous cruises using this type of CTD good results were found when SHIFT was run to advance the conductivity by +0.7 records. Tests were run on 2 casts with noisy T-S curves to ensure that the same setting works well for this data set and it did.

SHIFT was run on all casts advancing the conductivity by +0.7 records.

Fluorescence

Based on the results of 2006-20 the fluorescence was aligned by applying a shift of +42 records to all casts. The data was then examined to see if the results were satisfactory. The offset between downcast and upcast trace looks similar to that for temperature after this step

Dissolved Oxygen

This dissolved oxygen sensor is noted for poor time response and hysteresis. For this cruise the slow descent rate should be helpful in getting better data. To correct for the time response we generally shift the channel by 100 to 160 records, but that is with sampling at 24Hz. For this cruise the sampling is at 8Hz. For 2006-20 a setting of +40 records gave the best overall results while for 2006-27 with a somewhat higher descent rate the best choice was +50. For this cruise the descent rate was higher, but the sensor used was different. Tests were done using shifts of +30, +40, +50, +60 and +80. The best results overall were with +40 so that was applied to all casts.

- REVERSE

The downcast data is bad in the top 5 to 15m for some casts, so the upcasts were examined. The data is better there, though the pumps do seem to go off prematurely for some casts, so the top 2 to 5m will have to be examined closely. All casts were put through REVERSE. Both downcast and upcast will be processed; a number 9 will be placed in the first digit of the event number for the upcasts. So file 2006-10-0001 will contain downcast data and 2006-10-9001 will contain upcast data. Later a decision can be made about which upcast files need to be archived.

- DELETE

Both the downcast and the reversed files were put through DELETE using the following parameters:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00    Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.2m/s (calculated over 5 points) was deleted

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval taken from the header.

Pressure was not filtered.

COMMENTS ON WARNINGS: There were no warnings.

The upcast files were also put through DELETE using the same parameters.

- DETAILED EDITING

Page plots were produced and examined for spikes and instabilities and used to guide the use of CTDEDIT. The downcast data was edited in two passes, with the first displaying T, S and DO because the dissolved oxygen channel makes it clear whether the pump is operating or not. Records were removed where it is clear that the pump was off. Sometimes temperature and salinity traces look ok, but it takes a few more seconds for the DO data to reach positive values so those DO data points were replaced with pad values. Cast #1 had no useful downcast data. Some surface records were removed from all other casts. On the second pass T-S plots were displayed and further editing was applied to casts #3, 5 and 7.

For casts 1, 3, 5, 6, and 7 there was a significant loss of downcast data due to the pump coming on late, so upcast data will be processed for those casts. They were edited as described for the downcast data.

Upcast and downcast files were examined together to see if patching could have good results. In most cases that would result in a sudden shift in values and no patch is possible for cast #1 since there is only upcast data available. Surface data was removed from cast #1 below 9.8db; the deep waters of cast #1 could be anoxic explaining the low DO values, but the salinity is also very low so it is believed that the pump did not operate at all during the downcast and only came on at about 10m of the upcast. The pumps turned off near the surface for most upcasts but not for casts #11, 12 and 13, presumably because the surface conductivity for those casts was higher than the cut-off set for the pump.

All casts required some editing. Note was made of the editing details in the files.

- BIN AVERAGE

The following Bin Average values were applied to the ED1 files:

Bin channel = pressure 		Averaging interval = 1db

Minimum bin value =   .000	Average value will be used

Interpolated values are NOT used for empty bins.

- Comparison of bottles to CTD salinity and dissolved oxygen

Calibration sampling was done during 3 of these casts. The log suggests that a rosette was used, but no rosette files could be created. In fact there were 2 Niskin bottles mounted on the wire.

Information available about the depth of sampling is as follows:

| Sample # | Cast # | Sample Depth entered on rosette log sheet | Bottom Depth from CTD | Comment | Estimate of sample depth |
| --- | --- | --- | --- | --- | --- |
| 1 | 3 | Bot-5 | 113.3 | DO only | 108.3 |
| 2 | 3 | Bot-15 | 113.3 | DO only | 103.3 |
| 3 | 5 | 203 | 196.1 | DO/Sal | 191.1 |
| 4 | 5 | 191 | 196.1 | DO/Sal | 181.1 |
| 5 | 10 | 93 | 91.3 | DO/Sal | 86.3 |
| 6 | 10 | 83 | 91.3 | DO/Sal | 76.3 |

The depths given for casts #5 and 10 don’t make sense since there is no CTD data from as deep as that given for the deeper samples. Lucius Perreault says that the data were collected using Niskin bottles mounted about 1m and 11m above the CTD. So the depth is estimated based on maximum sampling pressure for each cast.

Salinity

The results, given in 2006-10-sal-comp.xls, are scattered and show the CTD to be high by 0.0042, 0.0063, 0.0027 and 0.0285. Excluding the last sample which was in a low transmissivity area, the average indicates the CTD is high by 0.0044. Low transmissivity has been associated with problems in analysis that lead to anomalously low titration results.

Dissolved Oxygen

There were DO samples from 3 casts only. The values from the ADD files were exported to spreadsheet 2006-10-DO-comp-leg1.xls and values were read from bin-averaged, edited downcast files. Because of the sampling method (CTD values taken while instrument was in motion) the differences will reflect a combination of errors due to calibration drift and time-response delays. So it is not readily comparable to other uses of this equipment.

Only 2 samples came from water with DO>0.1 ml/l and one of those was flagged and is out of line with the other. So we have 4 near-zero values and 1 other to analyze. Moreover, there is no sampling in areas of high DO. The trendline (excluding the flagged sample) is 1.135 * CTD:DO -0.054, but the significance of this is questionable. Moreover, it has been observed in the past that DO sensor calibration  drift during the next few casts after exposure to anoxic conditions. Heavy sampling is needed to resolve such a shift in calibration.

The results from 2006-11 and the second leg of this cruise have limited applicability because neither sampled any water with DO<2 ml/l. An attempt to combine the results of the two legs did not lead to a result that worked well for either leg.

The results of cruise 2006-10 Leg 2 using this DO sensor involved a two-step calibration involved first applying:

DO(corrected) = 1.0606 * DOX_CTD + 0.1138

followed by a second correction which was a simple offset of -0.017ml/l. The 2 corrections which can be applied in a single step:

DO(corrected) = 1.0606 * DOX_CTD  + 0.0968

When the Leg 2 correction was applied to the Leg 1 CTD data used in the comparison with bottles, the differences between the bottle samples and corrected CTD data were within 0.17ml/l. The CTD DO was high for the 4 bottles with very low DO values and too low for the bottle with DO ~3.5ml/l.

The offset would be of opposite sign for upcast files, hence the correction is:

DO(corrected) = 1.0606 * DOX_CTD  + 0.1308

- Intercomparisons

Sensor History

- Dissolved Oxygen

For 2006-11 the following correction was used to correct the error due to calibration drift and time response errors for the downcast files:

DO(corrected) = 1.0663 * DOX_CTD +0.0783

followed by a pressure-dependent time-response correction to the downcast files:

DO(corrected) = DOX_CTD -0.0006 * Pressure + 0.4532

For Leg 2 of 2006-10 the following calibration was applied:

DO(corrected) = 1.0606 * DOX_CTD +0.1138

followed by a simple offset of -0.017ml/l to correct for time response..

- Salinity

The conductivity sensor has been used on 2006-20 and 2006-27. During 2006-20 the differences between the CTD and bottles were very high, ~0.04. There was no salinity sampling during 2006-27.

Historic ranges –The only excursions from the local climatology are in the top 6db at which level the salinity is often below the historic minimum and the temperature sometimes above the historic maximum. This does not look like a calibration issue as the differences are large. Either surface conditions were unusual, the pump may not have been on long enough for the data to be reliable or the climatology does not have sufficient data at these depths. The upcast data are also outside the limits at the same places. These excursions are not considered evidence of calibration problems.

- Recalibration

DO: The results of Leg 2 will be applied but errors may be larger, and in particular the CTD DO is likely to be non-zero for the anoxic water. The correction for the downcast files as explained in section 16 is:

DO(corrected) = 1.0606 * DOX_CTD + 0. 0968

And for the upcast files the correction is:

DO(corrected) = 1.0606 * DOX_CTD + 0.1308

Sal – The results of 2006-20 are not trusted. The comparison based on the 3 of the 4 bottles of this cruise suggests that the CTD is high by 0.0044 from all casts. This is weak evidence, but the best we have. So the salinity values will be reduced by 0.0044, but if a large drift is found when this sensor is next recalibrated then 2006-10 and 2006-20 should be revisited.

After recalibration one cast with reliable data in both the upcast and downcast was examined by plotting  upcast and downcast together. The two dissolved oxygen traces were closer after recalibration than before. A check was made and no negative values were found in the dissolved oxygen channel.

- REMOVE

The following channels were removed from all casts: Scan_number; Conductivity:Primary, Descent Rate, Oxygen:Voltage:SBE and Flag.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

- HEADER EDIT

Header Edit was used to fix formats and channel names and to add the following comment:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The salinity recalibration is based on only 3 bottles.

The dissolved oxygen recalibration is based on sampling during the second

leg of this cruise when no low DO values were found (DO was >2ml/l.).

Because anoxic waters were sampled during Leg 1 it is estimated that the

errors will be larger for low DO and the following is a rough estimate:

•	±0.6ml/l from 0 - 20m

•	±0.2ml/l below 20m

For the files that came from upcast data an extra comment was placed in the headers as follows:

For cast #1:

***************************************************************************

WARNING: THIS DATA COMES FROM THE UPCAST PORTION OF THE DATA
FILE. THE DOWNCAST AND MOST OF THE UPCAST WERE NOT USEFUL BECAUSE
THE PUMP WAS NOT ON.

***************************************************************************

For casts #3, 5, 6 and 7:

***************************************************************************

WARNING: THIS DATA COMES FROM THE UPCAST PORTION OF THE DATA
FILE. THE DOWNCAST FILES WERE ALSO PROCESSED BUT ARE MISSING SOME NEAR-SURFACE DATA BECAUSE OF PUMP PROBLEMS.

***************************************************************************

- Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

c.) The conductivity, dissolved oxygen and pressure sensor history files were updated.

Particulars

1. Pump not have turned on until about 9db mark of upcast.

3/5. Bottle sample smelly – anoxic near bottom

3. Pump not on until 6.2db downcast.

5. Pump not on until 9.6db downcast.

6. Pump not on until 7.9db downcast.

7. Pump not on until 12db downcast.

13. CTD hit bottom

## Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2006-10 LEG 1

Dates: Start: 29 May 2006 End: 4 June 2006

Location: WCVI / Effingham Inlet

Vessel: Vector

Party Chief: Juhasz T.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334		Cruise ID#:		2006-10 LEG 1

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 290544 | 13May05 | Factory |  |  |
| Temperature | 4054 | 14Oct04 | Factory |  |  |
| Conductivity | 2754 | 15Oct04 | Factory |  |  |
| Dissolved Oxygen | 0047 | 10Feb05 | Factory |  |  |
| Transmissometer | 498DR | 28Mar06 | IOS |  |  |
| Fluorometer | 2228 | ? | ? |  |  |

LEG 2: June 6, 2006 – June 12, 2006

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained.

As noted in the log book the event numbers were assigned improperly with 3 numbers duplicating ones from the previous leg. The file names were all adjusted by adding 3 to the cast number section of the names.

The salinity data was divided into 2 spreadsheets for the two legs; event numbers were increased by 3.

Individual dissolved oxygen bottle files were obtained; flag channels and comments were included. File names were adjusted by adding 3 to each event number.

The CTD calibrations in file 2006-10-CTD0443.con were checked and no errors found. The pressure offset has been drifting and may need increasing.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

The files were converted to CNV files. A few files were checked and the pressures are never less than 0.9db. An initial look at the data showed that the secondary conductivity values were very low.

The descent rate was usually quite steady and reasonably high. Bottle stops were typically >30s.

The down and upcast temperature and conductivity channels are reasonably close though the primary conductivity has some spikes and the upcast traces are noisier than the downcasts.

Transmissivity looked fine with good agreement between downcast and upcast.

The dissolved oxygen looks much like during 2006-11 when it was noted to be better than usual with upcast and downcast offset vertically, but with less hysteresis than has been observed in the past. This sensor was repaired in November 2006. It performed well during 2006-11, but was used in anoxic waters during the early part of Leg 1 of this cruise. That may cause calibration drift but the latter section of Leg 1 was not in anoxic water so the sensor would probably have returned to its usual calibration by Leg 2.

The fluorescence was frequently off-scale (>15), but otherwise looks ok.

The altimetry is often extremely noisy but generally looks reliable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s. Those files were then converted to IOS HEADER format. The extensions were changed to BOT. All BOT files were plotted and a few significant outliers were found in the secondary salinity for cast #33. These were cleaned using CTDEDIT and the output copied to BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Two casts with a steady descent rate and no bottle stops were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5) were tried and the best results for both channels were with the latter two. CELLTM was run using (0.03, 9) for both conductivity channels, the same as the choice made for 2006-11.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 15 | 220 | -0.0025 | +0.00007 | +0.0017 | Steady |
| 29 | 250 | -0.0027 | +0.00007 | +0.0018 | Steady |

The results are very close to those seen during 2006-11. As noted during that cruise the temperature differences are much larger than usual even though the descent rate was very steady. At that time it was found that most noise in the temperature differences occurred at large gradient areas where the secondary temperature was spiky. Similarly all noise in conductivity and salinity differences were associated with noise in the secondary channels.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. The event numbers were correct in the headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header summary was checked against the log book entries and no errors were found.

The header check was run and no further errors were found.

Track plots were produced and added to the end of this report.

The average surface pressure is ~1.5db which is reasonable for the Vector, and where the surface value was low the salinity was also low.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

##### BOTTLE FILE PREPARATION

The salinity spreadsheet was edited to remove the Leg 1 samples and to change the event numbers by adding 3 to each one; the spreadsheet was simplified and headers changed to standard format. The spreadsheet was then converted to SAL files. No data had been flagged.

The ADD files for Leg 2 as prepared by the analyst were copied to the HYDRO folder and edited to change the file names and event #s by advancing both by 3, to reflect the changes made to other file names, and the comments format was fixed.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Sample #212 was assigned twice, but the first time was for a bottle that was not sampled (cast #55.) That line was removed from the addsamp file.

Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The SAL and ADD files were merged with the CST files in two steps. (Output: MRG1, MRG2.)

Those files were then merged with SAMAVG files (MRG) and then put through CLEAN to reduce the headers to File and Comments only, remove remarks and keep comments, remove SeaBird headers and comments from the secondary files. (MRGCLN)

##### 11. COMPARE

Salinity

COMPARE was run. All bottles were between 200 and 432db. When a few outliers were excluded the differences between CTD and bottles were fairly flat with pressure for both pairs. The primary was found to be high by an average of 0.0009 and the secondary high by 0.0031. There is a slight suggestion of a drift to higher differences with time, but there is too much scatter and too little data to put any confidence in that.

A comparison was done that excluded any bottom sampling and that showed the CTD to be low by an average of 0.0005 but includes only 8 points. When a comparison was done using only samples from the bottom of casts the CTD was found to be high by 0.008. All were within 13m of the bottom according to the altimeter. The transmissivity was as low as 35% for a couple of the stops, but this does not seem low enough to suggest the samples were contaminated by mud. The other possibility is that the gradient is higher at the bottom and given the CTD is deeper than the bottle the differences are likely to be positive. Whatever the case, the CTD is within 0.001 of the bottles so no recalibration is called for.

The outliers were from the bottom of casts #21 and 24. During those stops the pressure was very steady. During cast #21 the temperature and salinity varied more than enough to explain the difference between the bottle and CTD. The three samples for cast #24 were taken close together at the bottom but looking at the CTD data during the stop it is clear that a sudden change occurred in temperature and salinity towards the end of the stop. The salinity in the bottles went down by 0.006 between the first and third firing and the CTD values went down significantly at about the time of the 3rd bottle firing, so there appears to have been a lag. It seems likely that the bottles are fine, that the outliers are just a sign of rapid change. No salinity samples will be flagged.

Dissolved Oxygen

COMPARE was run. Plotting the differences versus CTD DO and pressure excluding any samples flagged “d”, only 1 major outlier was found. That was for the bottom sample of cast #40. There is little variation in the CTD data, and a plot of DO versus salinity makes the sample look out of line, not the CTD data. So that sample was flagged “d”.

The differences versus file pair number were examined first since the sensor was used in anoxic waters during the first leg of this cruise, which has been known to cause some calibration drift in the past. However, that occurred early in Leg 1, so is hoped any effect would have worn off before this leg began. The plot suggests that there is no notable time dependence. When samples from above 2db are removed the trendline is very flat though there is a lot more scatter towards the end. Given the last casts are well up Knight Inlet, this may not be surprising.

The differences versus pressure are similar to those seen during 2006-11, much tighter than in the past. The differences versus CTD DO show the most scatter, but when the shallowest samples are removed it is much tighter. Removing outliers by removing points with residuals <-0.4 and >0.4 produced a fairly tight relationship for both plots.

DOX-BOT = 1.0606 * DOX-CTD + 0.1138

DOX-BOT = DOX-CTD -0.0006 * Pressure +0.4513

This is quite close to the results of 2006-11:

DOX-BOT = 1.0663 * DOX-CTD + 0.0783

DOX-BOT = DOX-CTD -0.0006 * Pressure +0.4532

There is slight evidence of time dependence, but the range of DO values varies from place to place making it impossible to separate temporal drift from geographic variation.

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. A few slight outliers were investigated, but there was nothing to suggest that further flagging of bottle DO data was justified.

##### Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors were used during 2006-11. The primary sensors were found to be low by 0.0008 and the secondary high by 0.0014.

Both conductivity sensors were used for 2005-08. The primary salinity was within ±0.0013 and the secondary was high by about 0.0035.

The primary sensors were used during 2005-16. The comparison was very noisy and they were found to be low by 0.0047.

The secondary sensors were used for 2005-21, -22 and -25. They produced salinity high by 0.0027, 0.0021 and 0.0032, respectively.

During 2005-23 there was a lot of scatter and the secondary sensors were found to be high by 0.0012.

- The DO sensor has been used for 2006-08 and 2006-11 since its latest calibration. The first was a very complex processing job as an error was made in the identification of the sensor so calibration was done twice. However, an analysis of one cast shows that the relationship was something like: 	CTD-BOT = 1.08 DOX-CTD - 0.005   (2006-08)

And the results of 2006-11 were: CTD-BOT = 1.0663 * DOX-CTD + 0.0783

The same CTD sensor was also used during Leg 1 of this cruise, but the sampling was very                  limited.

Historic ranges – Local climatology was only available for a few casts. The only data that fell outside the historic ranges for those casts were around 10 to 20m in high gradient zones; these excursions are not believed to indicate problems with equipment.

##### 13. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were on the order of 1s. A shift of +24 records (1s) is the shift that has been used in most other cruises. All data was shifted by +24records. (Output: SHFFL)

Dissolved Oxygen

Tests were run using +80, +100, +110 and +120 records on 3 casts. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by +110 records (about 5s). The choice of +110 records was found best for 2006-08 and 2006-11. After this step a few casts were plotted and the results are good; the offset between downcast and upcast features in DO looks much like those for temperature. As noted for 2006-11 there is little discernable difference between down and up DO values other than accounted for by vertical separation.

Conductivity

Tests were run on two casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.2 for both sensors. Those settings were found best for both sensors during 2006-08 and for the primary during 2006-11, but the secondary looked best with no change during 2006-11.

Both conductivity channels were advanced by -0.2. (Output: *.SHFC).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salinity

Maximum Surface Pressure (relative): 10.00 Minimum Salinity: 5

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors are closer to the bottles and were used for 2006-08 and -11. There is more history for the secondary sensors and they have produced salinity that is consistently high, but the differences varied from 0.0012 to 0.0035. Both channels contain spikes.

Page plots were produced and used to guide the editing. On-screen plots of descent rate and pump status were also used. All casts required only light editing except for cast #52 which required no editing.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

The salinity does not require recalibration, but the dissolved oxygen does. From section 11 we have the following equation for recalibration of DOX:

DOX-BOT = 1.0606 * DOX-CTD + 0.1138

File 2006-10-recal1.ccf was prepared to apply the above correction to the DO channel in the SAM and MRG files. COMPARE was rerun for dissolved oxygen and when the same outliers were left out of the comparison as in the first run, the average difference is -0.000008ml/l indicating that the recalibration worked properly; there is little variation with pressure and time. (See 2006-10-dox-comp2.xls.)

The edited downcast files, EDT, were recalibrated using 2006-10-recal1.ccf. (Output:COR1)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time, and the comparison with titrated samples those due to drift in calibration. But there remains an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values.

The FIL files were bin-averaged using 0.5db and then thinned to usual bottle levels and the CTD DO values were compared with the upcast bottle DO values. When plotted against pressure there was a lot of noise near the surface. When points above 12db and a few other outliers were excluded the following correction was derived from the trendline:

CTD DOX (Corrected) = CTD DOX - 0.000008*Pressure - 0.017 (See 2006-10-dox-comp3.xls.)

There is a lot of noise in the comparison so results depend on what is excluded, but in general the pressure variation is negligible. Below 20db all points are within 0.2ml/l. (Above 20m the data comes from an area of sharp gradient so that the match of bottle depth with CTD is critical and large errors are to be expected.) The offset does not appear to be due to incomplete flushing of bottles since there is so little pressure dependence. So it is assumed to be an error due to time response of the CTD sensor. While the error is small, it still seems worth fixing since it is consistent with depth. This is the smallest correction ever found for this sensor. During 2006-11 it was about 0.03ml/l.

File 2006-10-recal2.ccf was prepared to apply an offset of -0.017ml/l to the SBE DO channel in the thinned files prepared for the comparison and to the metre-averaged, edited downcast files. (Output: THN2 and COR2) COMPARE was rerun and the results were satisfactory. (See 2006-10-dox-comp4.xls.) Note that this correction is not applied to bottle files as it concerns only “in-motion” errors.

There is a lot of scatter in the comparison but excluding a few outliers suggests that the dissolved oxygen data in the CTD files should be considered ±0.6ml/l to 20m, ±0.2ml/l from 20m to 175m and ±0.1ml/l below 175m.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.6ml/l from    0 -  20m

•	±0.2ml/l from   20 - 175m

•	±0.1ml/l below 175m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~85% to 144%. Cast #52 had the highest saturation. The fluorescence was off-scale near the surface for that and many other casts. A surface bottle confirms that the dissolved oxygen values are ok, so the high saturation values are assumed to be due to intense biological activity.

##### Final Bottle Files

The MRGCLN files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### 21. Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun; no errors were detected.

The sensor history was updated for the CTD sensors.

Particulars from logs

14. Altimeter slow to find bottom

26. No water in Niskin bottle

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2006-10 LEG 2 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 4 June 2006 End: 12 June 2006 |  |  |  |  |  |  |
| Location: Broughton Inlet |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Stucchi D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2006-10 LEG 2

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 01/02/05 | Factory “ |  |  |
| Conductivity | 1766 | 07/06/05 | “ |  |  |
| Secondary Temp. | 2106 | 08/07/05 | “ |  |  |
| Secondary Cond. | 1729 | 12/07/05 | “ |  |  |
| Transmissometer | 723DR | 28/03/06 | IOS |  |  |
| SBE 43 DO sensor | 0766 | 24/11/05 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |