## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected Salinity:Bottle lost during addition of HPLC. S.H. |
| ### 13 January 2021 | ### Added HPLC Data. S.H. |
| 18 January 2018 | Changed channel name “Fluorescence:CDOM:URU:Wetlabs” to “Fluorescence_CDOM:URU:Wetlabs” in underway file. J.L. |
| 1 Sept. 2006 | Fixed DO in CHE files for casts 60 and 62. G.G. |

## PROCESSING NOTES

Cruise: 2006-11

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG - JdeF

Party Chief: Pena A.

Platform: Vector

Date: April 14, 2006 – April 21, 2006

Processed by: Germaine Gatien

Date of Processing: 12 June 2006 – 31 July 2006

Number of original CTD casts: 72 (first cast divided into 2 files)

Number of CTD casts processed: 70 (1 upcast only and 1 had pumps off)

Number of rosette casts: 30	Number of rosette casts processed: 30

Number of TSG files: 1		Number of TSG files processed: 1

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#723DR), a PAR sensor (#4565), a surface PAR sensor (#16504) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#0766) was mounted on the primary pump and a Seapoint Fluorometer (#2356) with a 10X cable was mounted on the secondary pump. The deck unit was a model 911 (#0425). The salinometer was an 8400B model Autosal (S/N 68572).

There was a thermosalinograph (S/N 2487) mounted with a Wetlab Wetstar fluorometer (#WS3S-889P) and a Wetlab CDOM fluorometer.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book, rosette sheets and bridge log were available. The log book contained neither a record of the serial number of the thermosalinograph nor which sensors were mounted on it.

The record of loop sampling in the log book was excellent for the first few days with exact times of sampling, but late in the cruise times were occasionally either missing or the time zone was wrong and two loop sample numbers were used twice.

There were some errors in nutrient sample labels for one cast. The rosette log sheets are believed to be correct.

There were many errors in file names and station names in the data files. The log records about this were unclear but the chief scientist was aware of the problem and had a list of most of the errors and what the correct entry should be. In many cases the station name given is the same as the cast number. There is always room for confusion with this project because the station names are numbers that mostly overlap with the event numbers. A letter placed before or after the station name might help. Adding to the problem is the fact that the sample numbers were not always recorded in the Daily Log. The extra information in the Daily Log frequently helps sort out confusion over file names and sample numbers.

The SBE dissolved oxygen sensor performed very well, resolving sharp gradients better than in the past and needing only a small correction. Surface saturation was unusually high in Saanich Inlet (~130%) during both the downcast and upcast. A bottle confirms the dissolved oxygen values are correct, so this is assumed to due to intense biological activity.

The SBE dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l in the top 150m

- ±0.1ml/l below 150m

The stops for rosette sampling lasted 30s or more.

There appears to be an offset on the order of 1 mg/m^3 in the data from the Thermosalinograph fluorometer compared to both titrated chlorophyll from loop samples and the CTD fluorometer data.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log Book, Bridge Log and Rosette Log were obtained. There was a loop sample taken at the start of every shift, and the time was noted precisely. This is very useful for calibration.

Niskin Bottle #13 did not close on many occasions.

There were errors in file names for many of the early casts. The problem was not recorded in the log book, though there was a note that there was no event #12 and no file #16 which helped track down what had happened. A few errors were found and corrected early, but some were not fixed until conversion to IOS HEADERS.

Corrections were made to a few station names (casts 47 and 82) based on notes in the log book.

The nutrient, chlorophyll and salinity data were each obtained in spreadsheet format. Each spreadsheet was edited to remove information not needed for the chemistry files. There were flag and comment channels for all, but the only salinity value flagged had no comment entered. That information was obtained from the salinity analysis sheet and entered into the simplified spreadsheet. The salinity data was divided into 2 spreadsheets for loop and rosette samples.

Individual dissolved oxygen bottle files were obtained; flag channels and comments were included.

The CTD calibrations were checked and no errors found. The pressure offset is the same one used for 2006-08.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

The files were converted to CNV files. A few files were checked and the pressures are never less than 0.9db.

There were a few casts with non-standard names (an extra zero in the event number); those were corrected after conversion.

The file names were corrected for event numbers 7, 9, 13, 14, 15, 16 which were originally named 6, 7, 12, 13, 14 and 15. (Other errors were discovered later – see section 9.)

As usual for this area the descent rate was highly variable; it was very steady for some casts, very noisy and fairly low for others. Bottle stops typically lasted just over 30s.

The down and upcast temperature and conductivity channels are reasonably close.

Transmissivity looked fine with good agreement between downcast and upcast.

The dissolved oxygen looks better than usual with upcast and downcast offset vertically, but without the sort of hysteresis that is usually observed.

The fluorescence looks ok.

The altimetry is often extremely noisy but generally looks reliable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s. The names of files 6, 7 and 13 were changed to 7, 9 and 14.

All files were then converted to IOS HEADER format. The extensions were changed to BOT. All BOT files were plotted and a few significant outliers were found in the secondary salinity for casts #22, 23 and 60. These were cleaned using CTDEDIT and the output copied to BOT.

Sample Loop22 was entered in the log book twice; there is a note in the log saying that the second one should have been Loop23. In the bottle salinity file there is only a Loop 22. It will be assumed that this is correctly labelled.

There are a number of errors in the rosette sheets and Daily Log concerning sample numbers. The sample numbers are not always entered in the log, and the cast number is wrong or missing on some rosette sheets. Extra care will be needed to ensure the sample numbers go with the right data.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Two casts with a steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7) and (0.03, 9) were tried and there was little difference between several of them. There are significant differences between the upcast and downcast data so that the judgment of the optimal setting is difficult. A choice of (0.03, 9) is probably best for both channels. CELLTM was run using (0.03, 9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 11 | 400 | -0.0027 | +0.00002 | +0.0025 | Moderate, steady |
| 35 | 350 | -0.0027 | +0.00007 | +0.0016 | Moderate, steady |
| 87 | 290 | -0.0026 | +0.00003 | +0.002 | Moderate, steady |

The temperature differences are much larger than usual even when the descent rate was very steady. Cast #57, though shallow, was extremely well-mixed so it was examined closely. Most noise in the temperature differences occurred at large gradient areas where the secondary temperature was spiky. Similarly all noise in conductivity and salinity differences is associated with noise in the secondary channels.

There is no evidence of temporal drift in the differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header summary was checked carefully against the log book entries and more errors were found. The station name is wrong for casts #47, 73, 81, 82 and 84 and the file name was wrong for casts #24, 26, 45, 62, 64, 65 and 67. These were corrected in the IOS files and BOT files.

The header check was run and no further errors were found. Track plots were produced and added to the end of this report.

The average surface pressure is ~1.9db which is reasonable for the Vector.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

As mentioned in the log book the pumps were not on for cast #10, so it will not be processed further.

##### BOTTLE FILE PREPARATION

Because there were many inconsistencies over event numbers care was taken in checking the bottle files to be sure they are correctly identified.

The salinity and dissolved oxygen bottle files for the first cast were named #1, but the cast was divided into #1 (downcast) and #2 (upcast) so the bottle files were renamed as #2. The chlorophyll file was already named properly.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. When bottles did not fire, the bottle number gets out of step with the record on the log sheet. The addsamp file was edited to fix this (addsamp2.csv), but there is room for error and care will have to be taken to ensure the matches are done properly, especially given that there are errors in the sheets. The header names were changed to standard format, lines with pad values for sample #s were removed and the file was then converted to CST files to form the framework for the bottle files.

Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number. Because sample #184 was used twice the number was changed at sea to #9184 for cast #40. (There does not appear to have been any sampling from #184 during cast #39.)

The salinity data was provided in EXCEL format with flag and comment channels. No data had been flagged. Channel names were changed to standard format and the file was saved as 2006-11sal.csv. There were 5 cases for which the sample # does not correspond to sampling given on the rosette log sheets. In some cases there were 2 bottles fired at a single depth, so it does not matter, but in most cases the difference involved sampling from different depths, so all should be re-examined after running COMPARE.

For now, the rosette log is taken as correct for the following:

- For cast #17 the sample named #62 was changed to 60 to match the rosette log record.

- There is a salinity sample said to be #107 from cast #28. Sample #107 is said to be from cast #26 in the log book, but from cast #28 according to the rosette log sheet. The salinity and dissolved oxygen data from this sample look like they are from the first bottle of cast #28, but the chlorophyll is clearly a surface value so probably from cast #26. The sample number for cast #26 will be changed to #9107 which will only affect the chlorophyll and any other optical sampling. The sample numbers for cast #28 will be used in accordance with the rosette log as this causes the least confusion.

- For cast #28 there is a salinity value said to be from sample #119, but according to the rosette log sheet that bottle did not fire so sampling was from #120 (bottle14) at the same depth. This was changed in 2006-11sal.csv.

- For cast #39 the sample named #180 was changed to 181 to match the rosette log record.

- For cast #65 the sample named #268 was changed to 270 to match the rosette log record.

- For cast #71 the sample named #290 was changed to 293 to match the rosette log record.

The files were converted to individual SAL files.

The extracted chlorophyll data was in spreadsheet format and included a flag channel and comments. The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The resulting file is 2006-11-chl.csv. That spreadsheet was converted to individual files. The following comments from the analysis sheet was added to the comments that will later be added to the headers of the bottle files: “Average of two samples is reported unless stated otherwise. Variability is assessed as the % (std dev/mean*100).”  (Output: CHL)

The dissolved oxygen data was received in individual files with flag channel and comments.

The nutrient data was provided in spreadsheet format with flag channel and comments. Channel names were changed to standard format and records with no sampling were removed; the file was saved as 2006-11-nuts.csv. The data was reordered on sample number and then converted to individual files.

Note that sample #336 from cast #80 looked odd with DO and nutrient values out of line with those above and below it. On a TS curve it is clear that there was a stable intrusion at about 100m of warm, high DO water, and sample #336 would have been in the middle of the feature.

The SAL, CHL, ADD and NUTS files were merged with the CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4.)

Those files were then merged with SAMAVG files (MRG) and then put through CLEAN to reduce the headers to File and Comments only, remove remarks and keep comments, remove SeaBird headers and comments from the secondary files. (MRGCLN)

NOTE: It was discovered later that the sample numbers from cast #19 were probably incorrect. The rosette log sheet indicates that the last 4 samples were taken from Niskin bottles #11, 13, 15 and 17 which correspond to samples 74, 76, 78 and 80. The sample bottles for the last 4 nutrient samples were labelled #73, 74, 75 and 76. The sample numbers were changed, flag “c” entered for the values and a not of explanation put in the header. Melanie Quenneville investigated this and concurred with these conclusions.

##### 11. COMPARE

Salinity

COMPARE was run and problems were noted as follows:

- Sample #314, cast #77, Niskin #1: Salinity severe outlier and it looks like DO is bad too. Both look more like water from about 120db rather than 177db. Does not look like they were drawn from another bottle, more likely the bottle misfired. Salinity flagged “d” and message of warning sent to Wendy Richardson.

- Samples #385 and 400, cast #88, Niskins #1 and 16: Salinity values appear to be reversed. Changed in SAL file, but flagged “c”. DO looks ok, but there is little variation in DO.

When all bottles from below 50m were included the primary salinity was found to be high by about 0.008 but there was a lot of variability and the trendline was not flat. When outliers (differences >0.2) were removed a flat trendline was achieved and the primary was low by 0.0008; there were only 12 bottles in the comparison. Similarly, the secondary was high by 0.011 when all bottles below 50m were used, but when differences >0.02 were excluded it was found to be high by 0.0014 and the trendline was quite flat. The same bottles showed up as outliers for both channels, which suggests that they reflect problems with the bottle samples rather than problems in either pair of CTD sensors.

All four of these outliers are from Niskin #1 and were taken close to the bottom. It has been noted on other cruises in the Strait of Georgia that such samples are often not useful for salinity calibration. This may be due to turbidity due to a soft bottom. In one of the four outliers the transmissivity shows that turbidity could well be an issue. For the others there is no evidence of that, though it is still possible that mud can be an issue even when the transmissivity measured by the CTD is not particularly low. There may also be more salinity variability at the bottom. For whatever reason it is preferable to take samples at least 10m from the bottom. All the samples that were used in the comparison were from the bottom, raising some doubts about the results.

The four salinity outliers were flagged “c” and a note of explanation placed in the headers.

Dissolved Oxygen

COMPARE was run. Plotting the differences versus CTD DO and pressure excluding any samples flagged “d”, 10 outliers were identified by removing differences >2 and residuals <-.4 and >0.4. This produced a fairly tight relationship for both plots. This is unusual as the fit against pressure has been poor in the past. The only explanation for this may be that the DO sensor encountered no waters with DO <1.8ml/l. The following fits were found:

DOX-BOT = 1.0663 * DOX-CTD + 0.0783

DOX-BOT = DOX-CTD -0.0006 * Pressure +0.4532

The fit versus time shows no significant temporal drift. The fit against CTD DO is the tightest.

The following outliers were flagged “c”: #40 (cast 14), #76 (cast 19), #117 (cast28)

The following outliers were flagged “d”: samples #77 (314), #116 (cast 28) and #359 (cast 84).

Notes of explanation were place in the headers.

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. A few slight outliers were investigated, but there was nothing to suggest that further flagging of bottle DO data was justified.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. When all data is included the fluorescence is about 72% of the titrated chlorophyll values. When values with CHL>9 and 2 other outliers are excluded it is about 83%.

##### Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors were used for 2005-08 when the primary salinity was within ±0.0013 and the secondary was high by about 0.0035. The primary sensors were used during 2005-16 when the comparison was very noisy and they were found to be low by 0.0047. The secondary sensors were used for 2005-21, -22 and -25 when they produced salinity high by 0.0027, 0.0021 and 0.0032. During 2005-23 there was a lot of scatter and the secondary sensors were found to be high by 0.0012.

- Only one cruise has been processed that used this DO sensor since its latest calibration and it was a very complex processing job as an error was made in the identification of the sensor so calibration was done twice. However, an analysis of one cast shows that the relationship was something like::

CTD-BOT = 1.08 DOX-CTD - 0.005   (2006-08)

Historic ranges – The only excursion from the historic ranges in the Strait of Georgia was the deep salinity for cast #47 which was a little low, but this probably reflects inadequacies in the climatology. In Juan de Fuca Strait there were many low salinity values between 70 and 100m particularly close to the mouth. This is believed to reflect a real intrusion of cold, fresh water. There are also some low salinity values near the bottom for some of the Juan de Fuca casts. These excursions are not considered indicative of instrument calibration problems.

##### 13. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were on the order of 1s. A shift of +24 records (1s) is the shift that has been used in most other cruises. All data was shifted by +24records. (Output: SHFFL)

Dissolved Oxygen

Tests were run using +80, +100, +120 and +140 records on one cast. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by +100 or +120 records (about 4 to 5s). Since a choice of +110 records was found best for 2006-08 the same value was applied to the DO channel for all casts. After this step a few casts were plotted and the results are good; the offset between downcast and upcast features in DO looks much like those for temperature. There is no discernable difference between down and up DO values other than accounted for by vertical separation.

Conductivity

Tests were run on one cast with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.2 for the primary sensor which is the same result as found during 2006-08. The secondary conductivity looks best with no change, so no shift will be applied. The primary conductivity channel was advanced by -0.2. (Output: *.SHFC).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salinity

Maximum Surface Pressure (relative): 10.00

Minimum Salinity: 5

Surface Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #2 which was an upcast only.

##### DETAILED EDITING

The primary sensors are closer to the bottles and were also best during 2006-08. There is more history for the secondary sensors and they have produced salinity that is consistently high, but the differences varied from 0.0012 to 0.0035. There is a lot of noise in the secondary temperature and salinity.

Page plots were produced and used to guide the editing. On-screen plots of descent rate and pump status were also used.

The following casts required no editing: 41, 56, 60.

The following casts required heavy editing: 79, 80, 81, 82, 84.

All other casts required only lightly editing.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

The salinity does not require recalibration, but the dissolved oxygen does. From section 11 we have the following equation for recalibration of DOX:

DOX-BOT = 1.0663 * DOX-CTD + 0.0783

File 2006-11-recal1.ccf was prepared to apply the above correction to the DO channel in the SAM and MRG files. COMPARE was rerun for dissolved oxygen and when the same points were left out of the comparison as in the first run, the average difference is -0.00006ml/l indicating that the recalibration worked properly; there is little variation with pressure and time. (See 2006-11-dox-comp2.xls.)

The edited downcast files, EDT, were recalibrated using 2006-11-recal1.ccf. (Output:COR1)

The COR1 files were clipped to 100db and set aside to be processed later for Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time, and the comparison with titrated samples those due to drift in calibration. But there remains an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values.

The FIL files were bin-averaged using 0.5db and then thinned to usual bottle levels and the CTD DO values were compared with the upcast bottle DO values. When plotted against pressure and removing outliers identified by residuals, the following correction was derived from the trendline:

CTD DOX (Corrected) = CTD DOX - 0.000009*Pressure - 0.03 (See 2006-11-dox-comp3.xls.)

There is a lot of noise in the comparison so results depend on what is excluded, but in general the pressure variation is negligible and the offset around 0.03ml/l. The offset does not appear to be due to incomplete flushing of bottles since there is so little pressure dependence. So it is assumed to be an error due to time response of the CTD sensor. While the error is small, it still seems worth fixing since it is very consistent with depth. This is the smallest correction ever found for this sensor.

File 2006-11-recal2.ccf was prepared to apply an offset of -0.03ml/l to the SBE DO channel in the metre-averaged, edited downcast files. (Output: COR2) COMPARE was rerun and the results were satisfactory. (See 2006-11-dox-comp4.xls.) Note that this correction is not applied to bottle files as it concerns only “in-motion” errors.

There is a lot of scatter in the comparison but excluding a few outliers suggests that the dissolved oxygen data in the CTD files should be considered ±0.2ml/l to 150m and ±0.1ml/l below that.

##### Special Fluorometer Processing

All COR1 files were put through CLIP to produce files with data to 100db only for the use of A. Peña.

Those files were then bin-averaged (1/4db bins), put through REMOVE to remove extraneous channels, HEADEDIT to fix formats and channel names and recalibrated using 2006-11-recal2.ccf. The final files (FCTD) were saved in a separate directory. A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data was put through a median filter with fixed width 11, before bin-averaging. The recalibrated CTD bottle files (SAMCOR1) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.2ml/l from   0 - 150m

•	±0.1ml/l below 150m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~89% to 115% for all casts except the first one in Saanich Inlet. For that cast downcast saturation was ~136% and the upcast was ~126%. A surface bottle confirms that the dissolved oxygen values are ok, so the high saturation values are assumed to be due to intense biological activity.

##### Final Bottle Files

The MRGCLN files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

As explained in section 10 a probable error in nutrient sample numbers was discovered in the CHE file for cast #19. The sample numbers 73, 74, 75 and 76 were changed to 74, 76, 78 and 80 for nutrients only. Flags (“c”) were assigned to these values and a note of explanation entered into the header of the CHE file and the NUTS file.

##### 22. Thermosalinograph Data (Note: there was no intake temperature channel)

a.) Checking calibrations

There was 1 file containing TSG data. The con files had an error in the temperature calibration. The serial numbers are missing from the two fluorometers and they both have the same calibration constants, but one is identified as a CDOM. The CDOM fluorometer information was obtained from Jim Gower. For the other, the calibrations look like those for WS3S-889P, and Doug Anderson confirms that was the instrument used.

The value of the CDOM fluorometer in pure water was given as 0.7V in a factory calibration from 13 March 2002, and a concentration of 100ug/l was said to correspond to a voltage of 3.0V. From those values the scale factor was calculated as follows:

Scale factor = [CDOM]standard / (Vstandard - Vblank)

= 100/(3-0.07)

= 34.1297

So for the CDOM the scale factor was entered as 34.13 and the offset as 0.07. The corrected file was saved as 2006-11-TSG.con.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Salinity, Fluorescence:Wetlab:Weststar, Fluorescence:Wetlab:CDOM, Julian Days, Latitude, Longitude. Before this file could be converted to IOS HEADER format, the name of the CDOM channel was changed to Upoly1 since the CDOM name was not in the SeaBird parameter file.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

A time-series plot was produced and the only spikes noted were small ones.

An initial track plot was produced and showed that there was a bad position in one record. CTDEDIT was used to interpolate position for that record. A second pass through CTDEDIT was used to examine temperature, salinity and fluorescence (the Wetstar channel, not the CDOM); temperature and salinity were both cleaned in a single record and salinity was cleaned in 1 other area. A third pass through CTDEDIT was used to clean the CDOM channel for the single record that also had bad T and S.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or near 3.3db and exported to a spreadsheet. The TSG files were opened in EXCEL and reduced to the times when CTDs were run. (The TSG was not recording when the first cast was run.) Those files were combined in a spreadsheet matching times. The positions were compared and most were very close, with differences usually less than 0.0005º for both latitude and longitude. However, there was a segment between casts #47 and 57 in which the differences were quite a bit larger. This is in an area of large surface currents and it is believed that a small mismatch in time could produce a large difference in positions. The good correspondence in other regions supports the conclusion that the clock and GPS worked well. This spreadsheet will be used in the next step to compare temperature, salinity and fluorescence. (See 2006-11-tsg-ctd-comp.xls)

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 There was no intake temperature channel available.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. The salinity and fluorescence had been averaged by taking the median value over a 2 minute window. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. The plot of TSG temperature versus CTD temperature was created using all the casts available and a trendline was fitted. The offset of the trendline was adjusted until the slope was ~1 and the offset required was -0.22Cº. There were a lot of outliers and they were mostly on one side of the trendline. A subset of the CTD casts was chosen as well-mixed based on a mixed-layer depth calculated to be >4.5db. Two outliers were removed for temperature, salinity and fluorescence (not the same casts for each) leaving 32 points of comparison for each. (See 2006-11-ctd-tag-comp.xls and 2006-11-tsg-ctd-well-mixed.xls.)

The TSG fluorescence is higher than that of the CTD by an average of 2.2 with values of from 1.5 to 3.3 including outliers, but the trendline suggests a slope of about 1.3 and offset of ~1 when all data is used. When only casts that are well-mixed near the surface are included the fit is tighter and the slope is 1.5 with an offset of ~0.9.

The TSG salinity is lower than that of the CTD by an average of 0.019 based on the median of the well-mixed casts. The average of those values shows the TSG low by 0.03.

The lab temperature is higher than the CTD temperature by an average of 0.198 Cº and the median is 0.195Cº using just the well-mixed casts.

No time-dependence is noted but the scatter in the results combined with the large geographic variability, mean it is unlikely we would see any even if it existed.

e.) Loop Salinity and chlorophyll comparison

There were 19 loop salinity samples and 27 chlorophyll samples. A spreadsheet (2006-11-TSG-chl-sal-comp.xls) was prepared with the loop sample ID, the salinity and chlorophyll from the analysis spreadsheet and the date and time of collection as noted in the log book. There were problems with a number of the samples.

- Loop19 – Time in log looks like 12:44 but is between events at 19:25 and 20:13, probably time is PDT, so should be 19:44UTC

- Loop 22 – two samples with that label. On nutrient labels 22a and 22b were used. The times in the log for both samples appear to be in PDT.

- Loop 25 – no time in log – between events at 18:08 and 19:29, so time was guessed to be 19:00.

- Loop 26 – two samples with that label. On nutrient labels 26a and 26b were used.

The salinity and fluorescence data in the TSG file were averaged by doing a running median and the standard deviation found over 2 minutes (5 samples); the average value and standard deviation were found for the times of the loop samples and that data added to the loop analysis results. The TSG salinity is lower than the bottles by 0.017 (median 0.027) when all samples are included except for two extreme outliers. When only the 7 samples with the lowest standard deviations in TSG salinity are included it is low by an average of 0.021 (median 0.025). The difference from the one comparison with the least variability in TSG data was 0.021. (See file 2006-11-TSG-loop-comp.xls)

The TSG fluorescence was higher than the titrated chlorophyll by a factor of 2.5 if all data is included and 2.2 when values are excluded because of high variability in the fluorescence. The median of those casts with lower variability is 1.7. When plotted the trendline has a slope of ~1.1 with an offset of 1.3.

f.) Study of CDOM fluorescence

The two fluorescence channels were plotted together, but had very different patterns. The CDOM signal was insignificant in most areas, and the only noteworthy relationship noted is that it was high when salinity was low, in the area near the Fraser River. The chief scientist believes this is a reasonable result suggesting that the equipment was operating properly.

g.) Calibration history

The TSG was recalibrated in October 2005.During 2006-08 the temperature rise in the loop due to heating by the ship was found to be 0.22Cº but that was in February and on the TULLY. It is expected that there would be less heating on the Vector, though there is no history of use on this ship. Also the ambient waters are warmer so that should also lead to less change between intake and lab. The salinity was found to be low by 0.020.

h.) Conclusions

Temperature: The TSG temperature is higher than that of the CTD by about 0.20Cº. When last used on the Tully it was high by 0.22Cº. The ambient waters were cooler for the earlier cruise and the Tully is larger, both facts leading us to expect more heating in the loop, so the result looks reasonable.

Salinity: The TSG salinity was found to be lower than the loop samples by from 0.017 to 0.027 and lower than the CTD by about 0.019. Choosing a value of 0.22 is a reasonable compromise. When the TSG is next calibrated this should be re-examined.

Fluorescence: Both the comparison with CTD and loop samples suggests there may be an offset of about 1 in the TSG fluorometer. If we assume that is the case, then the TSG fluorescence is approximately 1.5 times that of the fluorometer on the CTD and 1.1 times the loop chlorophyll. No calibration will be applied to the fluorescence.

i.) Recalibration

File 2006-11-TSG-recal.ccf was used to apply offsets of -0.20 Cº and +0.022 to Temperature:Primary and Salinity:T0:C0.

j.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary and Flag.

HEADER EDIT was used to change the channel name Upoly1 to Fluorescence_CDOM:URU:Wetlabs, the DATA TYPE to THERMOSALINOGRAPH and to add the depth of sampling to the header and to add a general comment about the sampling method as well as the following remark:

Fluorescence is nominal. Temperature and salinity have been recalibrated

based on comparison with CTD data, salinity samples from the loop and the history of the sensors.

As a final check plots were made of the cruise track and data; no problems were noted.

##### 23. Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun; no errors were detected.

The sensor history was updated for the TSG and CTD sensors.

Particulars from logs

1/2. Cast stopped at bottom and started over.

8. Some problems with fluorescence sampling.

10. Pumps not on. Cast repeated.

12-15. Files were named wrong and should be 13 through 16. File names were changed after conversion.

12. Error message in WILDEDIT run (“required Header END line not found”); stray symbol in header removed and then program ran on this cast.

26/28. Bottle #5 cast #26 was originally named sample #107 but this was changed to #9107 to avoid confusion with bottle #1, cast #28 which was also named sampled #107.

38. Error message in SeaBird conversion but data looks fine. (“Unexpected end of file”)

39/41. Sample #184 was assigned twice so it was renamed as 9134 for cast #41. However, there was no sampling corresponding to #184, cast #39, so this was not really necessary.

47. Station name should be 57 (was entered as 47, was corrected after Sea-Bird conversion.)

51/52. Very well mixed.

78. Log notes that the program crashed after pump shut down. File looks fine.

82. Station name should be 73 (was entered as 82, was corrected after Sea-Bird conversion.)

# Institute of Ocean Sciences

CRUISE SUMMARY

| Cruise ID#: 2006-11 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 14 April 2006 End: 21 April 2006 |  |  |  |  |  |  |
| Location: SoG/JdeF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Pena A. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2006-11

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 01/02/05 | Factory “ |  |  |
| Conductivity | 1766 | 07/06/05 | “ |  |  |
| Secondary Temp. | 2106 | 08/07/05 | “ |  |  |
| Secondary Cond. | 1729 | 12/07/05 | “ |  |  |
| Transmissometer | 723DR | 28/03/06 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 24/11/05 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| PAR | 4565 | 23/12/04 |  |  |  |
| Surface PAR | 16504 | 02/01/04 |  |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2006-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 21/10/05 | Factory | 21/10/2005 |  |
| Conductivity | 2487 | 21/10/05 | “ | 21/10/2005 |  |
| Wetlab WetstarFluorometer | WS3S-713P | 10/01/01 | ? |  |  |
| Fluorometer, Wetlab CDOM | WSCD-817P | 13/03/2002 | ? |  |  |