## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2006-14

Agency: OSD

Location: North-East Pacific

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: September 12, 2006 – September 20 2006

Processed by: Germaine Gatien

Date of Processing: 30 January 2007 – 14 February 2007

Number of original CTD casts: 55	Number of CTD casts processed: 55

Number of bottle casts: 29		Number of bottle casts processed: 29

TSG files: 1

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea/Seatech transmissometer (#723DR), an SBE 43 DO sensor (#0047, on the primary pump) and a Seapoint Fluorometer (#2228) with a 10X cable (on the secondary pump). The deck unit was a model 911 (#0508) and the logging computer was PAC02328. A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P).

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order. There was confusion over sample numbers for one cast.

The salinity data are unusually spiky even when the descent rate was high and steady, suggesting problems with the flow rate. Most of the spikiness is 2-sided so errors should be minimized by metre-averaging. The differences between salinity channels are pressure dependent but those between the CTD and bottle salinity are much less so. This may also be evidence of pump/flow rate problems.

The thermosalinograph calibration was confusing. As has been observed during some other recent cruises the differences between the TSG and loop samples varied in different parts of the cruise, and the salinity differences are lower than during cruises before and after this one. The absence of the intake thermistor was unfortunate as it might help explain some of these observations.

The dissolved oxygen data in the CTD files are considered

- ±0.8ml/l from 0 –100db

- ±0.2ml/l from 100 –200db

- ±0.1ml/l from 200 – 1200db

- data below 1200db are considered unreliable by the manufacturer.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. There is no mention of a PAR sensor in the equipment list in the front of the book, but mention is made in the log when it was put on and off.

Titrated chlorophyll, nutrients and salinity data were obtained in spreadsheet format with flag channels.

The oxygen files were provided in individual OXY files which lacked flag and comment channels.

A note in the file mentions confusion over sample numbers for cast #23. The bottle files were all checked to ensure that they were in agreement and they all used samples #152 through 157 for that cast.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and errors were corrected in the CTD secondary conductivity and the TSG temperature and a pressure offset of +1.2db was added for the CTD; the fluorometer serial number was added to the TSG configuration file. The files were saved as 2006-29-ctd.con and 2006-29-tsg.con.

File 2006-29-ctd-par was prepared for the CTD casts with a PAR sensor.

##### Conversion of Raw Data

Data were converted using 2006-14-ctd-par for casts 2-9, 29-40, 53-57 and file 2006-14-ctd.con for all other casts.

A few casts were examined and all expected channels are present and look reasonable when plotted. The fluorescence seems less noisy than usual.

A note in the log mentions an error in the longitude for cast #4; the error is in the log record with 123º given where it should be 126º, but the header of the file is correct.

Rosette files were converted using a start time of -2s and duration of 5s.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT.

All BOT files were plotted and a few outliers were found in secondary salinity in cast #17. CTDEDIT was used to clean the salinity in that file and the results copied to 2006-14-0017.BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

This settings found to work best for CELLTM during 2006-24 before this cruise and 2006-15 and 2006-33 afterwards were all the same, namely (0.02, 7) for the primary and (0.03, 7) for the secondary conductivity. A test was run on two casts to see if the same settings look reasonable for this cruise and they do. CELLTM was run using (0.02, 7) for the primary and (0.03, 7) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on some deep casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 11 | 1480 | ~0 | -0.0005 | -0.006 | V.High, noisy |
| 44 | 500 |  | -0.0001 | -0.0005 | V.High, noisy |
| 44 | 1000 |  | -0.0002 | -0.0025 | V.High, noisy |
| 44 | 1480 | ~0 | -0.00043 | -0.0050 | V.High, noisy |
| 44 | 2000 |  | -0.00045 | -0.0055 | V.High, noisy |
| 69 | 1480 | ~0 | -0.0005 | -0.006 | V.High, noisy |

The differences are similar to those seen during 2006-15; the conductivity and salinity differences increase with depth. Cast #44 is particularly strange with a sudden shift in conductivity differences around 300db and 1100db with differences lower between those two levels than above and below.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run and no errors were found. The cruise track was plotted and no problems found. The average surface pressure is 3db, which is reasonable for the Tully

The altimeter readings from the header were exported to a spreadsheet and a few casts were checked and show that the algorithm worked well, though the altimetry itself is unusually noisy at the bottom near the end of the cruise.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

The salinity spreadsheet included loop data and rosette data. The file was edited to change headers to standard names and then saved in two files, one containing rosette data, 2006-14sal.csv, and the other loop data, 2006-14sal-loop.csv. The rosette file was then sorted on sample number and converted to individual SAL files.

Dissolved oxygen files (*.oxy) were provided. A flag channel was added using the IOS SHELL routine ADD CHANNEL. There were notes on the rosette sheets indicating problems with a few of the titrations, so “c” flags were added to those samples and the analyst’s notes were added to the header in the output ADD files.

The nutrient spreadsheet was simplified and saved as 2006-14nuts.csv. Extraneous columns were removed, header names were changed to standard format and the data were sorted on sample number. File 2006-14nuts .csv was then converted to NUT files.

Extracted chlorophyll data were obtained and saved as 2006-14CHL.csv. The file was edited to remove extraneous lines and columns, header names were changed to standard format, data were sorted on sample number and loop data were saved separately as 2006-14CHL-loop.csv. 2006-14CHL.csv was converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

DO was plotted against CTD Salinity. The only outlier was near the bottom for cast #5.

##### 12. COMPARE

Salinity

COMPARE was run with sample number as the reference channel. When data above 250db were excluded the differences between bottles and CTD were quite flat for both pairs of sensors. In section 7 it was noted that the differences between the two sensors did increase with pressure. If we estimate the differences between the primary and secondary salinity based on the trendlines in COMPARE then we find:

| Pressure | Primary Sal – Secondary Sal |
| --- | --- |
| 0db | 0.0040 |
| 500db | 0.0047 |
| 1000db | 0.0053 |
| 1500db | 0.0060 |
| 2000db | 0.0066 |

The pressure dependence is significantly less than noted in section 7. Whatever accounts for the much lower differences above 1000db is apparently not due to a faulty cell. Pump problems are a possible explanation.

The only outlier of note was sample #36 from cast #8. This was a near-bottom sample and those often give anomalous results. Looking carefully at the record of that cast the CTD was moving around a lot and the salinity was varying somewhat, but not by the 0.02 to 0.03 required to explain the differences from the bottle value. The bottle value will be flagged “c” with a note of explanation in the header. The other Niskin samples look ok. There is no suggestion of variability with time.

The primary salinity is high by an average of 0.0067psu and the secondary is high by an average of 0.0019psu.

Dissolved Oxygen

COMPARE was run using sample # as the reference variable. The fit of differences versus Bottle DO is not tight, seeming to have two general groups. This suggests that there may be some casts with drifting calibration due to anoxic sampling. From about cast #15 onwards there is no suggestion of this. It is known that when this sensor was returned to the factory the membrane was replaced. One sign of membrane trouble is an inability to deal well with anoxia. So the plug placed on the sensor in storage, and likely also present on the long steam to the 2nd cast has probably resulted in calibration problems.

The best fit was with the differences between bottles and CTD versus CTD DOX and dividing the casts into groups (#1-15 and #17-end). First, data from below 1250db were excluded from the fit and then outliers were excluded based on residuals; the fit found for casts 1 through 15 was:

CTD-BOT = 1.1341 DOX-CTD - 0.0044

And for casts 17 to the end:

CTD-BOT = 1.0341 DOX-CTD + 0.0425

Many combinations were considered to see if the fit should be divided into smaller or different ranges, but there did seem a distinct shift between #15 and #17. (See 2006-14-dox-comp1.xls.)

The flags originally entered were re-examined and one flag was added after COMPARE:

Cast #2, Sample 12 – flag “c” left unchanged as it is an outlier in COMPARE.

Cast #2, Sample 17 - flag “c” left unchanged though only a slight outlier when the early casts were plotted together.

Cast #5, Sample #27 – flag “d” added since the point was a severe outlier in COMPARE and an outlier in the DO versus salinity plot.

Cast #19, Sample #135 – flag “c” changed to “d” – severe outlier in COMPARE.

Fluorescence

COMPARE was run using the CTD Fluorescence and the Titrated Chlorophyll from bottles. When all data were plotted there was a lot of scatter with an average ratio of 1.4 for CHL/FL, but when the ratio was plotted against cast number it became clear that the ratio was on the order of 0.9 well offshore and 1.8 nearer shore, with values >2 very close to shore. Put another way the fluorescence is about 55% of the chlorophyll fairly close to shore and 110% well offshore. It is often observed that the ratios vary in this general way. Fluorescence is 81% of chlorophyll for CHL<2 and 55% for CHL>4.

##### 14. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The values found varied from 1.1 to 1.5s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Conductivity

Tests were run on 2 casts using shifts between -1.5s and +0.5s and T-S plots were prepared to compare the results. A setting of -0.8s worked best for the primary and -0.4s for the secondary. The primary setting is the same as for cruise 2006-15; the secondary is slightly higher.

All casts were put through SHIFT using -0.8 records for the primary conductivity and -0.4 records for the secondary. (Output *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on 2 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Values from +100 to +140 were tried and the best overall match of features was with a choice of +120 records. In previous use of this instrument settings of +110 to +150 have been used. All casts were shifted by +120 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #61 at the surface of the upcast.

##### DETAILED EDITING

The secondary sensors were chosen for the archive because the secondary salinity was closer to the bottles and has been used on other 2006 cruises using this equipment. Page plots were produced using (T1,S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

The descent rate was generally very noisy. Both salinity channels were unusually noisy, even during cast #1 when the descent rate was very steady. Such noise was also noted during 2006-15. The spikes were generally two-sided which suggests flow-rate problems. Editing was applied to salinity in cases where the spikes were unstable and not likely to be removed by metre-averaging, or not two-sided.

All casts required some editing. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

The salinity was found to be high by 0.0019 during this cruise and by 0.0009 and 0.0013 during 2009-24 and 2006-15, the cruises immediately before and after. The results of 2006-15 are based on many bottles and are mid-way between the results of 2006-24 and 2006-14, so those will be applied to this data set.

The dissolved oxygen data will be divided into two groups. Cast #15 fits well with the earlier group and #17 with the later group; the question is what to do with cast #16. Plotting SBE dissolved oxygen versus salinity suggests that #16 is probably closer to #15 than #17, though it is by no means clearly so.

An initial recalibration was run on the MRG and SAM files using file 2006-14-recal1.ccf to subtract 0.0013 from the secondary salinity and to apply the following correction to the dissolved oxygen:

CTD-BOT = 1.1341 DOX-CTD - 0.0044 (Casts 1 – 16)

CTD-BOT = 1.0341 DOX-CTD + 0.0425 (Casts 17 – 83)

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well. The salinity differences average ~0.0006 below 100db and ~0.0002 below 300db, and there is little pressure dependence especially below 300db. The dissolved oxygen fit looks excellent in plots of DO versus pressure and SBE DO. There is a slight suggestion of time dependence in the differences, but it looks as though a few near-shore casts near the end are biasing the trendline.

The same DO and Salinity calibration was applied to the EDT files. (See 2006-14-dox-comp2.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 100db. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity: These sensors were used for 2006-13, -24, -15 and -33, but the first had only 2 casts. The primary salinity was found to be high by 0.005, 0.0053, 0.0073 and 0.0144 and the secondary by 0.0015, 0.0009, 0.0013 and 0.0113. The results of 2006-33 were not trusted. The primary conductivity has shown some signs of pressure dependence.

2. Dissolved Oxygen: This sensor was used for the same cruises as the conductivity sensors. For 2006-24 and 2006-33 there were different calibrations used for different parts of the cruise due to some low DO conditions affecting the calibration. When it was returned to the factory the membrane was found to need replacing and one effect of that is to make the sensor more susceptible to calibration drift due to low DO values.

3. Pressure: The pressure sensor has been found to need an offset of +1.2db for the 2 cruises before this one and the two after.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The only excursion of salinity from those ranges was at cast #9 between 50 and 100db with salinity a little high. The temperature was high for offshore casts in the LB line below 500db. These excursions do not look indicative of calibration problems and more likely reflect real conditions not well represented in the local climatology.

Repeat Casts – There were no repeat casts.

##### Final Calibration of DO

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences for cast #1 are very noisy. The differences are very low for casts (2-15); for casts #17 to 83 the average difference indicates the dissolved oxygen is high by an average of 0.018ml/l. There is some evidence of pressure dependence in the second set of casts but this is based on very few bottles. For the first set there is only one bottle below 300m in the comparison so the absence of such pressure dependence is also unconvincing. It seems best to subtract 0.02ml/l from SBE:DO for casts #17 to the end. (See 2006-14-dox-comp3.xls.)

A second recalibration was applied to the downcast files only, casts #17 to 83. (AVG, THN and CLIP) to subtract 0.0ml/l. (Output: COR2, THN2 and CLIPCOR)

A final run of COMPARE was done to ensure the recalibration was done properly and it looks fine. (See 2006-14-comp4.xls.)

The clipped files were recalibrated, bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were filtered before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added and the channels reorder to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.8ml/l from 0 –100db

- ±0.2ml/l from 100 –200db

- ±0.1ml/l from 200 – 1200db

- data below 1200db are considered unreliable by the manufacturer.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The near-surface values were between 60% and 110% with the highest values well offshore. These results seem reasonable.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.MRGREM)

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. The standard comments for hydro files were adapted to reflect the fact that the chlorophyll was analyzed on board. The following comments were entered:

Chlorophyll samples are collected in calibrated polycarbonate

bottles (~330ml), then filtered onto 25mm MFS GF/75 filters at <70mmHg

vacuum. The filters are placed in glass scintillation vials and immediately

frozen at -20ºC until analysis. The chlorophyll is extracted in 10 ml of 90%

acetone for 24h at -20ºC, then brought to room temperature and the extract

is decanted and measured using a Turner Designs Model 10 fluorometer,

which has been calibrated with solutions of commercially prepared

chlorophyll.

Chlorophyll samples were analyzed on board ship by Marie Robert.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### 22. Thermosalinograph Data

a.) Checking calibrations

There was 1 file containing TSG data. A report was printed for the con file, the temperature and fluorometer calibrations were corrected and the serial number for the fluorometer was added and the resulting file was saved as 2006-14-TSG.con.

The history of the sensor was obtained.

b.) Converting to IOS Headers, adding position headers and time channels, preliminary checks

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format. There was no secondary temperature channel available.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that is consistent with the initial time noted in the log.

ADD TIME CHANNEL was used to add time and date channels in IOS SHELL format and the output files were named *.ATC.

Time-series plots were produced. The salinity data have many spikes that are not associated with spikes in any other channel including the flow rate. Editing will be required to interpolate or remove these spikes. It was also found that the first 19 records have very low flow rate, so the temperature, salinity and fluorescence will have to be replaced with pad values.

A preliminary track plot shows there is at least one error in positions. The problem was tracked down by plotting and fixed using a text editor to replace bad values with pad values in both latitude and longitude for one record. A preliminary comparison of positions/time between log book and TSG indicates that there is no problem with time such as was noted during 2006-33.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within .3db of 4db and exported to a spreadsheet. The TSG file was opened in EXCEL, median and standard deviations were calculated for temperature, salinity and fluorescence channels, and the file was then reduced to the times when CTDs were run. Those files were combined in a spreadsheet (2006-14-ctd-tsg-comp.xls). The positions were compared and were very close, with average differences for both latitude and longitude of <0.0001º and no difference greater than 0.0002º. This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are variations in alignment, but they are not systematic. This step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 There was no secondary temperature sensor on this cruise.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the TSG temperature and salinity with those of the CTD. There was a lot of scatter especially for casts in the first half of the cruise, probably due to rougher weather as evidenced in the descent rate of the CTD. The TSG salinity is noisy and was, at this stage, unedited which explains one extreme outlier in the comparison.

The standard deviations in the TSG data were used to eliminate 7 outliers in each of the temperature, salinity and fluorescence comparisons and to select the 5 casts with the lowest standard deviations in each channel. (See 2006-14CTD-tsg-comp.xls)

The following table summarizes the results of the comparisons, Ttsg-Tctd, Stsg-Sctd, and FLtsg/FLctd.

|  | Temp Diff | Sal Diff | FL ratio |
| --- | --- | --- | --- |
| average all | 0.2683 | -0.0379 | 2.1225 |
| average excluding 7 pts with highest std. dev. | 0.2417 | -0.0355 | 2.0090 |
| Max of lower Std Dev pts | 0.4557 | 0.0235 | 3.3309 |
| Min of lower Std Dev pts | 0.0729 | -0.1496 | 1.4901 |
| average 5 lowest std dev | 0.2139 | -0.0274 | 1.7958 |
| Average of 5 casts with best mixed surface | 0.1562 | -0.0364 | 2.1399 |

When the outliers are excluded there is still a lot of scatter. The TSG primary temperature was consistently higher than the CTD and the salinity lower in most cases. Both the TSG and CTD are subject to error and matching the depth of the CTD data to the intake of the TSG inexact. The latter matching is less critical if the surface waters are well-mixed. Choosing 5 CTD casts with well-mixed surface waters suggests that the TSG temperature is high by 0.16Cº and the salinity low by 0.036. Choosing the 5 data points with the lowest TSG variability suggests the TSG temperature is high by 0.21Cº and the TSG salinity low by 0.027. Unfortunately these two “best” sets don’t overlap except for cast #83 which has the lowest variability in TSG and is well-mixed. It indicates that the temperature is high by 0.25Cº and salinity low by 0.026.

The TSG fluorescence was higher than that of the CTD. The ratio of the two fluorometers did not vary with distance from shore. The average for the 9 casts closest to shore and the 11 farthest from shore were 2.27 and 2.31 respectively. When the ratio is plotted against CTD Fluorescence it does appear that the ratio varies inversely with CTD Fluorescence. The highest fluorescence values were in the middle sections of each line, near the 100m depth contour, and the ratio is ~1.5 at those casts.

- Loop Bottle Comparisons There were 11 loop bottles. The salinity values (using a median over a 2-minute window) were compared with the TSG files and there is a lot of scatter but for every bottle the TSG salinity is low, by from 0.02 to 0.13, with an average of 0.07 using all bottles and 0.05 for the 7 casts with low standard deviation. The chlorophyll-fluorescence comparison is even noisier. When 3 outliers are excluded (based on large standard deviation in the TSG data) the TSG fluorescence is higher than the chlorophyll by about 0.7ml/l. This difference looks fairly flat with fluorescence. The variability was very high even when the standard deviation is low. (See 2006-14-loop-comp.xls.)

Since it was noted during 2006-13 that differences during the outward journey were lower than during the inward leg, the data was examined to see if such a pattern stood out. There is a slight hint that when steaming faster the differences are higher, but those are generally also times when the ship probably was exposed to the weather on the port side.

- Calibration History

The TSG was recalibrated in October 2005. It was used for 2006-08, -11, -21, -15 and -33 when the temperature was found to be high by 0.22, 0.2, 0.18, 0.16 and 0.19Cº, respectively. The salinity was found to be low by 0.20, 0.022, 0.022, 0.19 and 0.23, respectively. During 206-21 the recalibration was based primarily on history. The 2006-24 results were not trusted, but temperature was found to be high by 0.13 and salinity low by from 0.19 to 0.28. During 2006-13 there were many loop samples but only 1 CTD cast for comparison; the temperature was found to be high by 0.19Cº and the salinity is low by between 0.13 and 0.23 with an average of -0.18; a notable difference was found between outward and inward legs of the cruise.

Conclusions

Using the 5 well-mixed casts, the TSG temperature is higher than the CTD by ~0.16Cº which is about what we expect for the TULLY at this time of year.

Using the 5 well-mixed casts the TSG salinity is lower than the CTD by 0.037. Compared to the loop samples it appears to be low by ~0.05 if we go with the casts with low standard deviation. However, if we use the mean it would be low by ~0.07 and that is probably a wiser choice. These results are out of step with the history of this instrument when it has generally been found low by about 0.2. The loop samples are mostly taken while steaming and not one of the loop salinity samples differs from the TSG by more than 0.13. The salinity data was analyzed at IOS and loop and rosette samples were run at the same time.

When we choose data based on the standard deviation being low, we are biasing the results to when the ship was stopped. The comparison with CTD casts has the same bias. Such a bias is unimportant if the equipment operates in the same way in either case. Using the data while stopped does seem to reduce the noise allowing a more reliable comparison. But it is not appropriate if something is really different while steaming. In that case the recalibration should be chosen to suit the record overall.

There does seem to be some mystery about the TSG system. We do not have sufficient information to calibrate different sections in different ways, and the history may not be relevant if the conditions were quite different from other cruises. So, corrections will be used that suit the overall data set. Corrections of +0.07 to salinity and -0.16Cº to temperature appear appropriate, but this could be revisited later if more is learned about what is causing the strange variability in the TSG data. It is a pity we did not have the intake temperature sensor on during recent cruises. It might help clear up the mystery of what is different while stopped or changing direction. If temperature is not involved then we need to focus on conductivity. Could there be more bubbles in the system under certain conditions?

f.) Editing

The time-series plots were examined and there was a lot of salinity spiking with no associated temperature spikes.

CTDEDIT was used to clean such salinity spikes and to remove the first 20 records because the flow meter indicates there was no flow for records #1 to 19. Record #20 looks like the flow had not settled down, but by #21 the data looks ok.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.16 Cº and +0.07psu to Temperature:Primary and Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary, Flag and UPloy0 (flow rate). The HEADER2 positions in the headers are wrong, presumably because of the bad position that was mentioned in section 22(b). CLEAN was rerun on the file to fix that.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

4. Event number repeated for Net and Rosette cast that followed. Error in longitude.

11. Computer crashed.

20. Bottom depth variable

21. Bottle #10 not tripped.

26. Tightened weights on rosette

30. Bottle #3 did not close; no sampling

31. Computer crashed beginning of cast

40. Bottles #1 and 2 tripped at same depth

43. Bottles #13 and 15 – valve may have been open

61. Delay in stopping acquisition

68. Altimeter not picking up bottom

70. Not sure of bottom depth

79. Upcast not logged

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2006-14 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 12 Sept. 2006 End: 12 Sept. 2006 |  |  |  |  |  |  |
| Location: WCVI |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Yelland D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550        Cruise ID#:		2006-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 18May06 | Factory “ |  |  |
| Conductivity | 2399 | 23May06 | “ |  |  |
| Secondary Temp. | 2038 | 07June05 | “ |  |  |
| Secondary Cond. | 2754 | 23May06 | “ |  |  |
| Transmissometer | 723DR | 28March06 | Factory |  |  |
| SBE 43 DO sensor | 47 | 10June2006 | Factory |  |  |
| Fluorometer | 2228 |  | IOS |  |  |
| Pressure Sensor | 75636 | 04/06/1999 | Factory |  |  |
| Altimeter | ? | ? | ? |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2006-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 21/10/05 | Factory |  |  |
| Conductivity | 2487 | 21/10/05 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? | ? |  |  |