## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |  |
| --- | --- | --- |
| ### 10 Dec 2020 | ### Added HPLC data. S.H. |  |
| 28Oct2019 | Corrections to header comments and flags for most CHE files. |  |
| 16-Aug-2017 | Changed suffix of 2006-15-surface0001.HED underway file to PCO2. Suspect suffix was HED because Header Edit was the last program used on this file and listed in the History section. |  |
| 24-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. |  |
| 10-Jun-2013 | Added Iron profile files with cast numbers 8xxx from Keith Johnson’s spreadsheet file which can be found in the cruise .DOC directory. |  |
| 11-Apr-2010 | Added Lisa Miller’s Dissolved Inorganic Carbon and Alkalinity data to the rosette files. J.L. |  |

## PROCESSING NOTES

Cruise: 2006-15

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: September 20, 2006 – October 9, 2006

Processed by: Germaine Gatien

Date of Processing: October 25, 2006 – 23 November 2006

Number of original CTD casts: 69

Number of CTD casts processed: 69

Number of bottle casts: 30 full rosette casts and 39 surface samples

Number of bottle casts processed: 63 (+5 files for UBC without chemistry data, 1 file empty)

TSG files: 1

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea/Seatech transmissometer (#723DR), a SeaBird DO sensor #0047 (on the primary pump) and a Seapoint Fluorometer (#2228) with a 10X cable (on the secondary pump). A Biospherical PAR sensor (#4656) was mounted for some casts. The deck unit was a model 911 (#0508) and the logging computer was PAC02328 Shuttley PACOSAP-FS03. Two salinometers were used: a Portasal 8410A A/N 68627 and an Autosal 8400B S/N 68572. A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P).

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order although two of the rosette sheets lacked the information that there had been nutrient sampling. Additional notes from the chief scientist were very helpful.

The ADD files had non-standard names.

The salinity data are unusually spiky even when the descent rate was high and steady, suggesting problems with the flow rate. Most of the spikiness is 2-sided so errors should be minimized by metre-averaging.

The descent rate of the CTD was generally kept quite high, on average. Those casts with an average of about 1.5m/s produced much cleaner data than those with similar noise levels in the pressure signal but lower average descent rate. The higher descent rate does reduce the amount of data collected but it looks to be of significantly higher quality.

The SBE dissolved oxygen sensor had a lot of noise during casts 39, 45, 48, 60, 74-77. This was noted at sea and the problem was resolved after a cable was replaced. Where obvious, the bad data were removed.

The thermosalinograph salinity was edited heavily to smooth many spikes in salinity and to remove salinity from large sections of the record where values looked unbelievable with sudden jumps in values which did not correspond to changes in temperature, fluorescence or flow rate. Comparisons with CTD data confirmed that many salinity data were bad. There remain some doubts about all the salinity data.

The Thermosalinograph temperature is considered ±0.2Cº.

The dissolved oxygen data in the CTD files are considered

- ±0.6ml/l from 0 – 250db

- ±0.25ml/l from 250 – 400db

- ±0.1ml/l from 400 – 1250db

- data below 1250db are considered unreliable by the manufacturer.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Together with the notes from Marie Robert this gave an excellent summary of problems encountered during the cruise.

Titrated chlorophyll, nutrients and salinity data were obtained with flag channels. Two salinometers were used, a Portasal at sea and an Autosal at IOS

The oxygen files lacked a flag channel, but the chief scientist discussed this with the analyst and no flags were required. The file names of the ADD files were changed to standard format (an extra zero had to be removed.)

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and errors were corrected in the CTD secondary conductivity and the TSG temperature; the fluorometer serial number was added to the TSG configuration file. Two CTD configuration files were prepared, one with PAR and one without.

##### Conversion of Raw Data

Data were converted using 2006-15-ctd-par for casts 4, 5, 34, 41, 61, 63 and 66 and file 2006-15-ctd.con for all other casts.

A few casts were examined and all expected channels are present. The secondary temperature has more structure than the primary though it is not obvious which is better. The two conductivity channels have very similar shapes, but are offset by about 0.0004S/m. Fluorescence was occasionally off-scale late in the cruise. Transmissivity looked fine. The dissolved oxygen trace has the usual offset between upcast and downcast. The PAR trace looks good.

Rosette files were converted using a start time of -2s and duration of 5s. The rosette file for cast #78 was empty. According to the log the bottle did not trip, so this rosette file will not be processed further.

Station names in the CNV and ROS headers were edited to correct the errors mentioned in the notes from the chief scientist.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT.

All BOT files were plotted and no significant outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on 3 casts with (alpha, 1/beta) set to (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03,9) and (0.0245, 9.5). For the primary the best setting was (0.02, 7) and for the secondary the best choice was (0.03, 7).

CELLTM was run using (0.02, 7) for the primary and (0.03, 7) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 10 | 1990 | -0.0002 | -0.00055 | -0.0065 | V.High, noisy |
| 26 | 1990 | ~0 | -0.00054 | -0.0055 | V.High, noisy |
| 53 | 1990 | -0.0002 | -0.00053 | -0.0063 | V.High, noisy |
| 106 | 1990 | -0.0002 | -0.00050 | -0.0058 | V.High, noisy |
| 10 | 1000 | -0.0005 | -0.0005 | -0.0055 | V.High, noisy |
| 26 | 1000 | -0.0001 | -0.00048 | -0.0056 | V.High, noisy |
| 53 | 1000 | -0.00025 | -0.00046 | -0.0053 | V.High, noisy |
| 106 | 1000 | -0.00045 | -0.00045 | -0.0050 | V.High, noisy |
| 10 | 500 | -0.0005 | -0.00048 | -0.0050 | V.High, noisy |
| 26 | 500 | -0.0004 | -0.00041 | -0.0044 | V.High, V.noisy |
| 53 | 500 | -0.0004 | -0.00038 | -0.0043 | V.High, V.noisy |
| 106 | 500 | -0.0006 | -0.00040 | -0.0042 | V.High, V.noisy |

There is significant pressure dependence in the conductivity and salinity with higher differences at depth. There appears to be some time-dependence in those two parameters as well.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run and no errors were found.

The cruise track was plotted and no problems found.

The average surface pressure is 2db, which is a little low for the Tully. The last cruises on record for this sensor are from 2004. For the last 3 uses offsets of +0.5, +1 and +1db were applied. Checking a few casts shows that there are many non-zero pressures with associated “in-water” values of salinity, transmissivity and fluorescence and the pumps were running for many of them. During the upcast of event #99 fluorescence, transmissivity and conductivity go to zero at -1.2db. The pumps were not running at that point but that would not affect the transmissivity. An offset is clearly needed for this data and doing it early will ensure that good data are not lost in applying DELETE.

CALIBRATE was run using calibration control file 2006-15-recal-pr.ccf to apply an offset of +1.2db to the pressure channel.

The altimeter readings from the header were exported to a spreadsheet and a few casts were checked and show that the algorithm worked well.

##### BOTTLE FILE PREPARATION

NOTE that the salinity analysis sheet mentions that there were two samples labelled #533. It looks as though one was later renamed as sample #529.

The chief scientist reported that Niskin Bottle #13 may not have performed well.

According to the log there were problems with spiking in the dissolved oxygen sensor channel and many steps were taken to try to resolve this. After cast #77 there were no such problems.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. For cast #41 Niskin Bottle #20 was fired but not sampled. The sample number assigned at sea was 226B; this was changed to 9226.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)  Some bottles (casts #7, 29, 48, 67 and 88) were fired only for UBC researchers, so there are no samples to be added; bottle files were prepared for the use of UBC, but will not be archived.

The salinity data were complex because they included loop data and rosette data and were created in two batches using different salinometers. There were 3 files provided:

- Salinity_summary.xls (created by Doug Yelland) contains samples analyzed at sea on the Portasal. The flags indicate that the variability was >0.0005 among the 3 readings.

- 2006-15_sal.xls (created by Hugh Maclean) contains samples analyzed at IOS with the Autosal. The standard deviations are included in that file and the maximum is 0.0003 so there is no suggestion of problem in the analysis.

- 2006-15 sal.xls (created by Marie Robert) is a compilation of both files, with cast #, sample #, station name, salinity, and some flags.

The 2006-15 sal.xls file was saved as 2006-15 sal.csv and then edited. Comments were added for the flagged values, header names were changed to standard names and the loop values were removed. (A separate file was created with just loop data to be used later during TSG processing.) The file was then converted to individual SAL files.

Dissolved oxygen files (*.add) lacked a flag channel, but it is known that the analyst found no problems that needed flagging. A flag channel was added using the IOS SHELL routine ADD CHANNEL.  There was an error in sample numbers for cast #113. That was corrected. An error in sample # was made in the OXY file for cast #81; the sample # said to be from 10m was actually from 5m so that was changed in the ADD file.

The nutrient spreadsheet was simplified and saved as 2006-15nuts.csv. The results for casts with only one bottle files were in a separate sheet in the analyst’s file, but were combined with the other Niskin data for this file. Extraneous columns were removed, header names were changed to standard format and the data were sorted on sample number. There were duplicate samples for sample #567 that had not been averaged; they were averaged and the flags changed to “f” – all other duplicates had already been averaged. File 2006-15nuts .csv was then converted to NUT files. The loop data were saved separately as 2006-15nuts-loop.csv.

Extracted chlorophyll data were obtained and saved as 2006-15CHL.csv. The results for casts with only one bottle were in a separate sheet in the analyst’s file, but were combined with the other Niskin data for the chemistry files. The file was edited to remove extraneous lines and columns, header names were changed to standard format, data were sorted on sample number and loop data were saved separately as 2006-15CHL-loop.csv. 2006-15CHL.csv was converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

Sample #92 cast #18 will be flagged “e” and the values replaced with pad values because all samples suggest the bottle fired at depth, not at the surface.

DO was plotted against CTD Salinity. The only outlier was near the surface for cast #45 but in an area of high gradient, so it is not clear that there is anything wrong with the bottle.

##### 12. COMPARE

Salinity

COMPARE was run with sample number as the reference channel. When data above 500db were excluded the differences between bottles and CTD were quite flat for both pairs of sensors, and especially flat for the secondary. However, the primary differences below 2000m look somewhat higher than between 500 and 2000db. The secondary salinity does not show that behaviour. In section 7 it was noted that the differences between the two sensors did increase with pressure. It appears likely that the primary salinity accounts for this.

There were 4 large outliers, three are from the surface and the other is in an area of high CTD salinity variability, so no flags were added. Mild outliers that were flagged “c” are:

- Cast #45 Sample #227, 4020db (From Niskin #1)

- Cast #45 Sample #232, 2000db (from Niskin #6)

- Cast #76 Sample #392, 2000db (from Niskin #13, outlier compared to 22 other bottles at same site)

The trend in differences against time have the same slope for both pairs of sensors suggesting the variations probably reflect something other than calibration drift.

COMPARE was run again using the Niskin bottle # as reference channel. When the outliers noted above are removed as well as values with pressure <500, no bottle stands out as a problem. The chief scientist has expressed some concern about Niskin #13. There was one major outlier, but it was at the surface in a high salinity gradient area. There was also the outlier mentioned above which was just slightly out of line. There is no evidence of major problems with the bottle.

The primary salinity is high by an average of 0.0073psu and the secondary is high by an average of 0.0013psu.

Dissolved Oxygen

COMPARE was run 3 times using sample #, pressure and Niskin # as the reference variable. The fits are not as tight as usual though the differences are relatively small compared to results before the recent calibration and repair. Dividing the casts into different areas and casts with or without well-mixed surface waters did not produce much difference. It is likely that the noise level is a reflection of the fact that the DO sensor is working better and showing more detail in the profiles than has been the case in the past with this sensor. The fit looks much like sensor #0766 since it was repaired.

The best fit was with the differences between bottles and CTD versus CTD DOX. First, data from below 1500db were excluded from the fit and then outliers were excluded based on residuals; the fit found was:

CTD-BOT = 1.0512 DOX-CTD + 0.0599

When plotted against file pair number there is no suggestion of temporal drift. A few casts were examined in detail to see why the scatter was so high. It was found that in many cases there was a subsurface maximum in DO and generally high variability that the SBE sensor would be unlikely to track well. No DO values were flagged. (See 2006-15-dox-comp1.xls.)

Because there were some casts with very noisy patches in the SeaBird dissolved oxygen channel a check was made that no such data were corrupting the comparison. Fortunately, the bad data do not occur during the upcast of the relevant casts.

Fluorescence

COMPARE was run using the CTD Fluorescence and the Titrated Chlorophyll from bottles. When all data were plotted there was tremendous scatter, but when the data divided into offshore and inshore (with casts from P9 to P26 and P26 to ARGO0 being treated as offshore) the plots are much simpler. The near-shore data had the slope of CHL versus FL as ~1, but for the offshore data the titrated chlorophyll is about 1/3 of the fluorescence. This fits observations from other cruises and indicates that the fluorescence sensor behaved normally.

##### 14. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The values found varied from 0.9 to 2s with the highest offsets being at casts with extremely noisy descent rates. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Conductivity

Tests were run on 2 casts using shifts between -1.5s and +0.5s and T-S plots were prepared to compare the results. A setting of -0.8s worked best for the primary and -0.2s for the secondary.

All casts were put through SHIFT using -0.8 records for the primary conductivity and -0.2 records for the secondary. (Output *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on 4 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Values from +100 to +150 were tried and the best overall match of features was with a choice of +140 records. In previous use of this instrument settings of +110 to +150 have been used.

All casts were shifted by +140 records. (Output: *.SHFO)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: None.

##### DETAILED EDITING

The secondary sensors were chosen for the archive. Both salinity traces had many spikes but the secondary salinity was closer to the bottles and less pressure-dependent. Page plots were produced using (T1,S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used. The salinity was unusually spiky even when the descent rate was high and steady. (The primary salinity was no better.) The spikes were generally two-sided which suggests flow-rate problems. Editing was applied to salinity in cases where the spikes were unstable and not likely to be removed by metre-averaging, or not two-sided. Most of the editing was in the high gradient region just below the mixed layer.

All casts required some editing. Note was made of the editing details in the headers of the relevant files.

After editing a small study was done to see how the descent rate affected the amount of data removed in editing and running DELETE. It was found that as the descent rate went from 1 to 1.5m/s the amount of data removed went down from 6 to 8% to less than 1%. Of course, using a higher descent rate meant there was a lot more data to start with in using the lower descent rate. So there is a trade-off – more data and more editing or less data and little editing. The balance is probably in favour of the higher descent rate since the data are in fact of higher quality. The thing that is more difficult to assess is the data removed by DELETE which is much higher with the lower descent rate. It would be very time-consuming to calculate how much DELETE removed from the downcasts. (See 2006-15-edit-study.xls.)

##### Initial Recalibration

An initial recalibration was run on the MRG and SAM files using file 2006-15-recal1-bot.ccf to subtract 0.0013 from the secondary salinity, to add 1.2db to the pressure and to apply the following correction to the dissolved oxygen:

CTD-BOT = 1.0512 DOX-CTD + 0.0599

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the same DO and Salinity calibration was applied to the EDT files but file 2006-15-recal1.ccf was used since the pressure had already been corrected for the full CTD files. (See 2006-15-dox-comp2.xls.)

##### Special Fluorometer Processing

The COR files were clipped to 100db. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

At this stage CTDEDIT was used to remove bad points in the SBE dissolved oxygen channel in casts 39, 45, 48, 60, 74, 75, 76, 77. The noise was noted at sea and the problem was resolved after cast #77 when a cable was replaced.

##### Other Comparisons

Previous experience with these sensors – These sensors were used for 2006-13 but there were only 2 casts and only one that was very deep. The primary salinity was found to be high by 0.005 and the secondary by 0.0015. The secondary conductivity sensor was also used on one other cruise since the last calibration but was on an SBE25 with a different temperature sensor and only 3 bottles. It was found to give salinity high by 0.004 at that time. There is no useful history for the dissolved oxygen sensor since it was repaired and recalibrated in June 2006 and the only other use was on the 2 casts of 2006-13. The results for those 2 casts were similar to those of this mission.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were a few isolated, minor excursions from those ranges, but nothing that looks indicative of calibration problems.

Repeat Casts – There were repeats of 4 deep casts at stations P16, P20, P26 and Z9. The bin-averaged files were plotted together on T-S surfaces and the differences found at depth were on the order of 0.001 to 0.005Cº in temperature and 0.0003 to 0.002 in salinity with the largest differences at P16. The results at P20, P26 and Z9 show much smaller differences as is expected based on local gradients.

##### Final Calibration of DO

Files were bin-averaged to 0.25m bins for just the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences were very noisy especially near the surface. This is to be expected because there was a subsurface maximum that was fairly sharp (seen in both bottle data and CTD data), so small mismatches between bottle and CTD will have a large effect. We are trying to match upcast and downcast so this is a significant factor. The sensor appears to be recording more detail than it did before a recent repair, so that is another factor to be considered. It would appear that the sensor is reading a little high, but the best we can do is average all the differences, which would lead to a correction of anywhere from -0.002 to -0.045ml/l depending on how many points are excluded as outliers. An offset of -0.03ml/l looks reasonable. It is close to the offsets used for 2006-13 and for cruises before the most recent recalibration. A second recalibration was applied to the downcast files only (AVG, THN and CLIP) to subtract 0.03ml/l. (Output: COR2, THN2 and CLIPCOR)

A final run of COMPARE was done to ensure the recalibration was done properly and it looks fine. (See 2006-15-dox-comp3.xls and 2006-15-comp4.xls.) A check was made again that the corrupted DO data had no effect on the comparison and it does not appear to have done so.

The clipped files were recalibrated, bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were filtered before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added and the channels reorder to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.6ml/l from 0 – 250db

- ±0.25ml/l from 250 – 400db

- ±0.1ml/l from 400 – 1250db

- data below 1250db are considered unreliable by the manufacturer.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The near-surface values were between 90% and 110% except for station P2 and 3 casts in Rivers Inlet which had values between 70% and 90% near the surface. These results seem reasonable.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.MRGREM)

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together. (Output: CHE1)

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

There were a few hydro files which contained no chemistry analyses. These were casts run for UBC researchers: events #7, 29, 48, 88, 106. The files were placed in a separate folder to be sent to UBC. Remarks were removed from the headers that referred to analyses methods.

The files for UBC were put through the same steps as other bottle files (except for the merges with chemistry data) and they were given a special header comment.

##### 22. Thermosalinograph Data

a.) Checking calibrations

There was 1 file containing TSG data. A report was printed for the con file, the temperature calibration was corrected and the serial number for the fluorometer was added and the resulting file was saved as 2006-15-TSG.con.

The history of the sensor was obtained.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format. There was no secondary temperature channel available. Two channels given as Voltage0 and voltage1 proved to be the fluorescence and flow rate channels. A pressure channel existed but contained only null data.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that is consistent with the initial time noted in the log.

ADD TIME CHANNEL was used to add time and date channels in IOS SHELL format and the output files were named *.ATC.

Time-series plots were produced. The salinity data have many spikes that are not associated with spikes in any other channel including the flow rate. Editing will be required to interpolate or remove these spikes.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within .3db of 4db and exported to a spreadsheet. The TSG file was opened in EXCEL and reduced to the times when CTDs were run. Those files were combined in a spreadsheet (2006-15-ctd-tsg-comp.xls). The positions were compared and were very close, with average differences for both latitude and longitude of <0.0001º and no difference greater than 0.0003º. This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are variations in alignment, but they are not systematic. This step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 There was no secondary temperature sensor on this cruise.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared comparing the two TSG temperature channels and the salinity with those of the CTD. There was a lot of scatter for casts near the coast, not surprising in Rivers Inlet. (See 2006-15CTD-tsg-comp.xls)

The TSG primary temperature was higher than the CTD by from 0.17 to 0.26Cº depending on what data are selected. The TSG salinity was lower than the CTD salinity by from 0.04 to 0.46psu. The ratio of the two fluorometers varied from about 1 near shore up to 3 well offshore.

Using a calculation of mixed layer depth based on salinity gradient, casts were selected that were well mixed to at least 10m. The averages were once again calculated. The primary TSG temperature was found to be high by 0.20 (all data) or 0.16Cº (casts up to #75). If casts 38 to 75 are used the temperature is high by 0.165Cº. The salinity is now found to be low by 0.15 (all data) to 0.04psu (up to #75). There is a sudden change in the salinity differences during the northern leg from P26 to Argo stations. Looking at the data it is quite clear that the salinity is suspicious, so this data should not be included in the comparison for calibration. (See 2006-15-ctd-tsg-comp.xls.)

- Loop Bottle Comparisons There were 29 loop bottles but no time was recorded for 7 of them. The salinity values (using a median over a 2-minute window) were compared with the TSG files and there is a lot of scatter with differences from -0.7 to -0.03psu. Standard deviation was calculated over the 2 minutes. The differences were plotted against the standard deviations and no obvious pattern emerged. The scatter is high even where the standard deviations are low. If only the 5 points with the lowest S.D. are included there is still a range from -0.15 to -0.29. This may be due to spikiness in the salinity. There is no obvious drift with time though the distribution of samples is limited with a few from Juan de Fuca, one at P1 and the rest all occurring after the ARGO stations on the way back. Using only the data from well offshore, excluding 1 outlier, the salinity was low by an average of 0.192psu. (See 2006-15-sal-loop-tsg-comp.xls.)

- Calibration History

The TSG was recalibrated in October 2005. It was used for 2006-08, -11 and -21 when the temperature was found to be high by 0.22, 0.2 and 0.18Cº, respectively. The salinity was found to be low by 0.20, 0.022 and 0.022 respectively. The TSG was also used for 2006-13 when the temperature was found to be high by 0.19Cº and the salinity is low by between 0.13 and 0.23 with an average of -0.18 with a notable difference between outward and inward legs of the cruise.

Conclusions

The TSG temperature is higher than the CTD by 0.16Cº compared to 0.18Cº during 2006-13. This is not unexpected since the temperature correction is usually smaller in September than in July because the heating by the ship is lower when ambient temperature is higher.

The TSG salinity is lower than the CTD by 0.04psu in the first part of the cruise and lower than the loop samples by 0.192. Later in the cruise the differences from CTD values goes up. Most of the loop samples are from a period during which the TSG salinity did not look reliable. However, 2006-13 found values between 0.13 and 0.23 with a lot of loop samples and no spikiness. Since the CTD comparison applies only to when the ship is stopped, it seems better to go with the loop results and history of the instrument and apply a larger correction, but it should be noted that the accuracy is in some doubt. It was learned during 2006-13 that the values can vary markedly during different parts of a cruise. These variations may somehow depend on the direction the ship is travelling, though no explanation is known for why that should be so.

Corrections of +0.19psu to salinity and -0.16Cº to temperature appear appropriate, but this could be revisited later if more is learned about what is causing so the strange variability in the TSG data..

f.) Editing

The time-series plots were examined and there was a lot of salinity spiking and some sections of very suspicious salinity data, mostly after cast #75. Two passes were made through CTDEDIT, the first to interpolate large spikes in salinity (1 or 2 points with no corresponding temperature spike) and the second to look carefully at the bad sections. A few more wild points were removed in the second pass and salinity points were removed in areas where the data look clearly wrong but there are two many points to justify interpolation.

In some cases it looks as though slight misalignments cause the conductivity to spike due to small changes in temperature, but in many cases the salinity drops off steadily and then suddenly resumes believable values. In the comparison of TSG and CTD it was noted that the differences increased markedly after cast #74.  In most areas the water was very well-mixed at the depth of the intake so it is hard to believe large variations. This is particularly true in the North-South leg where the worst noise was found. In the Rivers Inlet section there are larger gradients expected and the salinity looks reasonable, so little editing was done there. The flow rate and the fluorescence show no notable variations in the areas of bad salinity.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.16 Cº and +0.19psu to Temperature:Primary and Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary, Flag and UPloy0 (flow rate).

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

7, 11, 67 – large spikes in T and S.

18. Samples left in lab for 3 hours, NUTS & CHL not refrigerated.

24-26, 38-40, 45, 48, 75-77. Large spikes in DO sensor data.

60. Upcast interrupted – went from 400m back down to 550m and then up again.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2006-15 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 20 Sept. 2006 End: 9 October 2006 |  |  |  |  |  |  |
| Location: Gulf of Alaska |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Robert M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550        Cruise ID#:		2006-15

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 18May06 | Factory “ |  |  |
| Conductivity | 2399 | 23May06 | “ |  |  |
| Secondary Temp. | 2038 | 07June05 | “ |  |  |
| Secondary Cond. | 2754 | 23May06 | “ |  |  |
| Transmissometer | 723DR | 28March06 | Factory |  |  |
| SBE 43 DO sensor | 47 | 10June2006 | Factory |  |  |
| Fluorometer | 2228 |  | IOS |  |  |
| Pressure Sensor | 75636 | 04/06/1999 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2006-15

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 21/10/05 | Factory |  |  |
| Conductivity | 2487 | 21/10/05 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? | ? |  |  |