## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected Salinity:Bottle lost during addition of HPLC. S.H. |
| 13 January 2021 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2006-21

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG - JdeF

Party Chief: Masson D.

Platform: Vector

Date: June 12, 2006 – June 18, 2006

Processed by: Germaine Gatien

Date of Processing: 22 August 2006 – 13 October 2006

Number of original CTD casts: 63

Number of CTD casts processed: 63

Number of rosette casts: 26	Number of rosette casts processed: 26

Number of TSG files: 1		Number of TSG files processed: 1

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#723DR), a PAR sensor (#4565), a surface PAR sensor (#16504) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#0766) and a Seapoint Fluorometer (#2356) with a 10X cable were mounted on the pumps. The deck unit was a model 911 (#0425). The salinometer was an 8400B model Autosal (S/N 68572).

There was a thermosalinograph (S/N 2487) mounted with a Wetlab Wetstar fluorometer (#WS3S-889P).

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book and rosette sheets were available. The Thermosalinograph section of the log book was not completed, and since the TSG data had not been placed in the processing folder it could easily have been missed. The presence of loop samples was the only clue that there was something missing.

There are problems with the times in the SeaBird headers. The computer time disagrees with the NMEA time. In two cases the log entry is incorrect but note was made of this in the log comment column. The NMEA time looks ok.

Time and positions were not recorded in the log book for 3 loop samples; for another sample there was only an estimate of time and there are obvious errors in other records. TSG data was used to correct the loop times where possible.

There were errors in event numbers on the rosette log sheets for 2 casts.

The fluorescence dark value is on the order of 0.12ml/l at 200 to 300m. This is higher than expected and should be studied further by looking at future use of sensor #2356 in other regions.

The transmissivity was very noisy for one cast with many zero values. It was not obvious how to edit this data so it was left as is, with a note of warning in the header. (Event #62)

The SBE dissolved oxygen sensor performed well with some detail in the profile and good correspondence between downcast and upcast. Surface saturation was unusually high in Saanich Inlet (~155%) during both the downcast and upcast. A bottle confirms the dissolved oxygen values are correct, so this is assumed to due to intense biological activity.

The SBE dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l in the top 200m

- ±0.1ml/l below 200m

Most stops for rosette sampling lasted about 30s, but many were significantly shorter which may have contributed to unusually noisy comparisons between CTD data and bottle samples, though it does not explain the unusually large differences between them. Salinity has not been recalibrated due to uncertainties in the comparison.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained. There is note of problems with the time in the log for casts 10 and 11. There are two times in the SeaBird headers – System UpLoad Time and NMEA time. The latter appears to be correct in all cases. For casts 1 to 11 the System UpLoad Time appears to be in PDT. For casts 10 and 11 PDT times were entered in the log book. In all other cases the log book appears to be correct with UTC time. For casts 12 through 32 the System Upload Time is UTC +1 hour. From cast #13 on the times are in agreement. In checking the time it was noted that the headers are scrambled for cast #10 so those were fixed.

The nutrient, chlorophyll and salinity data were each obtained in spreadsheet format. Each spreadsheet was edited to remove information not needed for the chemistry files. There were flag and comment channels for all. The salinity and nutrient data were divided into 2 spreadsheets for loop and rosette samples. There were already 2 separate chlorophyll files. Individual dissolved oxygen bottle files were obtained; flag channels and comments were included.

The CTD calibrations were checked, a few small errors corrected and the file saved as 2006-21-CTD.con.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

The files were converted to CNV files. A few files were checked and the pressures are never <0.9db.

As usual for this area the descent rate was highly variable; it was very steady for some casts, very noisy and fairly low for others. Bottle stops typically lasted over 30s, but many were shorter.

The temperature and conductivity channels are fairly close during downcasts, but significantly different during upcasts.

Transmissivity shows reasonable agreement between downcast and upcast but looks very spiky at depth for some casts.

As has been observed with this particular sensor since it was last repaired, the dissolved oxygen looks good with upcast and downcast offset vertically, but without the sort of hysteresis that was always observed in the past.

The fluorescence looks ok near the surface, but the deep values are unusually high for some casts in both downcast and upcast. Cast #1 is shallow, the next 3 casts (2, 4 and 5) look odd with values >0.5 well below 100m, but this is an area of high mixing. The casts from the mouth of Juan de Fuca Strait look ok.

The altimetry is often extremely noisy but the data generally look reliable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format. CLEAN was used to add event numbers and the extensions were changed to BOT. All BOT files were plotted and a few significant outliers were found in the primary salinity for casts #62. The bad data was removed using CTDEDIT and the output copied to BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Three casts with a steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5) were tried. A choice of (0.02, 7) was best for both channels though there was not a huge difference between some of the settings. CELLTM was run using (0.02, 7).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 28 | 260 | -0.0021 | -0.000035 | +0.0015 | Moderate, steady |
| 42 | 265 | -0.0021 | ~0 | +0.0015 | Moderate, steady |
| 65 | 180 | -0.0024 | +0.00001 | +0.002 | Moderate, steady |

The temperature differences are larger than usual but similar to those of 2006-11. A deep cast (#39 to 425db) was also examined to see if the salinity differences show significant pressure dependence and they do not.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header summary was checked carefully against the log book entries and a few errors were found. The station name is wrong for casts #1, 11, 23, 24, 25. These were corrected in the CLN files and BOT files.

There are inconsistencies in times, but a note in the log indicates that there was a problem with the clock and the header times appear to be correct.

The header check was run and no errors were found.

Track plots were produced and added to the end of this report.

The average surface pressure is ~1.1db which is low for the Vector and there are some negative values. Examining these files suggests that a very small offset may be appropriate since the negative values are mostly associated with “in-water” conductivity values, but the pumps were not on and the CTD might have been dipping in and out of the water. The temperature and transmissivity values suggest that it was mostly out of water when negative pressures occur. According to the log book a pressure check was done in Saanich Inlet and the results is noted as “fine”. No adjustment will be made to the pressure.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets and the file was then converted to CST files to form the framework for the bottle files. Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

There was an error on the rosette sheet for event number #29 which was given as #30. This will have to be checked in the analysis results.

The salinity data was provided in EXCEL format with flag and comment channels. No data had been flagged. Channel names were changed to standard format and the file was saved as 2006-21-sal.csv. The files were converted to individual SAL files. The file named 2006-21-0030.sal was renamed 2006-21-0029.sal.

The extracted chlorophyll data was in spreadsheet format and included a flag channel and comments. The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The resulting file is 2006-21chl.csv. That spreadsheet was converted to individual files. The following comments from the analysis sheet was added to the comments that will later be added to the headers of the bottle files: “Average of two samples is reported unless stated otherwise. Variability is assessed as the % (std dev/mean*100).”  (Output: CHL)

The dissolved oxygen data was received in individual files with flag channel and comments. The only flagged value had format errors; those were fixed and the comment was edited to make clear that it applied to the DO channel. The file named 2006-21-0030.add was renamed 2006-21-0029.add.

The nutrient data was provided in spreadsheet format with flag channel and comments. The file was divided into loop samples and rosette samples. Channel names were changed to standard format and lines with no sampling were removed; the files were saved as 2006-21nuts-rosette.csv and 2006-21-nuts-loop.csv. Event number 30 was changed to 29. The rosette data were converted to individual files.

The SAL, CHL, ADD and NUTS files were merged with the CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4.)

Those files were then merged with SAMAVG files (MRG) and then put through CLEAN to remove SeaBird headers and comments from the secondary files. (MRGCLN)

##### 11. COMPARE

Salinity

COMPARE was run and one major outlier for both sensors was sample #1, cast #1, Niskin #1. It was flagged “d”.

There is a lot of scatter in the results with significant pressure dependence. There is some time dependence but given the nature of the cruise that could be a geographic issue rather than calibration drift. It is surprising to see so much pressure dependence in both sensors (and the distribution is very similar for the 2 sensor pairs) which raises the question of whether there can be a sampling problem. Few of the salinity samples came from the bottom, so there would not appear to be a concern with contamination by mud. However, the transmissivity does seem unusually low for many of the bottles where salinity samples were gathered. Removing bottles from the comparison for which transmissivity was less than 40% does produce a flatter result with pressure but there is still a lot of scatter.

Various runs of COMPARE were done to check possible explanations for the scatter. The same Niskin Bottle was used for most deep sampling so that is not a factor. Plotting differences against transmissivity values confused the issue. If only data from below 125m was included there was some indication that scatter increased as transmissivity decreased, but only if a few deep outliers were removed. There is a hint that the bottle salinity could be reading too low at lower transmissivity values, but this is not very convincing. There have been problems noted in the past with low transmissivity water sampling being associated with anomalously low bottle values.

When the standard deviations in the CTD salinity channels were low there was somewhat less scatter in the differences but they were still higher than expected, so noise in the CTD data would not seem to explain the scatter. Looking at the descent rate of the CTD during the stop shows that there was a lot of motion during 6 of the rosette casts, mostly in the early section of the cruise. For 2 of those casts the differences are large, but the others are not out of line, so this does not seem to be an explanation either.

The scatter does not appear to be related to the CTD calibration or noise in the signal. That leaves problems in sampling or salinity analysis. There could be a problem in Niskin bottle #2 or with how the water was drawn from it.

The primary is a little closer to the bottles than the secondary, but is still high by an average of 0.005 when outliers were removed. It is noteworthy that excluding a few outliers the difference between the two CTD sensors varies little. This does support the idea that the scatter is related to sampling or analysis.

A few stops were examined in detail to see if the stops were long enough. Some were very short (ex. casts #10, 25 and 62), but most were about 30s long. In some cases the CTD dropped throughout the stop, and the CTD salinity was quite noisy for others. Three stops that look excellent (long enough and little CTD motion) were identified. They had differences that varied from 0.008 to 0.017 for the primary salinity and the secondary is similar. Based on just the 3 examples (casts #29, 51 and 66) one could conclude that the differences are growing with time, but it seems unbelievable that both pairs are moving in exactly the same way.

Dissolved Oxygen

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. A few outliers were investigated, one had already been flagged by the analyst (sample #68) and another (sample #213) was clearly out of line as were the nutrients from the same bottle, so it was flagged “d”. Sample #266 looks a little out of line, but it was not flagged at this point.

COMPARE was run excluding any DO samples with “d” flags. Plotting the differences versus CTD DO and pressure and excluding cases where the differences were >10 and then residuals <-.2 and >0.2 resulted in the identification of a few more outliers. Those samples are only slightly out of line so will not be flagged. The fit was:

DOX-BOT = 1.0866 * DOX-CTD + 0.0412

The fit against pressure excluding the same points does not fit a straight line as well:

DOX-BOT = DOX-CTD -0.0008 * Pressure +0.4361

The fit versus time shows no significant temporal drift when the same outliers are excluded as for the other fits; however, it is noted that both samples from the first cast were excluded in the fit.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. When all data is included the fluorescence is about 91% of the titrated chlorophyll values with an offset of about 0.25. When the offset is set to 0, fluorescence is 103% of CHL. There is only one sample (cast #56, station 9) that looks like a severe outlier, but this looks like just a mismatch of CTD and rosette since the CTD does have values as high as the bottle slightly below the depth of the bottle. This probably means that the bottle has not flushed completely so that the water in the bottle derives from a little deeper than the bottle stop depth. Or there could just be small real variations over the metre between the two. At the surface the fluorescence gradient is very large, so a minor mismatch can produce huge differences. No flag will be applied to this sample. The fact that the offset in the fit appears to be driven by the higher CHL value points may be another sign that minor mismatches in CHL and FLUOR account for the difficulty fitting the same trendline throughout.

The dark value is anomalously high for a few casts (notably #2 and 76), but given the area this could be real. For other casts the dark value is on the order of 0.12ml/l. If we assume that the calibration of the instrument is perfect, and adjust the offset in the CHL vs FL plots to force a slope of 1, we find an offset of 0.06 to 0.09ml/l which is a little lower than the dark values observed.

##### Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors were used for 3 other cruises this year: 2006-08, 2006-11 and 2006-10 in January, April and June. The bottle comparisons indicated that the primary salinity was within 0.0013 of the bottles for the first, low by 0.0008 for the second and within 0.001 for the third. The secondary salinity was found to be high by 0.0035, 0.0014, 0.0031, respectively.

- This DO sensor was used for 2006-08, 2006-11 and 2006-10, but the first was a very complex processing job as an error was made in the identification of the sensor so calibration was done twice. However, an analysis of one cast shows that the relationship was something like:

CTD-BOT = 1.08 DOX-CTD - 0.005   (2006-08)

For the other two cruises the results were excellent with tighter fits than seen in the past

CTD-BOT = 1.0663 DOX-CTD + 0.0783 with  offset of -0.03ml/l applied as a second calibration (2006-11 April)

CTD-BOT = 1.0606 DOX-CTD + 0.1138 with offset of -0.017ml/l applied as a second calibration. (2006-10 June)

Historic ranges – There are many excursions from the historic ranges with temperatures a little high in the top 20m in the northern Strait of Georgia and in deeper waters in the southern Strait of Georgia and at one cast in the North-East part of Juan de Fuca Strait. The deep salinity for cast #47 which was a little low, but this probably reflects inadequacies in the climatology. In Juan de Fuca Strait there were many low salinity values between 70 and 100m particularly close to the mouth. This is believed to reflect a real intrusion of cold, fresh water. There are also some low salinity values near the bottom for some of the Juan de Fuca casts. These excursions are not considered indicative of instrument calibration problems.

##### 13. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were on the order of 1s which is the shift that has been used in most other cruises. All data was shifted by +24 records and a few casts were checked afterwards and found to be shifted appropriately. (Output: SHFFL)

Dissolved Oxygen

Based on the recent results of previous uses of this instrument, tests were run using +80, +100 and +120 records on two casts. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by between +100 or +120 records (about 4 to 5s). A choice of +110 records was found best for 2006-08 and 2006-11, so that shift was applied to the DO channel for all casts. After this step a few casts were plotted and the results are good; the offset between downcast and upcast features in DO looks much like those for temperature.

Conductivity

Tests were run on one cast with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.2 for the primary sensor which is the same result as found during 2006-08 and 2006-11. The secondary conductivity looks best with no change as was found during 2006-11, so no shift will be applied. The primary conductivity channel was advanced by -0.2. (Output: *.SHFC).

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range 10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: none

##### DETAILED EDITING

The primary sensors are closer to the bottles and were chosen for 2006-08 and 2006-10. There is more history for the secondary sensors and they have produced salinity that is consistently high, but the differences varied from 0.0012 to 0.0035.

Page plots were produced and used to guide the editing. On-screen plots of descent rate and pump status were also used.

The following casts required fairly heavy editing: 6-10, 12-13, 63 and 68-69.

All other casts required only lightly editing.

Note was made of the editing details in the headers of the relevant files.

For cast #62 there was bad transmissivity data; it was not clear what was good data and what bad, so no attempt was made to edit this channel.

##### Initial Recalibration and Fluorescence Filter

Salinity will not be recalibrated. Previous uses of the sensors suggest they are within 0.0013. From this cruise the average difference suggests that the salinity is high by 0.0025, but the scatter is huge and the average is very dependent on how outliers are identified. This decision should be revisited if future use provides evidence of a larger error.

From section 11 we have the following equation for recalibration of DOX:

DOX-BOT = 1.0866 * DOX-CTD + 0.0412

File 2006-21-recal1.ccf was prepared to apply the above correction to the DO channel in the SAM and MRG files. COMPARE was rerun for dissolved oxygen and when the same points were left out of the comparison as in the first run, the average difference is -0.00015ml/l indicating that the recalibration worked properly; there is little variation with pressure and time. (See 2006-21-dox-comp2.xls.)

The edited downcast files, EDT, were recalibrated using 2006-21-recal1.ccf. (Output:COR1)

The COR1 files were clipped to 100db and set aside to be processed later for Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time, and the comparison with titrated samples those due to drift in calibration. But there remains an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values.

The FIL files were bin-averaged (0.25db bins) & thinned to usual bottle levels. Then CTD DO values were compared with the upcast bottle DO values. Plots were made of differences against pressure and DOX and outliers were excluded. (See 2006-21-dox-comp3.xls.) The resulting fits were used to recalibrate the CTD DO data in the thinned files and COMPARE was rerun. (See 2006-21-dox-comp4.xls.) This process was repeated many times because the results were not satisfying. The best results were found using recalibration file 2006-21-recal2.ccf:

CTD DOX (Corrected) = CTD DOX + 0.00009 *Pressure - 0.041

There is a lot of scatter as was also the case during 2006-11 but at that time there was little pressure dependence:

CTD DOX (Corrected) = CTD DOX - 0.000009*Pressure - 0.03 (See 2006-11-dox-comp3.xls.)

File 2006-21-recal2.ccf was applied to the AVG files, but not to the bottle files since this error does not apply to data collected while stopped. (Output: COR2)

A possible explanation for the pressure dependence is that some stops were quite short and that will affect shallow samples more than deep ones due to higher gradients.

There is a lot of scatter in the comparison but excluding a few outliers suggests that the dissolved oxygen data in the CTD files should be considered ±0.2ml/l to 200m and ±0.1ml/l below that.

##### Special Fluorometer Processing

All COR1 files were put through CLIP to produce files with data to 100db only for the use of A. Peña.

Those files were then bin-averaged (1/4db bins), recalibrated suing 2006-21-recal2.ccf and put through REMOVE to remove extraneous channels. HEADEDIT was run to fix formats and channel names and the final files (FCTD) were saved in a separate directory.

A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data was put through a median filter with fixed width 11, before bin-averaging.

The recalibrated CTD bottle files (SAMCOR1) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.2ml/l from   0 - 200m

•	±0.1ml/l below 200m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~65% to 90% for Juan de Fuca, Haro Strait and Discovery Passage. It was higher in the Strait of Georgia (95 – 120%) and very high in Saanich Inlet (155%) and Boundary Bay (130%). For Saanich Inlet a surface bottle confirms that the dissolved oxygen values are ok, so the high saturation values are assumed to be due to intense biological activity.

As a final check on the files a track plot, cross-reference listing and HEADER CHECK were run. Errors were found in the station name of a few casts; these were fixed.

##### Final Bottle Files

The MRGCLN files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

The flags for cast #62, sample #213 were changed from “d” to “e” as all are consistent with a bottle misfire.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

As a final check a cross-reference listing, a track plot and HEADER CHECK were rerun. This turned up two errors:

Cast #24: No ADD file had been created so the DO data was not merged with the bottle file. The problem probably arose because of errors in the positions in the headers in the OXY file. This was corrected and then an ADD file created.

Cast #76:  the chlorophyll data had not been merged with the bottle file due to an error in naming the event number in the spreadsheet.

After these corrections MERGE was run to combine MRGCOR1 with ADD for cast #24 and with CHL for cast #68; the subsequent steps described above in this section were rerun for those two casts only.

##### 22. Thermosalinograph Data (Note: there were no intake temperature and flow rate channels)

a.) Checking calibrations

There was 1 file containing TSG data. The con file had an error in the temperature calibration and the fluorometer ID was missing and its date wrong. The file was corrected and saved as 2006-21-TSG.con.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Salinity, Fluorescence:Wetlab:Weststar, Julian Days, Latitude, Longitude. It was then converted to IOS SHELL format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that appears to be correct.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

A time-series plot was produced and the only spikes noted were small ones.

An initial track plot was produced and no problems were noted in positions. CTDEDIT was used to examine temperature, salinity and fluorescence but all spikes look real so no editing was done.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or near 3db and exported to a spreadsheet.

The TSG file was opened in EXCEL; temperature, salinity and fluorescence were averaged by taking the median value over a 2 minute window, and the data was then reduced to the times when CTDs were run. (The TSG was not recording when the first cast was run.) Those files were combined in a spreadsheet matching times. The positions were compared and all were very close, with differences less than 0.0003º for both latitude and longitude and average differences of 0.0001º for both. The good correspondence supports the conclusion that the clock and GPS worked well. This spreadsheet will be used in the next step to compare temperature, salinity and fluorescence. (See 2006-21ctd-tsg-comp.xls)

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 There was no intake temperature channel available.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were used to find the differences between the salinity, fluorescence and temperature channels from the CTD and the TSG. Graphs were prepared comparing the TSG temperature and salinity channels with those of the CTD. They show tremendous scatter and larger differences than are expected from this equipment. Based on median values the TSG temperature is high by ~0.5Cº and the salinity low by ~0.3units. The mixed layer was generally very shallow during this cruise, so a small mismatch in depth of intake with the CTD depth will be significant. Only the ratio of TSG fluorescence to CTD fluorescence (~2) is close to the expected value. (See 2006-21ctd-tsg-comp.xls.)

An estimate of surface mixing was made using the IOS SHELL routine SURFACE and the 11 casts with the best mixing were chosen. The information for those casts was extracted from 2006-21ctd-tsg-comp.xls and named TSG-CTD-com-best_mixed.xls. When the casts were arranged according to how well mixed they are, it becomes clear that mixing is a real issue. There is still tremendous variability, but the median of the 5 best-mixed casts indicates the TSG temperature is high by 0.24Cº, the salinity low by 0.11 and the fluorescence ratio ~2.3.

e.) Loop Salinity and chlorophyll comparison

There were 20 loop salinity samples and 19 chlorophyll samples. A spreadsheet (2006-21-TSG-chl-sal-comp.xls) was prepared with the loop sample ID, the salinity and titrated chlorophyll from the analysis spreadsheet and the date and time of collection as noted in the log book. There were problems with a number of the samples.

- Loop 1, 7, 13 have no time or position recorded in the log book.

- Loop 18 – time approximate, no position recorded in the log book.

- Loop 5, 6, 10, 12, 14 and 15 – time in log clearly in error, but positions available. The times were matched with those in the TSG data file, then the position compared with the log entry. All disagreed. So the TSG file was searched for locations where the positions matched the loop log entry, time noted and the salinity and fluorescence values compared with the loop samples. In most cases of time error there was a difference of an integral number of hours suggesting time zone errors. However, for Loop 5 that was not the case. The new times are recorded in file loop-time-position-check.xls.

- Loop 16 – two samples were taken 16 and 16a, but it is believed that they were not at exactly the same time. The first was assumed to be at the time given in the log for loop 16.

When the salinity bottle results are compared with the loop samples there is a lot of variability. To study why there is so much noise in the comparisons late in this cruise, the standard deviation was calculated over 5 records for temperature, salinity and fluorescence in the TSG file. When this was plotted it became clear that the only period with little variation was in the early part of the cruise up to June 14th at 7:45 UTC. (See 2006-21-TSG-variability.xls.)

The average of all samples (excluding Loop 9 and Loop 16) shows the salinity high by 0.06. (Those two samples were excluded because there is too much uncertainty about the time of Loop 16 and Loop 9 was a poor match in salinity and in a very high variability region.) But early in the cruise when the variability was relatively low the TSG salinity appears to be low by 0.03 to 0.06 in 5 out of 6 cases. The differences later in the cruise are large and vary in sign. The ratio of TSG fluorescence to bottle CHL is about 2.4 or 2.5 depending on what data was included. (See 2006-21-TSG-loop-comp.xls.)

f.) Calibration history

The TSG was recalibrated in October 2005.

During 2006-08 the temperature rise in the loop due to heating by the ship was found to be 0.22Cº (February and on the TULLY). During 2006-11 it was found to be high by about 0.20Cº (April, VECTOR).

During 2006-08 the salinity was found to be low by 0.02 and during 2006-11 it was low by 0.022.

g.) Conclusions

Temperature: The history of instrument would suggest that the TSG temperature is higher than that of the CTD by no more than 0.20Cº, and given the warm surface waters, probably less than that. An estimate of 0.18Cº will be made. The comparison with CTD casts is too noisy to be used as a basis for recalibration but the best-mixed of those casts are compatible with that choice.

Salinity: The history of the instrument leads us to expect the TSG salinity to be low by about 0.022; the comparison with loop samples and CTD suggest larger errors, but the comparisons do not look reliable. The TSG salinity will be lowered by 0.22 but this decision should be revisited if analysis of other uses of the instrument indicates problems with this estimate.

Fluorescence: The TSG fluorescence is approximately 2.3 times that of the CTD.

h.) Recalibration

File 2006-21-TSG-recal.ccf was used to apply offsets of -0.18Cº and +0.022 to Temperature:Primary and Salinity:T0:C0.

i.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary and Flag.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and to add the depth of sampling to the header and to add a general comment about the sampling method as well as the following remark:

Fluorescence is nominal.

Temperature and salinity have been recalibrated based primarily on the history of the sensors.

As a final check plots were made of the cruise track and data; no problems were noted.

##### 23. Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun; no errors were detected.

The sensor history was updated for the CTD sensors.

Particulars from logs

10/11 – Log notes problems in time entries

# Institute of Ocean Sciences  CRUISE SUMMARY

| Cruise ID#: 2006-21 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 12 June 2006 End: 18 June 2006 |  |  |  |  |  |  |
| Location: SoG/JdeF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Masson D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2006-21

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 01/02/05 | Factory “ |  |  |
| Conductivity | 1766 | 07/06/05 | “ |  |  |
| Secondary Temp. | 2106 | 08/07/05 | “ |  |  |
| Secondary Cond. | 1729 | 12/07/05 | “ |  |  |
| Transmissometer | 723DR | 28/03/06 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 24/11/05 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| PAR | 4565 | 23/12/04 |  |  |  |
| Surface PAR | 16504 | 02/01/04 |  |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2006-21

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 21/10/05 | Factory | 21/10/2005 |  |
| Conductivity | 2487 | 21/10/05 | “ | 21/10/2005 |  |
| Wetlab WetstarFluorometer | WS3S-713P | 10/01/01 | ? |  |  |