## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2006-26

Agency: PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, B.C.

Location: Queen Charlotte Sound

Project: Acoustics

Party Chief: Cooke K.

Platform: W.E. Ricker

Date: August 15, 2006 – August 29 2006

Processed by: Germaine Gatien

Date of Processing: December 17, 2006 – January 19, 2006

Number of original CTD casts: 15   	Number of CTD casts processed: 14 (1 file surface data only)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Seapoint Fluorometer (#2229) with a 10X cable. Based on what was used during 2006-09, it is assumed that the deck unit was a model 911 (#0471).

# SUMMARY OF QUALITY AND CONCERNS

There was an error in the pressure sensor configuration used at sea; an offset of -0.6db was entered whereas +0.4db would be appropriate.

There was no calibration sampling for salinity. The history of the instrument suggests the salinity data are reliable with errors on the order of ~±0.001, but there were much higher differences between the two salinity channels than usual and those differences varied greatly from cast to cast. Noisy data in the secondary channels may explain this variability.

Fluorescence data are nominal and unedited except that some records were removed in editing temperature and salinity.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

This cruise was part of a joint-ship experiment with the John P. Tully cruise 2006-25. Sheets are glued into the log book for 2006-25 which give details on the CTD casts of 2006-26. There were no event numbers given; the file names include the station name. Consecutive numbers were assigned and entered on the Daily Log sheets, events #1 through #15.

There was no calibration sampling.

The cruise summary sheet was completed.

The histories of the conductivity and pressure sensors were obtained. Calibration constants were checked. Errors were found and corrected in the conductivity and pressure calibration. As has been found many times for this CTD, the pressure offset was set to -0.6db whereas +0.4db has been found appropriate for the past 2 years, so +0.4db was entered for this cruise as well and the date of the pressure calibration was corrected. The new con file was named 2006-26-ctd.con.

##### Conversion of Raw Data

All data were converted using configuration file 2006-25-ctd.con.

A few casts were examined and all expected channels are present. The descent rate was high and quite steady for the casts checked. The two pairs of T and C channels look similar and the upcast data are reasonably close to the downcast data. Cast #8 contains no useful data; this was noted on the log sheet.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

CELLTM was run using (0.03, 7) for the primary and (0.03, 9) for the secondary conductivity for cruises 2006-09 and 2006-38 when the same equipment was used. Tests were done on a few casts to check that those settings worked well for this cruise as well and they did.

CELLTM was run using (0.03, 7) for the primary and (0.03, 9) for the secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Four casts were plotted to check for agreement between the pairs of T and C sensors. The following values are rough estimates from downcast data as the differences were extremely noisy:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 230 | ~0 | ~0 | ~0 | High, fairly steady |
| 9 | 400 | -0.0002 | -0.0012 | -0.013 | High, fairly steady |
| 10 | 400 | -0.0002 | -0.0008 | -0.009 | High, fairly steady |
| 13 | 130 | ~0 noisy | -0.0002 | -0.003 | High, fairly steady |

The conductivity and salinity differences are much larger after the first few casts than during 2006-09 and 2006-38. However, the changes seem random with time.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

Station names were added after conversion based on the log entries.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and a few errors in station names were found and corrected.

The header summary was run and one inconsistency was found. The first cast is given as 01:40 in the log book and 20:47 in the GPS time in the file. This can be explained if the log time is 01:40pm PDT which requires adding 12 hours plus 7 hours to get UTC time. All other times are in agreement. No change was made to the header as it is believed to be correct.

The track plot was produced and looks reasonable and in agreement with a plot in the cruise report for 2006-25, with the exception of the first cast which does not appear on the cruise report plot.

The average surface pressure is 2.8db. The range was from 1.4 to 5.1db. The lower values are associated with very low salinity but the pumps were not turned on at those depths. A few casts were checked and once the pumps were turned on the salinity assumed expected values, so there is no indication of problems with the pressure calibration.

##### 10. SHIFT

Conductivity

For 2006-09 and 2006-22 the primary sensors were selected for archiving and the conductivity was advanced by -0.5s to minimize unstable features in T-S plots. A check of a few casts shows that the secondary channels are very noisy, so once again the primary will be selected. (The noise in the secondary channel probably accounts for the odd differences noted in section 7.) Tests were run on two casts from this cruise and the same setting was found appropriate.

All casts were put through SHIFT to advance the primary conductivity by -0.5s.

Fluorescence

A few casts were examined to compare the offset between the upcast and downcast fluorescence with that of the temperature traces. That difference divided by the sum of the average descent rate and ascent rate gives an estimate of the shift needed to align the fluorescence. For this cruise estimates varied from 1.0s to 1.3s.  A shift of +24 records (1s) was applied as is usual for this equipment.

##### 11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Pressure Minimum

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### 12. DETAILED EDITING

The primary temperature and salinity were chosen for further processing since they have been selected for other recent uses of this equipment and the primary salinity is less noisy than the secondary.

Page plots were produced using (T0, S0). These were used to guide the editing.

On-screen plots of descent rate and pump status were also used to guide editing.

All casts required some editing.

Note was made of the editing details in the headers of the relevant files.

##### 13. Other Comparisons

Previous experience with these sensors –

- The primary sensors were used for many cruises during 2005 and for 5 during 2006. Salinity was found to be within 0.002 except for one cruise in 2005 when it was low by about 0.003; in most cases it was within 0.001. There was a lot of scatter and few deep bottles for all previous comparisons.

- The secondary sensors were used during 2005-16 and 5 cruises during 2006. There were few bottles and a lot of scatter in the comparisons but for 2006-22 it was believed that the errors were <0.001. Earlier cruises suggest larger errors ~±0.003.

Historic ranges – Profile plots were made of T and S with local climatology superimposed; all data fell within the ranges.

Intercomparison with 2006-25 – Cruise 2006-26 overlapped with this cruise but no casts were occupied by both. T-S plots were produced with nearby casts plotted together. Nearby casts compare reasonably well, with the distance from shore being generally more significant than closeness.

##### 14. Initial Recalibration

There was no calibration sampling and the history of the instrument suggests that the salinity is good to ±0.001. Recalibration is not justified.

##### 15. Fluorescence Processing

The EDT files were put through a median filter, size 11, applied to the fluorescence channel only. One cast was examined before and after, and the filter was found to have worked well. (Output:FIL)

##### 16. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. No problems were noted.

Temperature, Salinity and Fluorescence profiles were plotted on-screen and no problems noted.

##### 17. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

HEADER EDIT was used to fix formats and channel names, to correct the geographic description and to add the following comments:

Fluorescence data are nominal and unedited except that some

records were removed in editing temperature and salinity.

There was no calibration sampling. Based on the history of the

conductivity sensor, salinity data are considered to be reliable,

and likely to be within +/- 0.001.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

An final examination of page plots showed that there was a surface record in cast #11 that was meant to be removed by CTDEDIT. This was removed using a text editor and the file was put through CLEAN to reset the headers; the output was then renamed as 2006-26-0011.ctd.

##### 18. Producing final files

A cross-reference listing was produced for CTD files.

HEADER CHECK was rerun and a cruise track plotted and no errors found.

The sensor history was updated for the conductivity and pressure sensors.

Particulars (notes from log sheets)

8. File empty

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2006-26 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 15 August 2006 End: 29 August 2006 |  |  |  |  |  |  |
| Location: Queen Charlotte Sound |  |  |  |  |  |  |
| Vessel: W.E.Ricker Party Chief: Cooke K. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0585 | No | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585         Cruise ID#:		2006-26

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19/03/05 | Factory |  |  |
| Conductivity | 3038 | 03/03/05 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 2102 | 07/06/05 | “ |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/2000 | Factory |  |  |