## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2006-27

Agency: IOS, Ocean Sciences Division, Sidney, B.C.

Project: Drifter Studies

Chief Scientist: Stuchhi D.

Platform: Buckshot

Date: May 1, 2006 – May 5, 2006

Processed by: Germaine Gatien

Date of Processing: 3 August 2006 – 17 August 2006

Number of original CTD casts: 20

Number of casts processed:  20

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (#0334) was run with pressure sensor 290544; an SBE43 dissolved oxygen sensor (#0047), a transmissometer (S/N 498DR) and a fluorometer (SeaPoint S/N 2228) were also mounted on the CTD.

# SUMMARY OF QUALITY AND CONCERNS

As noted at sea the data could not be converted with the sensor numbers entered correctly. It is believed that the CTD and sensors were the same as used for Legs 2 and 3 of 2006-20. The hex files were edited to enter the expected sensor serial numbers and the converted data looks reasonable.

There was no calibration sampling during this mission. Results of other cruises were used.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

The Daily Log was obtained. It is believed that the CTD configuration was the same as for Legs 2 and 3 of 2006-20, but the chief scientist noted in the log that the hex files could not be converted using the configuration file prepared for the cruise. As was noted during 2006-20 Leg3 there was a problem in the headers so that the temperature and conductivity sensors were entered incorrectly. A text editor was used to change the sensor entries in the headers and then conversion was possible. The appropriate configuration file was saved as 2006-27CTD.con.

A log note indicates the syringe was dry for cast #1. The CTD hit bottom during cast #14.

The times in the log book are in PDT.

The file names were in standard format.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

All hex files were converted using 2006-27CTD.con.

A few casts were plotted. The pressure signal has steps of about 0.2db as is usual for this CTD; the manufacturer states the resolution is 1db.

All variables look normal. The descent rate is sometimes low, and occasionally noisy.

4. WILDEDIT

WILDEDIT was run on all casts on pressure, temperature and conductivity channels using 2,20,25,0 for “Standard deviations for pass 1” and “Standard deviations for pass 2”, scans per block and “Keep data within this distance of the mean”.

5. ALIGN

This step was skipped. SHIFT will be used to align T and C.

6. WFILTER

Based on the results of many other cruises using this equipment, the SeaSoft routine WFILTER was run for all casts to apply a cosine filter, size 5, to the pressure, temperature and conductivity.

7. CELLTM

Based on the results of 2006-20 and other recent uses of this equipment CELLTM was run on all casts using α = 0.03, 1/β = 9.0.

8. DERIVE

Program DERIVE was run to calculate salinity.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert Sea-Bird ASCII data to IOS Headers.

The time is said to be UTC in the headers but is, in fact, PDT according to the log book. As found during 2006-20 the date was ahead by 1 day. To correct for these two errors, ADD TIME CHANNEL was used to subtract 17 hours from the times (-24 for the date error +7 to convert to UTC).

At this point a text editor was used to add entries for Event Number, Station Name, Latitude, Longitude and Water Depth based on log book entries.

10. Checking Headers

Header Check was run and turned up no obvious errors.

A Header Summary was produced and checked against the log.

The Surface Check produces an average of 0.22db.

Track plots were produced and added to the end of this report.

Plots were examined on-screen of all variables. The temperature and salinity seem reasonably close for upcast and downcast. The fluorescence looks odd for a number of casts and should be examined more closely later. The DO range is small, but the usual sort of offsets are seen near the surface; for one cast the DO looks wrong in the same range as fluorescence behaved strangely. The transmissivity is noisy and frequently quite low at the surface and near the bottom.

11. SHIFT

Conductivity

On previous cruises using this type of CTD good results were found when SHIFT was run to advance the conductivity by +0.7 records. Tests were run on 2 casts with noisy T-S curves to ensure that the same setting works well for this data set and it did.

SHIFT was run on all casts advancing the conductivity by +0.7 records.

Fluorescence

Based on the results of 2006-20 the fluorescence was aligned by applying a shift of +42 records to all casts. The data was then examined to see if the results were satisfactory. The offset between downcast and upcast trace looks similar to that for temperature after this step. Some features that previously looked like noise look believable after the shift. One interesting cast was Event #13. There was a spike in the fluorescence during the downcast. Before running the SHIFT routine it looks out of step with the transmissivity and upcast fluorescence. But after SHIFT it corresponds to some spikes in each of those channels and lines up with a sharp step in temperature. The downcast spike is large, probably better defined. A lower SHIFT setting was not sufficient to line up these features. For cast 14 there is an unbelievable spike, but it occurs during the upcast, as do several instances of zero values near the surface. For cast #10 there is a spike deep in the downcast.

Dissolved Oxygen

This dissolved oxygen sensor is noted for poor time response and hysteresis. For this cruise the fairly low descent rate may be helpful in getting better data. To correct for the time response we generally shift the channel by 100 to 160 records, but that is for an SBE 911+ with a higher sampling rate. For 2006-20 a setting of +40 records gave the best overall results but the descent rate was low during that mission. For this cruise a higher setting is expected. Tests were done using shifts of +40, +50, +60 and +80. The results varied with the depth, probably reflecting the descent rate being very low at the bottom of the casts. The best results overall were with +50 so that was applied to all casts.

- DELETE

The files were put through DELETE using the following parameters:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00    Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted

Sample interval taken from the header.  Pressure was not filtered.

COMMENTS ON WARNINGS: There were no warnings.

- DETAILED EDITING

Page plots were produced and examined for spikes and instabilities and used to guide the use of CTDEDIT. All casts required some editing. Most involved removing data from near the surface before the pumps started operating. Note was made of the editing details in the files.

It was later found that cast #12 needed further editing. CTDEDIT was used to remove 1 more surface record, and then all other steps from this point onwards were rerun.

- BIN AVERAGE

The following Bin Average values were applied to the ED1 files:

Bin channel = pressure 		Averaging interval = 1db

Minimum bin value =   .000	Average value will be used

Interpolated values are NOT used for empty bins.

- Intercomparisons

Sensor History

- Dissolved Oxygen

During 2005-32 the following correction was found:

DO(corrected) = 1.4614 * DOX_CTD – 0.032

A second correction was pressure-dependent:

DO(corrected) = CTD DOX +0.0004*Pressure – 0.1019

- Salinity

There were 6 salinity samples gathered using this equipment during 2006-20 between April and May 2006. The CTD salinity was found to be higher than the bottles by 0.041. It was believed unlikely that the CTD could be that far off, leaving some doubts about the sampling. 
During 2006-10 in May 2006 there were only 4 salinity calibration samples and one was an outlier in the comparison. The average of the other 3 samples showed the CTD salinity to be high by 0.0044. This result was used for recalibration of both 2006-10 and 2006-20.

Historic ranges –There is no local climatology available for this area.

- Recalibration

There was no calibration sampling during this cruise so recalibration files 2006-27-recal1.ccf and 2006-17-recal2.ccf were prepared based on previous cruises:

Dissolved Oxygen – use results of cruise 2005-32 (2-part correction.)

Salinity – use results of cruise 2006-10 Leg1.

At this stage plots were produced of dissolved oxygen versus salinity to check for outliers. The only problem noted was at the surface of cast #12. The edited file was examined and it appears that one more data record should have been removed due to pumps being off (as indicated by negative dissolved oxygen voltage.) Cast #12 was edited again and then bin-averaged and recalibrated.

- REMOVE

The following channels were removed from all casts: Scan_number; Conductivity:Primary, Oxygen:Voltage:SBE, Descent Rate and Flag.

- HEADER EDIT

Header Edit was used to fix formats and channel names and to add the following comment:

There was no calibration sampling during this mission.

Salinity calibration is based on calibration sampling during 2006-10 Leg 1.

Dissolved oxygen calibration is based on the results of 2005-32 for which

errors were estimated to be ±0.5ml/l between 0 and 50m, ±0.3ml/l between

50 and 150m and ±0.1ml/l below 50m. There are likely to be significant

differences for this cruise due to calibration drift and a lower descent

rate for the CTD.

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

A note was added to cast #1 to mention the log note about the syringe being dry.

Standards Check was run and Header Edit rerun until all problems were resolved.

- Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

c.) The conductivity and pressure sensor history files were updated.

Particulars

1. Syringe was dry.

14. CTD hit bottom.

## Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2006-27

Dates: Start: 1 May 2006 End: 5 May 2006

Location: Broughton Archipelago

Vessel: Buckshot

Party Chief: Stucchi D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334		Cruise ID#:		2006-27

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 290544 | 13May05 | Factory |  |  |
| Temperature | 4054 | 14Oct04 | Factory |  |  |
| Conductivity | 2754 | 15Oct04 | Factory |  |  |
| Dissolved Oxygen | 0047 | 10Feb05 | Factory |  |  |
| Transmissometer | 498DR | 28Mar06 | IOS |  |  |
| Fluorometer | 2228 | ? | ? |  |  |