## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.3 | Run at 2007/01/12 13:33:06 | Compare Excel Add-in Version 2.1
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 0.2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2006-29-0002.mrgtem | 2006-29-0002.sam | 121 | 200.248 | 200.248 | 0.035739 | 0.00019836 | 32.1969 | 32.1305 | 0.0085776 | -0.066402 | 32.1969 | 32.1319 | 0.0083647 | -0.064991
2006-29-0011.mrgtem | 2006-29-0011.sam | 82 | 150.437 | 150.445 | 0.11756 | 0.0084992 | 33.9001 | 33.8985 | 0.0002889 | -0.0016213 | 33.9001 | 33.9013 | 0.00023117 | 0.0011597
2006-29-0012.mrgtem | 2006-29-0012.sam | 30 | 200.868 | 200.845 | 0.10693 | -0.023392 | 33.9465 | 33.9433 | 0.00026063 | -0.0032463 | 33.9465 | 33.9462 | 0.00027001 | -0.00026321
2006-29-0015.mrgtem | 2006-29-0015.sam | 19 | 200.605 | 200.581 | 0.11365 | -0.023834 | 33.9685 | 33.967 | 0.00028125 | -0.001545 | 33.9685 | 33.9696 | 0.00021785 | 0.001091
2006-29-0023.mrgtem | 2006-29-0023.sam | 103 | 100.065 | 100.067 | 0.11653 | 0.0018539 | 32.9723 | 32.8974 | 0.0029654 | -0.074921 | 32.9723 | 32.9008 | 0.0018385 | -0.071484
2006-29-0027.mrgtem | 2006-29-0027.sam | 70 | 100.481 | 100.531 | 0.09938 | 0.050438 | 32.4119 | 32.2631 | 0.0025384 | -0.14883 | 32.4119 | 32.2661 | 0.002939 | -0.14585
2006-29-0032.mrgtem | 2006-29-0032.sam | 117 | 200.167 | 200.163 | 0.099142 | -0.0043945 | 31.7276 | 31.717 | 0.0034301 | -0.010639 | 31.7276 | 31.7203 | 0.0024164 | -0.0072842
2006-29-0042.mrgtem | 2006-29-0042.sam | 120 | 153.342 | 153.344 | 0.11378 | 0.0019684 | 30.938 | 30.948 | 0.002881 | 0.010029 | 30.938 | 30.9504 | 0.0020069 | 0.01243
2006-29-0044.mrgtem | 2006-29-0044.sam | 121 | 299.986 | 299.986 | 0.09043 | -0.00015259 | 31.2592 | 31.2484 | 0.00083674 | -0.010838 | 31.2592 | 31.2521 | 0.00065691 | -0.0071449
2006-29-0048.mrgtem | 2006-29-0048.sam | 84 | 350.205 | 350.188 | 0.098986 | -0.016968 | 31.4145 | 31.4027 | 0.0013091 | -0.01182 | 31.4145 | 31.4048 | 0.0010248 | -0.0096512
2006-29-0051.mrgtem | 2006-29-0051.sam | 121 | 300.406 | 300.406 | 0.10421 | -0.00018311 | 31.2459 | 31.2425 | 0.0002086 | -0.0033531 | 31.2459 | 31.2452 | 0.00022453 | -0.00068283
2006-29-0055.mrgtem | 2006-29-0055.sam | 121 | 250.234 | 250.234 | 0.070324 | -0.00024414 | 31.208 | 31.1965 | 0.0012038 | -0.01152 | 31.208 | 31.1995 | 0.0011619 | -0.0085373
2006-29-0058.mrgtem | 2006-29-0058.sam | 121 | 343.089 | 343.089 | 0.036924 | -0.00042725 | 31.2449 | 31.2759 | 0.00022196 | 0.030954 | 31.2449 | 31.2766 | 0.00021882 | 0.031713
2006-29-0060.mrgtem | 2006-29-0060.sam | 121 | 300.39 | 300.39 | 0.036521 | 9.1553e-05 | 31.1938 | 31.1894 | 0.00036679 | -0.0043507 | 31.1938 | 31.1902 | 0.00026756 | -0.0035858
2006-29-0062.mrgtem | 2006-29-0062.sam | 121 | 175.212 | 175.212 | 0.041375 | -0.00045776 | 30.951 | 30.9329 | 0.0002802 | -0.018103 | 30.951 | 30.934 | 0.00019136 | -0.016991
2006-29-0065.mrgtem | 2006-29-0065.sam | 121 | 150.712 | 150.712 | 0.072317 | 0.00038147 | 30.8976 | 30.9133 | 0.00080453 | 0.015747 | 30.8976 | 30.9147 | 0.00051598 | 0.01708
2006-29-0068.mrgtem | 2006-29-0068.sam | 121 | 250.479 | 250.479 | 0.082767 | 9.1553e-05 | 30.8095 | 30.805 | 0.00025051 | -0.0044766 | 30.8095 | 30.8061 | 0.00023092 | -0.0034485
2006-29-0070.mrgtem | 2006-29-0070.sam | 121 | 300.295 | 300.295 | 0.037683 | 0.00027466 | 30.802 | 30.8001 | 0.00028453 | -0.0018711 | 30.802 | 30.8006 | 0.00023156 | -0.0014458
2006-29-0073.mrgtem | 2006-29-0073.sam | 121 | 125.085 | 125.085 | 0.035247 | 0.00020599 | 30.5 | 30.4359 | 0.00167 | -0.064123 | 30.5 | 30.4371 | 0.0015792 | -0.062946
2006-29-0080.mrgtem | 2006-29-0080.sam | 121 | 351.79 | 351.79 | 0.037555 | -3.0518e-05 | 31.2726 | 31.2975 | 0.00063321 | 0.024942 | 31.2726 | 31.2984 | 0.0003145 | 0.025795
Create Fit Sheet

## File List
Number of Files= |  | 20
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2006-29-0002.mrgtem | 2006-29-0002.sam | 12 | 12 | 2006-29-0002.mrgtem & 2006-29-0002.sam
2 | 2006-29-0011.mrgtem | 2006-29-0011.sam | 14 | 14 | 2006-29-0011.mrgtem & 2006-29-0011.sam
3 | 2006-29-0012.mrgtem | 2006-29-0012.sam | 16 | 16 | 2006-29-0012.mrgtem & 2006-29-0012.sam
4 | 2006-29-0015.mrgtem | 2006-29-0015.sam | 18 | 18 | 2006-29-0015.mrgtem & 2006-29-0015.sam
5 | 2006-29-0023.mrgtem | 2006-29-0023.sam | 20 | 20 | 2006-29-0023.mrgtem & 2006-29-0023.sam
6 | 2006-29-0027.mrgtem | 2006-29-0027.sam | 22 | 22 | 2006-29-0027.mrgtem & 2006-29-0027.sam
7 | 2006-29-0032.mrgtem | 2006-29-0032.sam | 24 | 24 | 2006-29-0032.mrgtem & 2006-29-0032.sam
8 | 2006-29-0042.mrgtem | 2006-29-0042.sam | 26 | 26 | 2006-29-0042.mrgtem & 2006-29-0042.sam
9 | 2006-29-0044.mrgtem | 2006-29-0044.sam | 28 | 28 | 2006-29-0044.mrgtem & 2006-29-0044.sam
10 | 2006-29-0048.mrgtem | 2006-29-0048.sam | 30 | 30 | 2006-29-0048.mrgtem & 2006-29-0048.sam
11 | 2006-29-0051.mrgtem | 2006-29-0051.sam | 32 | 32 | 2006-29-0051.mrgtem & 2006-29-0051.sam
12 | 2006-29-0055.mrgtem | 2006-29-0055.sam | 34 | 34 | 2006-29-0055.mrgtem & 2006-29-0055.sam
13 | 2006-29-0058.mrgtem | 2006-29-0058.sam | 36 | 36 | 2006-29-0058.mrgtem & 2006-29-0058.sam
14 | 2006-29-0060.mrgtem | 2006-29-0060.sam | 38 | 38 | 2006-29-0060.mrgtem & 2006-29-0060.sam
15 | 2006-29-0062.mrgtem | 2006-29-0062.sam | 40 | 40 | 2006-29-0062.mrgtem & 2006-29-0062.sam
16 | 2006-29-0065.mrgtem | 2006-29-0065.sam | 42 | 42 | 2006-29-0065.mrgtem & 2006-29-0065.sam
17 | 2006-29-0068.mrgtem | 2006-29-0068.sam | 44 | 44 | 2006-29-0068.mrgtem & 2006-29-0068.sam
18 | 2006-29-0070.mrgtem | 2006-29-0070.sam | 46 | 46 | 2006-29-0070.mrgtem & 2006-29-0070.sam
19 | 2006-29-0073.mrgtem | 2006-29-0073.sam | 48 | 48 | 2006-29-0073.mrgtem & 2006-29-0073.sam
20 | 2006-29-0080.mrgtem | 2006-29-0080.sam | 50 | 50 | 2006-29-0080.mrgtem & 2006-29-0080.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 20
2 | (Sal1 - Sal_Bot) | Pressure | 20
3 | (Sal0 - Sal_Bot) | File Pair # | 20
4 | (Sal1 - Sal_Bot) | File Pair # | 20

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.00088695 |  | X | CALC Y
m= | 2.6081e-06 | 3.3085e-05 |  | Std Dev of Resid |  | 0.0043308 |  | 200.163 | -0.0059054
b= | -0.0064275 | 0.0085898 |  | Resid Sum of Squares |  | 0.00013129 |  | 300.406 | -0.005644
# of Values Extracted = |  |  | 20
# of Values Used in Fit = |  |  | 9
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 200.248 | 32.1305 | 0.0085776 | -0.066402 | -0.0059052 | 0.060497 |  | 200.845 | -0.0032463
2 | NO | 150.445 | 33.8985 | 0.0002889 | -0.0016213 | -0.0060351 | -0.0044139 |  | 200.581 | -0.001545
3 | YES | 200.845 | 33.9433 | 0.00026063 | -0.0032463 | -0.0059037 | -0.0026574 |  | 200.163 | -0.010639
4 | YES | 200.581 | 33.967 | 0.00028125 | -0.001545 | -0.0059043 | -0.0043594 |  | 299.986 | -0.010838
5 | NO | 100.067 | 32.8974 | 0.0029654 | -0.074921 | -0.0061665 | 0.068754 |  | 300.406 | -0.0033531
6 | NO | 100.531 | 32.2631 | 0.0025384 | -0.14883 | -0.0061653 | 0.14266 |  | 250.234 | -0.01152
7 | YES | 200.163 | 31.717 | 0.0034301 | -0.010639 | -0.0059054 | 0.0047338 |  | 300.39 | -0.0043507
8 | NO | 153.344 | 30.948 | 0.002881 | 0.010029 | -0.0060275 | -0.016056 |  | 250.479 | -0.0044766
9 | YES | 299.986 | 31.2484 | 0.00083674 | -0.010838 | -0.0056451 | 0.0051925 |  | 300.295 | -0.0018711
10 | NO | 350.188 | 31.4027 | 0.0013091 | -0.01182 | -0.0055142 | 0.0063056
11 | YES | 300.406 | 31.2425 | 0.0002086 | -0.0033531 | -0.005644 | -0.0022909 |  |  | -0.00576
12 | YES | 250.234 | 31.1965 | 0.0012038 | -0.01152 | -0.0057748 | 0.0057456
13 | NO | 343.089 | 31.2759 | 0.00022196 | 0.030954 | -0.0055327 | -0.036487
14 | YES | 300.39 | 31.1894 | 0.00036679 | -0.0043507 | -0.005644 | -0.0012934
15 | NO | 175.212 | 30.9329 | 0.0002802 | -0.018103 | -0.0059705 | 0.012132
16 | NO | 150.712 | 30.9133 | 0.00080453 | 0.015747 | -0.0060344 | -0.021782
17 | YES | 250.479 | 30.805 | 0.00025051 | -0.0044766 | -0.0057742 | -0.0012977
18 | YES | 300.295 | 30.8001 | 0.00028453 | -0.0018711 | -0.0056443 | -0.0037732
19 | NO | 125.085 | 30.4359 | 0.00167 | -0.064123 | -0.0061013 | 0.058022
20 | NO | 351.79 | 31.2975 | 0.00063321 | 0.024942 | -0.00551 | -0.030452
NO
NO
YES
YES
NO
NO
YES
NO
YES
NO
YES
YES
NO
YES
NO
NO
YES
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.011987 |  | X | CALC Y
m= | -8.225e-06 | 2.8223e-05 |  | Std Dev of Resid |  | 0.0036943 |  | 200.163 | -0.0030193
b= | -0.0013729 | 0.0073274 |  | Resid Sum of Squares |  | 9.5537e-05 |  | 300.406 | -0.0038438
# of Values Extracted = |  |  | 20
# of Values Used in Fit = |  |  | 9
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 200.248 | 32.1319 | 0.0083647 | -0.064991 | -0.00302 | 0.061971 |  | 200.845 | -0.00026321
2 | NO | 150.445 | 33.9013 | 0.00023117 | 0.0011597 | -0.0026103 | -0.00377 |  | 200.581 | 0.001091
3 | YES | 200.845 | 33.9462 | 0.00027001 | -0.00026321 | -0.0030249 | -0.0027616 |  | 200.163 | -0.0072842
4 | YES | 200.581 | 33.9696 | 0.00021785 | 0.001091 | -0.0030227 | -0.0041137 |  | 299.986 | -0.0071449
5 | NO | 100.067 | 32.9008 | 0.0018385 | -0.071484 | -0.002196 | 0.069288 |  | 300.406 | -0.00068283
6 | NO | 100.531 | 32.2661 | 0.002939 | -0.14585 | -0.0021998 | 0.14365 |  | 250.234 | -0.0085373
7 | YES | 200.163 | 31.7203 | 0.0024164 | -0.0072842 | -0.0030193 | 0.0042649 |  | 300.39 | -0.0035858
8 | NO | 153.344 | 30.9504 | 0.0020069 | 0.01243 | -0.0026342 | -0.015064 |  | 250.479 | -0.0034485
9 | YES | 299.986 | 31.2521 | 0.00065691 | -0.0071449 | -0.0038403 | 0.0033046 |  | 300.295 | -0.0014458
10 | NO | 350.188 | 31.4048 | 0.0010248 | -0.0096512 | -0.0042532 | 0.005398
11 | YES | 300.406 | 31.2452 | 0.00022453 | -0.00068283 | -0.0038438 | -0.0031609 |  |  | -0.0034779
12 | YES | 250.234 | 31.1995 | 0.0011619 | -0.0085373 | -0.0034311 | 0.0051062
13 | NO | 343.089 | 31.2766 | 0.00021882 | 0.031713 | -0.0041948 | -0.035908
14 | YES | 300.39 | 31.1902 | 0.00026756 | -0.0035858 | -0.0038436 | -0.0002578
15 | NO | 175.212 | 30.934 | 0.00019136 | -0.016991 | -0.002814 | 0.014177
16 | NO | 150.712 | 30.9147 | 0.00051598 | 0.01708 | -0.0026125 | -0.019693
17 | YES | 250.479 | 30.8061 | 0.00023092 | -0.0034485 | -0.0034331 | 1.5388e-05
18 | YES | 300.295 | 30.8006 | 0.00023156 | -0.0014458 | -0.0038428 | -0.0023971
19 | NO | 125.085 | 30.4371 | 0.0015792 | -0.062946 | -0.0024017 | 0.060545
20 | NO | 351.79 | 31.2984 | 0.0003145 | 0.025795 | -0.0042664 | -0.030061
NO
NO
YES
YES
NO
NO
YES
NO
YES
NO
YES
YES
NO
YES
NO
NO
YES
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0060834 |  | X | CALC Y
m= | 5.9445e-05 | 0.00028719 |  | Std Dev of Resid |  | 0.0043195 |  | 3 | -0.0062091
b= | -0.0063875 | 0.003356 |  | Resid Sum of Squares |  | 0.00013061 |  | 18 | -0.0053175
# of Values Extracted = |  |  | 20
# of Values Used in Fit = |  |  | 9
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 200.248 | 32.1305 | 0.0085776 | -0.066402 | -0.006328 | 0.060074 |  | 3 | -0.0032463
2 | NO | 150.445 | 33.8985 | 0.0002889 | -0.0016213 | -0.0062686 | -0.0046473 |  | 4 | -0.001545
3 | YES | 200.845 | 33.9433 | 0.00026063 | -0.0032463 | -0.0062091 | -0.0029628 |  | 7 | -0.010639
4 | YES | 200.581 | 33.967 | 0.00028125 | -0.001545 | -0.0061497 | -0.0046047 |  | 9 | -0.010838
5 | NO | 100.067 | 32.8974 | 0.0029654 | -0.074921 | -0.0060902 | 0.06883 |  | 11 | -0.0033531
6 | NO | 100.531 | 32.2631 | 0.0025384 | -0.14883 | -0.0060308 | 0.1428 |  | 12 | -0.01152
7 | YES | 200.163 | 31.717 | 0.0034301 | -0.010639 | -0.0059713 | 0.0046679 |  | 14 | -0.0043507
8 | NO | 153.344 | 30.948 | 0.002881 | 0.010029 | -0.0059119 | -0.015941 |  | 17 | -0.0044766
9 | YES | 299.986 | 31.2484 | 0.00083674 | -0.010838 | -0.0058525 | 0.0049851 |  | 18 | -0.0018711
10 | NO | 350.188 | 31.4027 | 0.0013091 | -0.01182 | -0.005793 | 0.0060268
11 | YES | 300.406 | 31.2425 | 0.0002086 | -0.0033531 | -0.0057336 | -0.0023804
12 | YES | 250.234 | 31.1965 | 0.0012038 | -0.01152 | -0.0056741 | 0.0058463
13 | NO | 343.089 | 31.2759 | 0.00022196 | 0.030954 | -0.0056147 | -0.036569
14 | YES | 300.39 | 31.1894 | 0.00036679 | -0.0043507 | -0.0055552 | -0.0012046
15 | NO | 175.212 | 30.9329 | 0.0002802 | -0.018103 | -0.0054958 | 0.012607
16 | NO | 150.712 | 30.9133 | 0.00080453 | 0.015747 | -0.0054363 | -0.021183
17 | YES | 250.479 | 30.805 | 0.00025051 | -0.0044766 | -0.0053769 | -0.00090035
18 | YES | 300.295 | 30.8001 | 0.00028453 | -0.0018711 | -0.0053175 | -0.0034463
19 | NO | 125.085 | 30.4359 | 0.00167 | -0.064123 | -0.005258 | 0.058865
20 | NO | 351.79 | 31.2975 | 0.00063321 | 0.024942 | -0.0051986 | -0.030141
NO
NO
YES
YES
NO
NO
YES
NO
YES
NO
YES
YES
NO
YES
NO
NO
YES
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.031196 |  | X | CALC Y
m= | -0.00011547 | 0.00024322 |  | Std Dev of Resid |  | 0.0036582 |  | 3 | -0.0026055
b= | -0.0022591 | 0.0028422 |  | Resid Sum of Squares |  | 9.3679e-05 |  | 18 | -0.0043376
# of Values Extracted = |  |  | 20
# of Values Used in Fit = |  |  | 9
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 200.248 | 32.1319 | 0.0083647 | -0.064991 | -0.0023745 | 0.062616 |  | 3 | -0.00026321
2 | NO | 150.445 | 33.9013 | 0.00023117 | 0.0011597 | -0.00249 | -0.0036497 |  | 4 | 0.001091
3 | YES | 200.845 | 33.9462 | 0.00027001 | -0.00026321 | -0.0026055 | -0.0023423 |  | 7 | -0.0072842
4 | YES | 200.581 | 33.9696 | 0.00021785 | 0.001091 | -0.002721 | -0.003812 |  | 9 | -0.0071449
5 | NO | 100.067 | 32.9008 | 0.0018385 | -0.071484 | -0.0028364 | 0.068647 |  | 11 | -0.00068283
6 | NO | 100.531 | 32.2661 | 0.002939 | -0.14585 | -0.0029519 | 0.1429 |  | 12 | -0.0085373
7 | YES | 200.163 | 31.7203 | 0.0024164 | -0.0072842 | -0.0030674 | 0.0042168 |  | 14 | -0.0035858
8 | NO | 153.344 | 30.9504 | 0.0020069 | 0.01243 | -0.0031828 | -0.015613 |  | 17 | -0.0034485
9 | YES | 299.986 | 31.2521 | 0.00065691 | -0.0071449 | -0.0032983 | 0.0038466 |  | 18 | -0.0014458
10 | NO | 350.188 | 31.4048 | 0.0010248 | -0.0096512 | -0.0034138 | 0.0062374
11 | YES | 300.406 | 31.2452 | 0.00022453 | -0.00068283 | -0.0035293 | -0.0028464
12 | YES | 250.234 | 31.1995 | 0.0011619 | -0.0085373 | -0.0036447 | 0.0048926
13 | NO | 343.089 | 31.2766 | 0.00021882 | 0.031713 | -0.0037602 | -0.035474
14 | YES | 300.39 | 31.1902 | 0.00026756 | -0.0035858 | -0.0038757 | -0.00028987
15 | NO | 175.212 | 30.934 | 0.00019136 | -0.016991 | -0.0039912 | 0.013
16 | NO | 150.712 | 30.9147 | 0.00051598 | 0.01708 | -0.0041066 | -0.021187
17 | YES | 250.479 | 30.8061 | 0.00023092 | -0.0034485 | -0.0042221 | -0.00077362
18 | YES | 300.295 | 30.8006 | 0.00023156 | -0.0014458 | -0.0043376 | -0.0028918
19 | NO | 125.085 | 30.4371 | 0.0015792 | -0.062946 | -0.0044531 | 0.058493
20 | NO | 351.79 | 31.2984 | 0.0003145 | 0.025795 | -0.0045685 | -0.030364
NO
NO
YES
YES
NO
NO
YES
NO
YES
NO
YES
YES
NO
YES
NO
NO
YES
YES
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE