## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2006-33

Agency: OSD

Location: North-East Pacific

Project: Effingham, LaPerouse, ECOHAB, Aquaculture, Victoria Sill

Party Chief: Juhasz T.

Platform: John P. Tully

Date: October 10, 2006 – October 20, 2006

Processed by: Germaine Gatien

Date of Processing: January 19, 2007 –  February 1, 2007

Number of original CTD casts: 88   	Number of CTD casts processed: 88*

(*Cast #12 contains only pressure and transmissivity since pump was off)

Number of bottle casts: 32                  	Number of bottle casts processed: 32

TSG files: 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea/Seatech transmissometer (#723DR), a SeaBird DO sensor #0047 (on the primary pump) and a Seapoint Fluorometer (#2228) with a 10X cable (on the secondary pump). The deck unit was a model 911 (#0508) and the logging computer was a Shuttle Case. The salinometer used was an Autosal 8400B S/N 68572. A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P).

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order except for a lack of details on the TSG system.

The pumps were not turned on for cast #12. The pumps did not perform well during cast #17 due to the CTD hitting the bottom and wood chips entering the pump and to avoid similar problems the pumps were turned off for the upcast of #20.

While most bottle stops were at least 30s long, there were some that were significantly shorter. The descent rate of the CTD was generally kept quite high.

The differences between bottle salinity and CTD salinity are out of line with the history of these sensors. Recalibration of salinity was based on cruise 2006-15, pending a post-cruise calibration of the sensors.

There were a number of problems with the data from the SeaBird dissolved oxygen sensor. The calibration varied through the cruise probably due to an aging membrane that did not respond well when sampling anoxic waters. There were many errors in the dissolved oxygen titration data. The Oxygen:Dissolved:SBE channel was removed from some casts because there was insufficient information to enable recalibration. For these reasons the error estimates are larger than usual.

It is recommended that when an anoxic plug is used on the dissolved oxygen sensor that note be made in the CTD Daily Log book when the plug is taken off or put on. This information will help in interpreting shifts in calibration and may provide early evidence that the membrane is beginning to fail.

The dissolved oxygen data in the CTD files should be considered, roughly

•	±0.8ml/l from 0 – 75db

•	±0.4ml/l from 75 – 175db

•	±0.1ml/l below 175db

•	The manufacturer considers the sensor unreliable below 1000db

The thermosalinograph temperature and salinity data look fine, but the fluorescence increased markedly through the cruise until the instrument was brushed. The ratio of TSG fluorescence to CTD fluorescence confirms that the data are bad before the cleaning. The fluorescence data were removed from the final files up to the point at which the brushing was done. Frequent brushing is clearly essential for good data.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and read.

Dissolved oxygen data were obtained in individual cast files with flag channel and comments.

There was no nutrient or chlorophyll sampling.

The salinity data were in a spreadsheet but lacked a flag channel, the analyst says that no flags are required.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and errors were corrected in the CTD secondary conductivity and an offset was added to the pressure calibration (+1.2db) based on the history. The fluorometer serial number was added to the TSG configuration file. The resulting configuration file was saved as 2006-33-ctd.con.

##### Conversion of Raw Data

Data were converted using 2006-33-ctd for all casts. File 2006-33-0test.cnv contains only a little surface data and will not be processed further.

A few casts were examined and all expected channels are present. The dissolved oxygen trace has the usual offset between upcast and downcast. For most casts the traces of all variables look reasonable, but for some of the early casts there are serious problems:

- Cast #12 looks bad throughout the downcast and part of the upcast and this was found to be due to the pumps not being turned on. The downcast will not processed. All values in pumped channels should be replaced with pad values in the rosette cast for levels at which the pumps are off. The full file can be put through REVERSE and the upcast used down to 80db, but this was a rosette cast with many stops so the data are likely to be  too corrupted to be useful.

- For casts #17 and 20 all pumped channels look bad for most of the upcasts – the pumps were off for the upcast of #20 but not for #17. These casts were in Effingham Inlet and during #17 the CTD hit bottom. From observations on deck it is believed that bits of wood got into the pumps. For #20 the pumps were turned off deliberately while the CTD was at the bottom, probably to avoid getting chips in the pump, because it was believed that the CTD had touched bottom. The transmissivity does not suggest that this was the case, but spikes in altimetry might have indicated a problem.  For both casts the downcasts look fine. The rosette casts should have temperature, fluorescence SBE:Oxygen:Dissolved and conductivity channels removed.

- For cast #21 the secondary conductivity is shifted relative to the primary conductivity from 80db upwards during the upcast. This will affect the bottle data but should not be a problem for the downcast files. The primary data look ok. The secondary channels are fine for the CTD file but the primary channels should be selected for the rosette cast.

- For cast #22 the secondary conductivity is shifted relative to the primary throughout the downcast but suddenly shifts to expected values partway through the upcast. The pumps were on. The primary data look ok. The primary should be selected for both CTD and rosette files.

Rosette files were converted using a start time of -2s and duration of 5s. As explained above file 2006-0test.ros will not be processed further. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT.

All BOT files were plotted and a few problems were noted, but most concern the casts mentioned above (17, 20, 21). For the bottom bottle of cast #12 there were significant variations, but they appear to be real, not noise.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on 3 casts with (alpha, 1/beta) set to (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03,9) and (0.0245, 9.5). For the primary the best setting was (0.02, 7) and for the secondary the best choice was (0.03, 7).

CELLTM was run using (0.02, 7) for the primary and (0.03, 7) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 11 | 500 | -0.0005 | -0.0003 | -0.0025 | V.High, noisy |
| 36 | 500 | -0.0008 | -0.0003 | -0.003 | V.High, noisy |
| 77 | 350 | -0.0002 | -0.0003 | -0.003 | Very steady, high |
| 82 | 170 | -0.0007 | -0.00035 | -0.003 | Very steady, high |

There is no significant pressure dependence in the conductivity and salinity but there is a lot of noise in the differences in the temperature differences, much more than usual.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The Header Check was run and led to the discovery that 3 of the files had scrambled time/position SeaBird headers. Those headers were fixed using a text editor and were reconverted and put through CLEAN. Another run of Header Check showed that cast #13 lacked a station name, so one was added to the file based on the CTD Daily Log entry.

The Header Summary was run and a few errors were found and fixed in station names. The station name for cast #1 was changed from SI-TEST to SI as that is the standard name for that station and the fact that it was a test is irrelevant to the station name.

The cruise track was plotted and no problems were found; track plots were added to the end of this report.

The surface pressure check was run including a mixed-layer depth calculation based on salinity differences compared to 4db. The average surface pressure is 2.3db, which is a little low for the Tully. A few casts with surface pressure <1db were checked; typically it was found that the CTD sat just above 1db with the pumps off for 30 or 40s and conductivity gradually moved towards expected values for that depth. The pumps were not turned on until after that initial soak and lowering to about 2.5db. The mixed-layer depth is very low in most cases, so the surface salinity samples will not be very useful for calibration purposes.

The altimeter readings from the header were exported to a spreadsheet and a few casts were checked and show that the algorithm worked well.

##### BOTTLE FILE PREPARATION

No salinity analysis sheets were available and there was no flag channel in the salinity spreadsheet. The analyst confirmed that no flags were required and the Calculated Salinity StandDev was < or = to 0.0001. The spreadsheet was edited by adding a flag channel, changing the headers to standard format and removing unnecessary columns and rows. The rosette samples and 4 loop samples were saved separately as 2006-33-sal.csv and 2006-33-tsg-sal.csv. One sample had no sample number and careful checks against rosette sheets and CTD log did not turn up any missing samples so it was dropped from the spreadsheets. It is likely a surface sample given the low salinity value.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

The salinity spreadsheet was saved as 2006-33-sal.csv and then edited. A flag channel and comments were added, header names were changed to standard format and the loop values were removed. (A separate file was created with just loop data to be used later during TSG processing.) The file was then converted to individual SAL files.

Dissolved oxygen files (*.add) had an error in the format, with an extra 0 in the cruise number section. The names were edited to change them to standard format. In the course of trying to merge the ADD files with the CTD files a few other errors were found in the ADD files. A text editor was used to correct sample numbers in 2 files, fix the format where one quality flag was entered and reorder one file to ensure increasing sample number.

The SAL and ADD files were merged with CST files in two steps. (Output: MRG1, MRG3), MRG3 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

DO was plotted against CTD Salinity. The only outlier was near the surface for cast #45 but in an area of high gradient, so it is not clear that there is anything wrong with the bottle.

##### 11. COMPARE

Salinity

COMPARE was run. When data above 100db and a few outliers were excluded the differences between bottles and CTD were quite flat for both pairs of sensors. During 2006-15 it was noted that the differences between the two sensors increased with pressure and it was thought likely that the primary salinity accounted for that. In this comparison there are no data from deeper than 300db so we cannot assess whether there is significant pressure dependence for this cruise.

Three large outliers were investigated:

- Cast #22 sample #97, 100db (From Niskin #3). The bottle differs significantly from the secondary salinity only, so the problem is clearly with the secondary CTD data.

- Cast #23 sample #107, 50db (from Niskin #4). The sample looks much more like CTD salinity at 75db and the rosette sheet says that the sample came from Niskin #3 at 75db, sample #106. The MRG file was adjusted to move the salinity value to sample #106.

- Cast #37 sample #180, 200db (from Niskin #2). The bottle stop was quite short but the local gradient was low and there is little salinity variation at that level. The sample value is higher than any salinity values during the CTD cast, so this does not appear to be a question of faulty labelling. The value was flagged “d”.

- The trend in differences against time have the same slope for both pairs of sensors suggesting the variations probably reflect something other than calibration drift. (See 2006-33-sal-comp1.xls)

COMPARE was run again after fixing the samples from casts #23 and 37. After removing the outliers it becomes a little clearer that there may be some pressure dependence. Two minor outliers identified were:

- Cast 63, sample #252, 125db (from Niskin #2) The CTD data are noisy. Sample is probably ok.

- Cast 74, sample #285, 200db (from Niskin #2). The rosette log does not indicate that there was any salinity sampling for this cast. The CTD data are noisy during this bottle stop and continues to decrease throughout the stop which was unusually long. Note was made in the header that the sample may not be in equilibrium with the CTD, but even with the problems noted the differences were not very large so no flag was assigned to the sample.

A third run of COMPARE used Niskin Bottle # as the reference channel. This showed similar results from Bottles #1 through 4, with the majority of samples being from #2. Two shallow samples from bottles #10 and 16 showed much smaller differences. (See 2006-33-sal-comp1-Niskin.xls)

The primary salinity is high by an average of 0.0144 and the secondary is high by 0.0113.

Dissolved Oxygen

An initial run of COMPARE turned up some errors and plots of DO versus CTD salinity were used to help diagnose the problems. The following were investigated:

- Cast #1- There were two sets of sample numbers assigned for comparison of two analysis processes. To match the sample numbers in the addsamp file and the salinity sampling, the samples numbers for the DO samples were changed from 1 – 10 to 11 – 20 in MRG3. Merges were rerun to create a new MRG file. A note of explanation was put in the headers.

- Cast #16 - Dissolved Oxygen sample #51 was an outlier in COMPARE and in DO vs Salinity plots. It was flagged “c”.

- Cast #17 – Comparison looks bad, but it is known the SBE pumps were not on; the bottle values seem reasonable.

- Cast #18 – The value from the bottom bottle was entered as 6.011. This is clearly wrong and it was changed to 0.011 which is close to nearby values; the sample was flagged “c” with an explanation.

- Cast #19 – The surface value recorded on the rosette log was not entered into the ADD file. It does not appear to be particularly out of line, so this is assumed to have been a typo. The value was entered in the MRG file, but not to the ADD file since analysis details were missing.

- Cast #24 – One bottle was flagged “c” by the analyst because of problems in the titration. It was a severe outlier in Compare so the flag was changed to “d” and the value changed to -99.

- Cast #31 – The value of the bottom bottle was originally entered in the file as 3.349. Replacing it with 2.349 produces better results and is consistent with nearby results. The value on the rosette sheet is entered just as “2”. The value was flagged “c”.

- Cast #63 – Slight outlier in Sal vs DO, no flag assigned.

COMPARE was rerun twice after fixing various items including the above samples; casts #12, 17 and 20 were not included in the final run since either the pumps were not working or were working badly. The results were puzzling. Many of the samples from the early casts show up as outliers. It is known that during casts #17 and 18 anoxic waters were sampled, but there are no obvious problems with casts 1, 3, 14 and 16. The worst data are in the open ocean where no anoxic waters were sampled. While the results of low DO sampling in Saanich Inlet might affect the first few casts, it seems most unlikely to still be a factor at event #16. Moreover, DO values were much lower for cast #18 and 19 then for #1, yet most of the Effingham Inlet data (other than #18 and 19 themselves) are similar to casts from the end of the cruise.  Nonetheless it is noted that the results for casts #1 to 16 taken as a group resemble those of #18 and 19 if a few outliers are removed. So this may mean anoxia is relevant. It just does not seem that the sampling of anoxic bottom waters can account for it.

According to Doug Anderson an anoxic plug was used before the cruise and during the steaming out to open ocean. This may account for the unusual fit. It is also known that the membrane needed repairs after this cruise. Near the end of its life the membrane does not respond as well to anoxic conditions. Given the problems with SBE Dissolved Oxygen, is it possible to recalibrate the data with any confidence? COMPARE was used to examine the data in groups to study this issue.

When casts #21 to the end were used and severe outliers removed the best fit was with the differences between bottles and CTD versus CTD DOX:

CTD-BOT = 1.0671 DOX-CTD + 0.0253 (Fit 1)

The casts in Effingham Inlet (excluding #17 and 20 which had bad CTD data) had the following fit if a few outliers were excluded:

CTD-BOT = 1.1469 DOX-CTD - 0.1373 (Fit 2)

The two casts with anoxic bottom waters in Effingham (18 and 19) had the fit:

CTD-BOT = 1.4888 DOX-CTD - 0.0523 (Fit 3)

The other Effingham Inlet casts (#21 through 29) without anoxic bottom waters had the fit:

CTD-BOT = 1.0547 DOX-CTD + 0.0190 (Fit 4)

Using just casts #1, 3, 14 and 16 (excluding 3 outliers) the fit is:

CTD-BOT = 1.5623 DOX-CTD + 0.0377 (Fit 5)

The Broughton Inlet / QCS casts have the fit:

CTD-BOT = 1.0437 DOX-CTD + 0.0492 (Fit 6)

Finally casts #31 to the end (after Effingham Inlet) have the fit:

CTD-BOT = 1.0664 DOX-CTD + 0.0365 (Fit 7)

which is very similar to Fit #1. Finally the two anoxic casts had fits:

CTD-BOT = 1.5307 DOX-CTD - 0.0772 (Cast #18)

CTD-BOT = 1.4539 DOX-CTD - 0.0126 (Cast #19)

When Fit 1 is compared to Fit 4, the differences are no larger than 0.1ml/l.

Treating the Broughton/QCS casts separately leads to only slight differences, so Fit 1 will be used for casts #21 to the end.

For the early part of the cruise, the difference between Fit 5 and Fit 3 are significant, so casts 1 through 16 must be calibrated separately from casts #18 and 19.

Casts 18 and 19 show significant differences from each other, so there is insufficient information to calibrate them with confidence; therefore, the SBE Dissolved oxygen channel should be removed from the files destined for the archives.

Casts #17 and 20 had bad data during the upcast, but the downcasts are usable. The upcast bottles from that cast were compared with the downcast CTD DO data. We don’t expect a good fit because the downcast data are from a moving CTD when time response is an issue and for cast #17 all bottles had been flagged “c” due to problems noted with the Niskin bottles. But we can get some idea of how close they are to other casts:

CTD-BOT = 1.4314 DOX-CTD + 0.1404 (Cast #17)

CTD-BOT = 1.2293 DOX-CTD + 0.3440 (Cast #20)

Cast #17 is reasonably close to Fit 5 so will be treated with the 1-16 group. However Cast #20 does not look like any other fit so the SBE:DO channel should be dropped since there is insufficient information to enable recalibration.

In conclusion

- Fit 5 will be used for casts 1 through 17.

- DO channel will be removed from casts 18, 19 and 20

- Fit 1 will be used for casts 21 to the end.

(See 2006-33-dox-comp1-fix2.xls.)

A separate run of COMPARE with Niskin Bottle # as the reference channel indicates that there was no problem associated with a particular bottle. (See 2006-33-dox-comp1-Niskin.xls.)

##### 12. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The values found were all on the order of 1s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

Conductivity

Tests were run on 2 casts using shifts between -1.5s and +0.5s and T-S plots were prepared to compare the results. Settings of -0.8s worked best overall for the primary and -0.2s for the secondary, so all casts were put through SHIFT using those settings. (Output *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on 2 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Values from +100 to +150 were tried and the best overall match of features was with a choice of +120 records. In previous use of this instrument settings of +110 to +150 have been used.

All casts were shifted by +120 records. (Output: *.SHFO)

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning pertained to surface data from the upcast of event #49.

For cast #12 the SHFO file was put through REVERSE and then DELETE.

##### 14. Other Comparisons

Previous experience with these sensors – These sensors were used during the following 2006 cruises and the sensors were found to be, on average:

| Cruise # | Primary Sensors | Secondary Sensors |
| --- | --- | --- |
| 2006-13 (July) | +0.005 | +0.0015 |
| 2006-24 (July) | +0.0053 | +0.0009 |
| 2006-15 (Sept.) | +0.0073 | +0.0013 |

The dissolved oxygen sensor was repaired and recalibrated in June 2006 and was used for 2 casts during 2006-13. During 2006-24 2 different calibrations were needed due to anoxic sampling and an offset of -0.04ml/l was applied. During 2006-15 with many bottles there was the usual 2-part calibration with a first correction of  CTD-BOT = 1.0512 DOX-CTD + 0.0599 followed by an offset of -0.3ml/l.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were some excursions from those ranges, but nothing that looks indicative of calibration problems. Much of the deep Juan de Fuca salinity was high, and the salinity at QCS1 was high, but that probably reflects inadequacies in the range data.

Repeat Casts – There were no repeat casts.

##### 15. DETAILED EDITING

It is not obvious which sensors to use. The differences from bottles are large for both, though the secondary is closer, but given the history there are grave doubts about the comparison with bottles. Some problems have been noted in the secondary channels during upcasts, but this does not seem to be a factor in downcasts. In the past the secondary sensors have proven to be better, so those were selected for editing.

The secondary sensors were chosen for the archive. Page plots were produced using (T1,S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

The following casts required heavy editing: 5-14, 31-32, 37, 40-43, 45-46.

All other casts required only light editing. Note was made of the editing details in the headers of the relevant files.

The altimeter reading was adjusted for casts #20 and 22 to reflect that more than 1db of data was removed from the bottom of the cast.

The reversed file for cast #12 was examined in CTDEDIT to see if the data from 75m upwards are useful. As expected it is heavily corrupted by stops for rosettes and a very noisy ascent rate between those stops so that it does not look useful. The CTD data in the rosette files will be retained from 75m upwards.

##### 16. Initial Recalibration

The results of COMPARE are out of step with the history of these sensors with both showing much larger differences than noted in other recent uses. There are notes in the log book about open spigots for one cast. If low salinity water leaked into the bottles this would have the effect of making the CTD salinity look too high. The two surface samples show smaller differences, more like what has been seen in the past. Given these doubts it was decided to do the salinity recalibrations based on the results of 2006-15. When these sensors are used again or recalibrated at the factory this decision can be revisited.

An initial recalibration was run on the MRG and SAM files using file 2006-33-recal1-bot.ccf to subtract 0.0073 from the primary salinity and 0.0013 from the secondary salinity and to apply the following correction to the dissolved oxygen:

CTD-BOT = 1.5623 DOX-CTD + 0.0377  (Casts 1-17)

CTD-BOT = 1.0671 DOX-CTD + 0.0253 (Casts 21 to 88)

COMPARE was then rerun to check that the results were as expected. The results indicate the recalibrations worked well, so the same DO and Salinity calibration was applied to the EDT files. (See 2006-33-dox-comp2.xls and 2006-33-sal-comp2.xls.)

##### 17. Special Fluorometer Processing for Angelica Peña

The COR1 files for casts #2 to 11, 13-16 and 27-52 were put through CLIP to produce files with data to 100db only for the use of A. Peña.

Those files were then bin-averaged (1/4db bins), recalibrated suing 2006-29-recal2.ccf and put through REMOVE to remove extraneous channels. HEADEDIT was run to fix formats and channel names and the final files (FCTD) were saved in a separate directory.

A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data were put through a median filter with fixed width 11, before bin-averaging. Plots indicate the filter worked as expected.

There was no CHL sampling, so no special bottle files were prepared

##### 18. Fluorescence Filter and BIN AVERAGE of CTD files

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

Profile plots were also examined and no problems were found in SBE:DO, Fluorescence and Transmissivity.

##### 19. Final Calibration of DO

The alignment of DO corrected for transit time errors and the first calibration attempted to correct for drift in the sensor calibration. The final correction to DO addresses the poor time response of the sensor and this correction applies only to the CTD files. The bottle files are fine since the CTD stops and allows equilibration if the stops are long enough. We will compare the downcast DO values with those of the bottles, accepting that we are assuming no significant differences between upcast and downcast.

Files were bin-averaged to 0.25m bins for just the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. There is a lot of noise and significant time-dependence in the results.

When all casts are included except for #18 and 19 and differences >1ml/l are excluded the average of the differences indicate that the SBE Dissolved Oxygen is high by 0.06 ml/l. However, when we break the comparison down into groups of casts we find pressure dependence and noisiness in fits varies greatly. This probably just reflects real changes between downcast and upcast which are more significant in Broughton than offshore, for example. An offset of +0.2 looks good for the surface of Effingham Inlet while +0.12ml/l is about right for deep in Broughton Archipelago. The plot of differences against file pair number suggests that the offset should be -0.13ml/l early in the cruise and -0.01 at the end. The average for casts 31 to 78 is -0.03ml/l. In the recent past offsets of -0.03 have been applied to this sensor twice and -0.04ml/l once. For casts 1 to 29 we could use the average offset (-0.12ml/l) or use a pressure-dependent correction. Given the noisiness in the results only the average seems justified.

There is too much noise to justify breaking down the recalibration into small groups, but 2 groups looks defensible. We know that the sensor behaved differently during the first half of the cruise due to its response to anoxia. So for casts #1 through 29 an offset of -0.12ml/l will be applied. For casts #30 through 88 an offset of -0.03ml/l will be applied.

A second recalibration was applied to the downcast files only (AVG, THN1) to subtract 0.12ml/l for casts #1 through 29 and 0.03ml/l for casts 30 through 88. (Output: COR2 and THN2)

A final run of COMPARE was done to ensure the recalibration was done properly and it looks fine with an average difference of +0.0006. During the early casts the average is +0.0015 and for the later ones -0.00005ml/l. (See 2006-33-dox-comp3.xls and 2006-33-comp4.xls.)

##### 20. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts except #12, 18, 19,20 and #22: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The following channels were removed from cast #22: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The following channels were removed from cast #12: Scan_Number, Temperature:Primary, Temperature:Secondary, Salinity:T0:C0, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Fluorescence:URU:Seapoint, Oxygen:Dissolved:SBE, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The following channels were removed from casts #18, 19 and 20: Scan_Number, Temperature:Primary,  Salinity:T0:C0,  Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.MRGREM)

A second SBE DO channel was added and the channels reorder to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

Recalibration of SBE Dissolved Oxygen:

The SBE sensor calibration drifted significantly during the cruise, so

the data were divided into 2 sections for recalibration. Where there were

insufficient data to enable reliable recalibration the channel was removed.

The dissolved oxygen data in the CTD files should be considered, roughly

•	±0.8ml/l from 0 – 75db

•	±0.4ml/l from 75 – 175db

•	±0.1ml/l below 175db

•	The manufacturer considers the sensor unreliable below 1000db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The near-surface values varied from about 60% to 150% with the highest values in the northerly part of the WCVI casts. The lowest values were near shore in the southern part of the WCVI section. Offsetting any suggestion that this variation is due to the inadequacy of the recalibration is the observation of great variability in the later part of the cruise in Queen Charlotte Strait and Broughton Archipelago. Plots were produced to check surface values of DO titrations and SBE fluorescence. Those plots were also used to look for evidence of active mixing. The lowest saturation values were in Queen Charlotte Strait and are associated with low fluorescence and very well mixed surface water. The higher saturation rates are in the archipelago and are associated with high fluorescence and large surface gradients. There were no titrated DO values for the Strait. In the inlets surface titrated DO values are slightly lower than the SBE DO by from 0.2 to 0.5ml/l. It must be kept in mind that comparisons at the surface in high gradient areas will not be very close and this observation is based on only a few samples.

##### 21. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts except #17, 18, 19, 20, 21 and 22 to remove: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.MRGREM)

REMOVE was run on casts #17 and 20 to remove: Scan_Number, Temperature:Primary, Temperature:Secondary, Salinity:T0:C0, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Fluorescence:URU:Seapoint, Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.MRGREM)

REMOVE was run on casts #18 and 19 to remove: Scan_Number, Temperature:Primary,  Salinity:T0:C0,  Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Status:Pump, Descent_Rate, Altimeter and Flag. (Output: *.MRGREM)

REMOVE was run on casts #21 and 22 to remove: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Cast #12 was then edited to enter pad values for all pumped channels from 80 meters downwards.

Notes of explanation were added to the headers of the REM files of casts 12 and 17-22.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together. (Output: CHE1)

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### 22. Thermosalinograph Data

a.) Checking calibrations

There were 2 files containing TSG data. The first contains a second temperature channel, but the second does not. In both the serial number for the fluorometer was added and the fluorometer calibrations fixed and an error in the temperature calibration fixed; the resulting files were saved as 2006-33-TSG1.con and 2006-33-TSG2.con.

The history of the sensor was obtained.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to CNV files using a SeaSoft routine.

The channels converted for cast #1 were: Scan_Number, Temperature:Primary, Temperature:Secondary, Temperature Difference, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian.

The channels converted for cast #2 were: Scan_Number, Temperature:Primary, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian

These files were then converted to IOS HEADER format. The data in the secondary temperature channel of the first cast were bad, explaining why the con file was changed at sea. So the data were reconverted without that channel and the Temperature:Difference channel.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers. The times are said to be in UTC and that is consistent with the initial time noted in the log.

ADD TIME CHANNEL was used to add time and date channels in IOS SHELL format and the output files were named *.ATC. (This step was later repeated with a time offset.)

Time-series plots were produced. Temperature and salinity look reasonable. The fluorescence drops very suddenly late in the cruise.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within 0.3db of 4db and exported to a spreadsheet. A few casts were removed that had no pumped data at 4db.

The TSG files were opened in EXCEL. A plot was made of flow rate and while it is noisy, there is no evidence of serious problems with the flow.

The first file lasts for only 2 hours and there are no overlaps with CTD casts.

Columns were added to the second file with calculations of the median and standard deviation over a 2-minute window (5 records) for both temperature and salinity.This file was then reduced to the times when CTDs were run. Those files were combined in a spreadsheet (2006-33-ctd-tsg-comp.xls).

The positions were compared and did not compare well. The average latitude difference was small, but the variation large and the average longitude difference was 0.0047º. It would appear that the clock of either the TSG or the CTD is wrong and since the log book confirms the CTD times it must be the TSG. Matching 2 CTD positions to the TSG record shows a difference of approximately 25 minutes. This file was saved as

ADD TIME CHANNEL was rerun on the CLN files with the choice of adding 0.4167 hours to each record. Once again the resulting file was opened in EXCEL and reduced to the CTD times and placed in the same spreadsheet as the CTD data. When compared this time the latitudes and longitudes were all within 0.0003º of each other with an average difference of 0.0001º in both latitude and longitude. This file was saved as 2006-33-ctd-tsg-comp.xls. This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are variations in alignment, but they are not systematic. This step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 There was no secondary temperature sensor on this cruise.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Graphs were prepared of these differences versus cast #. There was a lot of scatter for casts in Effingham Inlet and the inlets of Broughton Archipelago where the near-surface gradients were quite high. (See 2006-33CTD-tsg-comp.xls) (The original comparison with the bad time was saved as 2006-33-ctd-tsg-comp-time-wrong.xls.)

When mixed-layer depth was calculated earlier for the CTD casts it was found that only 6 looked adequately mixed to be reliable in calibrating the TSG. The TSG primary temperature was higher than the CTD by an average of 0.21 Cº for those 6 casts; the range was from 0.15 to 0.25Cº. If one of those values is rejected (it had a large standard deviation in the TSG 2-minute window) then the average is +0.19Cº. Picking out a section during which the differences are less variable than in most of the record (casts 33 through 46) the TSG is high by about 0.29Cº. If all data are averaged the TSG is high by 0.39Cº.

The TSG is lower than the CTD salinity by from 0.22 in the casts 33 through 46 section where  variability is low and by 0.26 for the 6 well-mixed casts and again by 0.26 for the best 5 casts among those, based on the standard deviation in TSG Salinity over a 2-minute window.

The ratio of the two fluorometers increased throughout the cruise starting with a range of 2 to 7 during the first day and reaching a value of about 50 on October 16. On October 17 it was about 37 when the TSG fluorometer was brushed. At that point the fluorescence values dropped suddenly. From then on there is little variation in the ratio of the two fluorometers with an average of 1.67 and a range of 1.1 to 3.1 with no obvious trend through the final two days. This is a remarkable example of the value of brushing the TSG fluorometer regularly. During 2006-15 there was no such large variability in this ratio. It did rise near the end of the cruise but was never >4.

- Loop Bottle Comparisons There were 4 loop salinity bottles. The salinity values from the TSG files (using a median over a 2-minute window) were compared with the loop data. The differences were 0.232, 0.226, 0.191 and 0.280 with the TSG low in every case. The standard deviations were much lower for the first two cases than for the latter two. The average of all four gives the TSG salinity low by 0.23. (See 2006-33-sal-loop-tsg-comp.xls.) There was no chlorophyll sampling from the loop.

- Calibration History

The TSG was recalibrated in October 2005. It was used for 6 other cruises but several had recalibration based mainly on the previous history. For 2006-08, -11, -13 and 2006-15 when the temperature was found to be high by 0.22, 0.2, 0.16 and 0.16Cº, respectively. For the 1st, 2nd and 4th of these cruises the salinity was found to be low by 0.20, 0.022 and 0.019 respectively. For 2006-13 the salinity differences were hard to interpret varying from being high by 0.13 to 0.23 with an average of -0.18 and a notable difference between outward and inward legs of the cruise. There were many samples and perhaps we would always see this sort of pattern if we had more loop samples.

Conclusions

Comparing the TSG temperature to the CTD leads to different conclusions depending on which data are considered. The choice of the 5 well-mixed CTD casts that have correspondingly quiet TSG record leads to the conclusion that the TSG temperature is higher than the CTD by 0.19Cº. This compares to 0.16Cº during 2006-15. It is not unexpected that the temperature correction is slightly larger in October than in September because the heating by the ship is larger when intake water temperature is lower. The estimates based on many casts indicate higher corrections are needed, but those don’t seem justified based on the historic records.

The TSG salinity is lower than the CTD by from 0.22 or 0.26 depending on which data are chosen for the comparison. The TSG salinity is lower than the loop samples by 0.23. This is not out of line with the history of the instrument. While slightly higher than the correction for 2006-15 this will be partially due to the larger temperature error.

f.) Editing

The time-series plots were examined and the only problems noted are 1 spike in salinity and a sudden drop in fluorescence. The latter is understood. The fluorescence data up to ~0300 on October 17th are unreliable and should be removed. CTDEDIT was used to clean 1 spike in temperature and salinity and to remove all fluorescence data from the first file and from scans #1-14916 for the second file.

g.) Recalibration

CALIBRATE was used to apply offsets of -0.19 Cº and +0.23psu to Temperature:Primary and Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary, Flag and UPloy0 (flow rate).

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH change the GEOGRAPHIC AREA and add the depth of sampling to the header.

A track plot was produced and added to the end of the report.

The TSG sensor history was updated.

As a final check plots were made of the data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor histories were updated for CTD and TSG sensors.

Particulars:

Hake fishing forced abandonment of casts at LB10 and LB7 and LA7

1. Test cast duplicate OXY samples to compare WIN 98 and WXP AutoOxy Program

12. Pumps not turned on. For the sake of the bottle samples keep Pressure and Transmissivity only from the SBE data.

17. CTD hit bottom and wood chips got in pump. All pumped upcast channels unusable. Spigots on 1 and 3 open, probably also 2 and 4.

18/19. Sampled anoxic bottom waters.

20. Pumps off for upcast. All pumped upcast channels unusable. Samples 82 and 83 both from Niskin Bottle #4.

21. Wrong date on NMEA readout in SeaSave. Secondary data bad on upcast.

22. -1.2 salinity difference noted; corrected itself. Secondary data bad for downcast and most of upcast.

23. NMEA Date correct after restart SeaSave

24. Note in log says: “sample numbers need reversing with titration in file”

25. Bottle #3 did not fire.

31 Bottle order reversed – Niskin #2 at fired at 75db and #3 at 100m

35. Altimeter data lost around 5m above bottom.

Oct. 17 0313: Brushed TSG Fluorometer

78. Duplicate OXY samples; comparison cast WIN98 to WXP AutoOxy Program

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2006-33 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 10 Oct. 2006 End: 20 October 2006 |  |  |  |  |  |  |
| Location: WCVI |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550        Cruise ID#:		2006-33

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 18May06 | Factory “ |  |  |
| Conductivity | 2399 | 23May06 | “ |  |  |
| Secondary Temp. | 2038 | 07June05 | “ |  |  |
| Secondary Cond. | 2754 | 23May06 | “ |  |  |
| Transmissometer | 723DR | 28March06 | Factory |  |  |
| SBE 43 DO sensor | 47 | 10June2006 | Factory |  |  |
| Fluorometer | 2228 |  | IOS |  |  |
| Altimeter | 1252 |  |  |  |  |
| Pressure Sensor | 75636 | 04/06/1999 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2006-33

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 21/10/05 | Factory |  |  |
| Temperature2 | 2416 | 5/08/2004 |  |  |  |
| Conductivity | 2487 | 21/10/05 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Flow Meter | ? | ? | ? |  |  |