## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 23 March 2007 | Error in DO recalibration corrected |

## PROCESSING NOTES

Cruise: 2006-35

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG - JdeF

Party Chief: Masson D.

Platform: Vector

Date: November 28, 2006 – December 2, 2006

Processed by: Germaine Gatien

Date of Processing: 27 February 2007 – 17 March 2007

Number of original CTD casts: 71

Number of CTD casts processed: 68 (1 test cast, 1 cast was run 3 times – only 1 processed)

Number of rosette casts: 23

Number of rosette casts processed: 20 (1 cast was run 3 times – only 1 processed; 1 file produced but no sampling so not processed)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#953DR), a PAR sensor (#4656) and an Altimeter OA-916D (#1252). A SeaBird dissolved oxygen sensor (#1117) was mounted on the primary pump and a Seapoint Fluorometer (#2845) with a 10X cable was mounted on the secondary pump. The deck unit is unknown. The pumps were #2467 and 2454. The salinometer was an 8400B model Autosal (S/N 68572).

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book and rosette sheets were available. Both include helpful comments on problems encountered. The dissolved oxygen and salinity bottle data were received without quality flags and comments. Flags were added to the dissolved oxygen data based on comments on the rosette sheets.

The BL file for cast #35 was prepared by hand without a record of the exact time of bottle firing, so the CTD data are from the approximate time of firing and should be treated with care. Th

The stops for rosette sampling were generally very short, with intervals of ~10s between the stop and firing time. This may explain why the comparison between bottles and CTD salinity and dissolved oxygen look anomalous. The results of cruise 2007-01 were used to calibrate this data.

The SBE dissolved oxygen sensor was new and performed well with some detail in the profile and reasonably good correspondence between downcast and upcast. As mentioned above the calibration was based on 2007-01, which introduces some uncertainty. The SBE dissolved oxygen data in the CTD files should be considered, roughly:

- ±0.4ml/l from the surface to 150m

- ±0.2ml/l below 150m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained. The Daily Log includes useful comments on problems encountered including freezing of equipment and problems with rosette firing during event #35. It is also mentioned that the NMEA time did not match the computer clock. A preliminary check shows the log times in agreement with the header times in the data files. During cast #30 there was a problem with GPS positions but this resolved itself on start-up.

The nutrient, chlorophyll and salinity data were each obtained in spreadsheet format with flag and comments columns.

Individual dissolved oxygen bottle files were obtained with flag channels and comments.

The CTD calibrations were checked. The date of the secondary conductivity calibration, the offsets of the transmissometer and PAR sensor calibrations and the serial number for the dissolved oxygen sensor were wrong. Those errors were corrected and the pressure offset was changed to 3.8db to reflect recent cruise results; the resulting configuration file was saved as 2006-35-ctd.con. The dissolved oxygen sensor was a new one. The entry in the log book gives its serial number as 117, but it is believed to have been #1117, though the calibrations in the con file used at sea do not correspond to those of any sensor on file.

The cruise summary sheet was completed.

The histories of the conductivity and pressure sensors were obtained.

##### Conversion of Raw Data

The files were converted to CNV files. A few files were checked and all data look reasonable. There is some noise in transmissivity that looks suspicious but this is probably all in upcasts. There are 3 files for cast #35 but the log book indicates which one should be used, 2006-35-0035, not 2006-35-0035a or 2006-35-0035b. There is also a test file that contains no useful data and will not be processed further.

As usual for this area the descent rate was highly variable; it was very steady for most of the casts but noisy near the mouth of Juan de Fuca Strait.

Bottle stops typically lasted for 10 to 15s, much lower than the recommended 30s. In some cases this may make little difference because the descent rate was quite steady and the gradients low.

The fluorescence looks ok; the minimum values vary from ~0.05ug/l to ~0.15ug/l.

The altimetry is mostly less noisy than usual, but as noted in the log there were often big spikes just when the data were needed, near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format.

CLEAN was used to add event numbers and the extensions were changed to BOT.

All BOT files were plotted and a few significant outliers were found in the secondary salinity for cast #32. The bad data were cleaned using CTDEDIT and the output copied to BOT. Note was made in the headers of the editing done.

Header Check was run on the BOT files and no errors were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Three casts with a steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5) were tried. No choices improved the data, possibly because the temperature gradients are unusually low. This step was skipped.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 17 | 325 | -0.0003 | +0.00025 | +0.0027 X Noisy | Steady but low |
| 43 | 325 | ~0 X Noisy | +0.00021 | +0.0021 Very Noisy | Very steady |
| 57 | 325 | -0.0004 | +0.00018 | +0.0020 Very Noisy | Very steady |
| 58 | 325 | -0.0004 | +0.00024 | +0.0026 | Very steady |
|  | 390 | -0.0004 | +0.00020 | +0.0024 | Very steady |

There is no suggestion of temporal drift. There is also no evidence of pressure dependence.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and no errors were found.

The header summary was checked against the log book entries and no errors were found.

Track plots were produced and added to the end of this report.

The average surface pressure is ~2.1db which is normal for the Vector.

At this point it was realized that the altimetry data had not been converted. Separate CNV and ROS files were created with just altimetry and pressure and these were converted to IOS format and put through CLEAN to add event numbers. From these files the altimeter values were exported to spreadsheets; a few casts were examined and the algorithm was found to have worked well.

A text editor was used to insert the altimetry information from the headers of the files to the CLN files for the downcast data and to the MRG files for the bottles.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. This was saved as addsamp.csv. There is only one bottle in the file for cast #36 and it is from the surface when the pumps were not operating.

For cast #35 there is no BL file, so one was created based on plotting the upcast data and picking out the scan number just before the CTD headed upwards again. This was checked against the rosette log sheet. This will at best be a rough approximation. Conversion of the rosette file was successful and the addsamp file was prepared again. There is a bottle file for cast #36 but neither the log book nor the analyses show any sampling and the pumps were off, so that line was removed from the addsamp file.

The addsamp file was then converted to CST files to form the framework for the bottle files. Sample numbers were added to the BOT files which were then bin-averaged (SAMAVG) on bottle number.

The salinity data were provided in EXCEL format but lacked flag and comment channels. The channels were added and the file saved as 2006-35-sal.csv. There are no notes on the salinity analysis sheet to indicate that any flags or comments should be entered. There were no event numbers and sample numbers were combined with station names and Niskin bottle # in a single column; those were separated.  The files were converted to individual SAL files.

The extracted chlorophyll data were in spreadsheet format and included a flag channel and comments. The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The resulting file is 2006-35chl.csv. That spreadsheet was converted to individual files. (Output: CHL)

The dissolved oxygen data were received in individual OXY files without flag channel and comments. One of the files had an error in the file name, so that was fixed. ADD CHANNEL was used to add the flag channel and flags and comments were added based on those in the rosette log sheets. Errors were noted in one sample number in cast #5 and two in cast #40; these were fixed. A value of 0 was entered for two samples (#61 and 213) because the probe was not in; those samples were rerun but were out of order, so the zero values were removed and were replaced by the values from the second run. Notes about bubbles in the flask or on the lens were found for 2 samples (106 and 129), so “c” flags were entered for those. Those flags should be reviewed after the comparison with CTD data.

The nutrient data were provided in spreadsheet format with flag channel and comments. The samples were stored cool and analyzed within 1 week. Channel names were changed to standard format and lines with no sampling were removed; the data were reordered on sample_number and the file was saved as 2006-35nuts.csv. The data were converted to individual files NUT files.

The SAL, CHL, ADD and NUTS files were merged with the CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4.) The MRG4 files were cleaned to reduce the headers to File and Comments. (Output:MRGCLN1) and those files were merged with the SAMAVG files to produce MRG files. A note was added to the header of 2006-35-0035.MRG to indicate that the CTD data should be treated with care there was no BL file from that cast, so exact bottle firing times are not known.

As mentioned in section 9, the altimetry header information was inserted into the MRG headers.

The MRG files were put through CLEAN, to remove SeaBird headers and comments from the secondary files. (Output: MRGCLN2)

##### 11. COMPARE

Salinity

COMPARE was run and when samples from above 150m and differences >0.02 were excluded a flat trend line was achieved for both salinity channels when plotted against time or pressure. The following significant outliers were identified:

- Cast #19, sample #68, Niskin #2: Short stop with shed wake corruption. CTD looks like it settled down, but possibly Niskin contents affected. Flagged “c”.

- Cast #28, sample #94, Niskin #2: CTD data look ok. Flagged “c”.

- Cast #35, sample # 127, Niskin #2 – this is the cast with the fabricated BL file, so the CTD data must be considered approximate. Flagged “c”.

Both salinity channels were found to be low, the primary by 0.0059 and the secondary by 0.0036, and there is no significant pressure dependence. There is no suggestion of time dependence.

Dissolved Oxygen

Note: After running COMPARE on Salinity and DO it was found that data from above 50m were missing from the SAM file for cast #8. The preparation of this file and the MRG file for that cast were rerun. This will not affect the salinity comparison, but there was DO data missing. A quick check was made of that cast and the only outlier was one sample below 50db.

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. A few outliers were noted.

COMPARE was run and outliers removed judging by plots of differences against pressure, CTD DO and file pair number. The fit is much flatter than we are used to seeing, but there is a lot of scatter.

The following significant outliers were identified:

- Cast 2, samples 8 and 10: The first is an extreme outlier and the second a mild outlier and in the DO vs SAL plots. Flag the first “d” and the second “c”.

- Cast 8, sample #32: also outlier in DO vs SAL plot. Flagged “d”.

- Cast 28, sample #106 & 107 & 108: mild outliers in COMPARE, DO vs Sal plot ok, CTD noisy. No flags assigned but #106 was already assigned a flag based on log notes, so that will be left.

- Cast 32, sample #129: Flag assigned based on log notes, but looks fine in COMPARE so removed.

- Cast 35, sample #137: CTD data very noisy. No flag assigned.

- Cast 37, sample #144 &152: Mild outliers. DO vs Sal ok. CTD data noisy. No flag assigned.

- Cast 40, sample #160: CTD data looks bad. No flag assigned.

- Cast 42, sample #172: Mild outlier, DO vs Sal plot odd at bottom. Flagged “c”

- Cast 45, sample #189: Mild outlier, DO vs Sal plot ok, CTD noisy. No flag assigned.

- Cast 50, samples 216 and 217: Look ok in DO vs SAL plot. CTD dissolved oxygen data are clearly bad during part of the upcast so bottle probably ok. No flag assigned.

- Cast 67, sample #250, 252 & 253: 252 is clearly an outlier in COMPARE. The other two are mild outliers. DO vs Sal plots make all three a bit odd, but range is small and gradients large so difficult to judge. CTD data are noisy for these bottles (especially 253) with short stops that may also compromise the bottle data. Apply “c” flags as warning, but not clear which data are at fault.

Outliers were removed on the basis of residuals and while fits against CTD DO, pressure and event number were all quite flat, that against pressure looks flattest with an offset of ~-0.047ml/l.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. When CHL was plotted against fluorescence the fit was noisy with the trendline:  SBE Fluorescence = 0.6898 * Titrated CHL + 0.0173

The values are very low, so this fit is probably of no significance.

##### Other Comparisons

Previous experience with these sensors –

- The primary conductivity sensor has been recalibrated since the last use. The secondary has been used for 5 cruises but for two of them there were only 1 or 2 casts, and they were mounted on a different CTD system. There is some suggestion of pressure dependence below 2000db. The salinity was considered high by 0.0053, 0.0067 and 0.0073 during 2006-24, 2006-14 and 2006-15. The latter cruise included a lot of deep sampling. The data from this sensor have never been selected for the archive.

- This DO sensor is new.

- The pressure sensor has been used many times; the calibration is known to be drifting.

Results from later cruise –

- The primary sensors were used as primary sensors during 2007-01 which had many salinity samples and deep casts. The salinity was found to be high by 0.0035 with no significant temporal or pressure dependence. This suggests that the bottle values from this cruise are too high by about 0.01. Possible explanations for that are poorly sealed bottles, problem with the salinometer, poor flushing of Niskin bottles. Poor flushing could be related to the very short bottle stops. The short stops would also affect the CTD data, but it would probably settle down more quickly.

- For dissolved oxygen sensor #1117, the best fit during 2007-01 in February 2007 was against CTD DO values though the fit against pressure was good as well. However it was not flat with pressure, so the following fit was chosen for 2007-01:

CTD-BOT = 1.0107 DOX-CTD - 0.0142

The average difference of points in that fit gave the CTD values as high by 0.55ml/l. This is quite different from the results of 2006-35.

Why are the results of 2007-01 so different from 2006-35. Is the problem short bottle stops?

The following plot shows the data from one randomly selected bottle stop. The bottle was fired at about scan #28890, just 6 seconds after stopping. Immediately after stopping the salinity increases sharply to >30.1, presumably due to a shed wake carrying water from 6 or 7m below. The salinity then settles down quickly though there is still some variability at firing time. Dissolved oxygen goes down a little later, having a longer response time. If salinity with a value of >30.1 entered the bottle it would be likely to persist longer and might well lead to anomalously high values in bottles.  Poor flushing of Niskin bottles would also lead, on average, to anomalously low Dissolved Oxygen values in the bottles. Yet we see that the CTD DO values look lower than the bottles, not higher.  However, the picture is not clear with DO because of the CTD DO poor response time for these sensors. The CTD Dissolved Oxygen values would probably be even lower than that of the bottles.

Examining a few bottle stops in a high DO gradient, the CTD DO was found to be lower by about 0.1ml/l at the time of firing compared to 10s later. (In low DO gradients this error would be much lower.) If we add 0.1ml/l to the CTD values to allow for time response errors in high gradient zones, we would then estimate that the CTD DO is higher than the bottles by about 0.05ml/l rather than lower by 0.05ml/l as found in COMPARE. If there is poor flushing of the bottles the difference could be a little higher. For 2007-01 it was found that the CTD was higher than the bottles by about 0.07ml/l near the surface and by 0.03ml/l at 300m. While this is hardly a proof, it does seem that the results of 2006-35 are not inconsistent with 2007-01 when the short bottle stops are accounted for.

Historic ranges – There are many excursions from the historic ranges with temperatures and salinity high near the bottom in the northern Strait of Georgia. Juan de Fuca Strait has some excursions, mostly temperatures at depth being a little high on the south side. Salinity was high in the top 90m on the northern side of the mouth of Juan de Fuca. These excursions are not considered indicative of instrument calibration problems since both pairs of sensors behave the same way and since this region is noted for large variations that are not well represented in the climatology used for the plots.

##### 13. SHIFT

Fluorescence

The usual method for finding what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset.

The fluorescence values were low for this cruise, with just a few profiles from around Sechelt proving useful for this calculation. A shift of 1s has been used in most other cruises and looks appropriate for 2006-35. All data were shifted by +24 records. (Output: SHFFL)  A few casts were examined after this step and the shift worked well.

Dissolved Oxygen

This is a new sensor. A variety of advancements were used and +90 records looked best, though it did not suit every feature. This is an unusual cruise with smaller range of values than usual and some very well-mixed casts.

SHIFT was used to advance the DO channel by +90 records (~3.8s).

After this step a few casts were plotted and the results are good; the offset between downcast and upcast features in DO looks much like those for temperature.

Conductivity

Tests were run on 3 casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.5 for the primary sensor and -0.6 for the secondary and SHIFT was run using those settings.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range 10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: none

##### DETAILED EDITING

It was not obvious which sensors to pick. The results of COMPARE do not fit the results of other cruises. This cruise does not include any sampling below 400m and bottle stops were short, so it is believed that the results of 2007-01 are more reliable than the COMPARE results from 2006-35. Only the primary sensors were used for 2007-01 which was 2 months after this cruise; there were many deep bottles and the sensors appear to have performed well. The history of the secondary sensors shows temporal drift and pressure-dependence, so they look less reliable. So, the primary T and S channels were selected for editing.

The following casts required fairly heavy editing: 4-9 and 13-16

All other casts required only lightly editing mostly at the surface and bottom of casts.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

Salinity will be recalibrated by applying an offset of -0.0035 based on the results of cruise 2007-01.

Dissolved oxygen will also be recalibrated using the results of 2007-01 as follows:

CTD-BOT = 1.0107 DOX-CTD - 0.0142

After calibration of the bottle files COMPARE was rerun. The results are as expected.

The edited downcast files, EDT, were recalibrated using 2006-35-recal1.ccf. (Output:COR1)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the filtered files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time, and the comparison with titrated samples those due to drift in calibration. But there remains an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values. It must be kept in mind that the bottle values are believed to be a little low due to the short bottle stops as explained earlier. This will most affect the bottles in high gradient zones.

The FIL files were bin-averaged (0.25db bins) & thinned to the usual bottle levels. Then CTD DO values were compared with the upcast bottle DO values. Plots were made of differences against pressure and DOX and outliers were excluded. There is a lot of scatter in the fit. At depth the differences are very small, <0.02ml/l and the CTD values are generally lower than the bottles. This may mean that the CTD is reading slightly low, or since the bottles are mostly from the bottom of the casts any flushing problem would lead to higher values in bottles. Near the surface the average differences are on the order of +0.05 to +1ml/l but when points with high standard deviation in the CTD DO are removed it is reduced to about +0.02ml/l with the CTD high. This is consistent with short bottle stops since the bottles would be low if they had not equilibrated. The differences are not large and we lack sufficient information about bottle quality to justify further recalibration. No further recalibration will be applied. (See 2006-35-dox-comp3.xls.)

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

Recalibration was based on the results of cruise 2007-01 when the same

equipment was used.

The SBE dissolved oxygen data in the CTD files should be considered:

•	±0.4ml/l from surface to 150m

•	±0.2ml/l below 150m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~80% to 100% with the higher values in the Strait of Georgia with the exception of one well-mixed cast (#52) with a surface value of ~60%.

##### Final Bottle Files

The MRGCLN files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Descent_Rate and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Note was made that the nutrients were stored “COOL” and analyzed within 1 week.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

The files were named CHE.

##### 23. Producing final files

The following steps were used to look for processing errors:

a) Track plots were produced using the final CTD and CHE files.

b) Cross-reference listings and header checks were produced for the CTD and CHE files.

This led to the discovery that the PAR channel had not been converted for the downcast files and 2 of the bottle files. New CNV and ROS files were created with pressure and PAR channels only and converted to IOS format.

- The downcast files were put through DELETE, bin-averaged and merged with the CTD files. The output files were named CTD and used to replace the previous ones.

- The bottle files were averaged on bottle number, sorted on pressure, merged with the CHE files and then reordered. The output files were renamed CHE and used to replace the previous ones.

The sensor history was updated for the CTD pressure, conductivity and dissolved oxygen sensors.

March 23, 2007. An error in the recalibration of DO was corrected by multiplying all CTD DO data by 0.97883 (0.9893/1.0107). The slope of the correction originally applied was 1.0107 and should have been 0.9893.

May 24, 2007. Study of the post-cruise calibration of the conductivity and temperature sensors shows that the observed drift is compatible with the salinity recalibration applied.

Particulars

30. GPS position problem before casts resolved itself on start-up.

35. Problems with rosette bottle firing – 3 files created. Bottles fired from deck unit.

36. Conversion produced a ROS file but there was no sampling. Log says CTD cast not ROS. No CHE file prepared.

52. Lots of current – large wire angle.

56. Hit bottom

60. T/S increasing near bottom but transmissometer decreased.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2006-35 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 28 November 2006 End: 02 December 2006 |  |  |  |  |  |  |
| Location: SoG/JdeF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Masson D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443         Cruise ID#:		2006-35

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 31/05/06 | Factory “ |  |  |
| Conductivity | 2424 | 23/05/06 | “ |  |  |
| Secondary Temp. | 2668 | 18/05/06 | “ |  |  |
| Secondary Cond. | 2399 | 23/05/06 | “ |  |  |
| Transmissometer | 953DR | 21/03/06 | Factory |  |  |
| SBE 43 DO sensor | 1117 | 17/10/06 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |
| PAR | 4656 | 11/02/03 |  |  |  |
| Fluorometer | 2845 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |