## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 9 Jan. 2007 | 15 missing casts found and processed – see note at end |

PROCESSING NOTES

Cruise: 2006-36

Agency: PBS, SAFE, Nanaimo, B.C.

Project: Sea-Lice Monitoring

Area: Broughton Archipelago

Chief Scientist: Hargreaves B.

Platform: Clupea

Date: 21 June 2006– 29 June 2006

Processed by: Germaine Gatien

Date of Processing: 27 November 2006 –14 December 2006

Number of original CTD casts: 153 (but 15 files missing)

Number of casts processed: 138

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#1294) was used.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available.

The times in the log book are in Pacific Daylight Time. To match the headers in the archived files add 7 hours to the log times.

There was no salinity calibration sampling. No recalibration was applied

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients.

The data for casts 138 through 152 were not available and are believed to have been lost.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard. 
The dates in the log book are in Pacific Daylight Time, so 7 hours need to be added to each to get UTC.

Configuration file 1294.con contains the most up-to-date pre-cruise calibration coefficients for this instrument; all sensors were calibrated in late 2005. The temperature and conductivity coefficients were checked and are correct.

The log sheets contain records of 152 casts; the event numbers were originally named from 0 to 110, then restart at 1 and go to 41. These will be changed by adding 1 to the first 111 casts and 111 to the rest.

- Conversion of Raw Data

- The raw data were converted using conversion file 1294.con. A few casts were checked and no problems were found. The times in the headers appear to be in PST, so 1 hour later than the entries in the log book.

- All files were converted to IOS SHELL format in order to obtain a cross-reference listing to aid in renaming the files. At this point it was discovered that 15 files were missing. Efforts were made to find the data without success; it is believed they are permanently lost.

- The file names were renamed in standard format reflecting the corrected event numbers. (Note that if the missing files are found then they will have to be renamed, and then put through the steps that follow.)

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature. The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

5. ALIGNCTD

Tests were run on a few casts using advancements of from 0.5s to 1.5s. The results varied as is expected when the descent rate of the CTD varies, but +1s gave had the best overall effect.

6. DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

The times are wrong and the station names, bottom depths and positions are missing from the headers. That information is available from the log sheets, so a text editor was used to add details and fix the time.

8. Checking Headers

Track plots and header checks turned up some errors in positions. These were corrected and then a header summary was produced as well as a header check, cross-reference listing and cruise track plots. A few more errors were detected and fixed.

Track plots were added to the end of this report including full cruise track and plots of the casts from either 1 or 2 day periods to make the track easier to see.

The surface check was run. The average surface pressure is +0.2db and the negative values were associated with very low salinity. No adjustment will be made.

9. Test Plots and Fix Time

Profiles were plotted for all casts and up and down traces compared. No problems were noted. ADD TIME CHANNEL was used to add 8 hours to the start times so they are in UTC.

10. SHIFT

Alignctd was used earlier to align temperature. At this point tests were run to see if alignment of conductivity was required. Date was examined on a T-S surface after shifts of conductivity by from -0.7 to +0.2 records. The best overall results were with no shift; values of -0.3 and +0.3 sometimes removed some unstable features, but more often they added instabilities. No shift will be applied.

11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were warnings for casts #18 and 50 that there was less than 10m worth of data. This fits the log record for cast #18, but not for cast #50. There is no indication that data was lost in the processing since the files as originally converted also contain less than 10m worth.

12. CTDEDIT

Downcast page plots were produced and were examined for spikes and instabilities to guide the use of CTDEDIT.  Plots of descent rate were also checked. Some casts required no editing but most needed removal of some records (mostly at the surface) and some cleaning of salinity.

As mentioned in section 11, upcast data was selected for casts 145 – 152.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and no problems noted. No further editing will be done.

14. Intercomparisons

COMPARE - No data available for salinity calibration.

Previous Use of CTD –None since most recent recalibration in late 2005.

Historic ranges – No local climatology available.

15. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The HEADEDIT routine was used to fix the chief scientist’s name and to add the following comment to the headers:

There was no salinity calibration sampling.

Salinity for this cruise has been edited but large errors likely remain.

SeaCat salinity is prone to error due to mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

These errors are likely to be much larger than calibration errors. SeaCat

salinity errors are expected to be as high as 0.05units in areas of high

gradients.

The standards check routine was run and no further problems found.

16. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Particulars:

57. Latitude in log book obviously wrong. Presume should be 50 52.278N.

138-152. Data missing. (FOUND IN JAN. 2007)

153. This cast is not listed in log book, and occurred about 3 weeks after the other casts. Position obtained from Brent Hargreaves’ bridge log and confirmed to be part of this mission.

9 January 2007 – The fifteen missing casts were found and put through all steps.

In the editing process it was found that casts 145 – 152 had poor conductivity data in the top 5 to 15db. The upcast data looks better so those casts were put through REVERSE and then put through DELETE again with output extensions DELREV to distinguish them from the downcast files, *.DEL.

The cruise track maps were updated in the report and at the end of this report.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2006-36

Dates: 	Start:	21 June 2006	    End:	29 June 2006

Location: Broughton Archipelago                 Vessel:    Clupea

Party Chief: Hargreaves B.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/19 SEACAT / 1294             Cruise ID#:		2006-36

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 29 Nov. 2005 | Factory |  |  |
| Conductivity | 1294 | 29 Nov. 2005 | Factory |  |  |
| Pressure Sensor | 163223 | 06 Dec. 2005 | Factory |  |  |