## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2006-38

Agency: Salmon and Freshwater Ecosystems

Location: North-East Pacific

Project: Beamish

Party Chief: Sweeting R.

Platform: W.E. Ricker

Date: July 2, 2006 – July 22 2006

Processed by: Germaine Gatien

Date of Processing: December 13, 2006 – December 16, 2006

Number of original CTD casts: 24   	Number of CTD casts processed: 23 (1 file empty)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Seapoint Fluorometer (#2229) with a 10X cable. Based on what was used during 2006-09 which immediately preceded this cruise, it is assumed that the deck unit was a model 911 (#0471).

# SUMMARY OF QUALITY AND CONCERNS

There was an error in the pressure sensor configuration used at sea; an offset of -0.6db was entered whereas +0.4db would be appropriate.

There was no calibration sampling for salinity but the history of the instrument suggests the salinity data are reliable with errors on the order of ~±0.001.

The fluorescence data are nominal.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

No Log Book was available.

There was no calibration sampling.

The cruise summary sheet was completed. The histories of the conductivity and pressure sensors were obtained. Calibration constants were checked. Errors were found and corrected in the conductivity calibration. As has been found many times for this CTD, the pressure offset was set to -0.6db whereas +0.4db has been found appropriate for the past 2 years, so +0.4db was entered for this cruise as well and the date of the pressure calibration was corrected. The new con file was named 2006-38-ctd.con.

##### Conversion of Raw Data

All data were converted using the con files given above. File 2006-38-0021.dat contains no data.

A few casts were examined and all expected channels are present. The descent rate was high and quite steady for the casts checked. The two pairs of T and C channels look similar and the upcast data are reasonably close to the downcast data. The fluorescence maxima were very high for some casts and look likely to have gone offscale for cast #19. There was no chlorophyll sampling available, but in the past fluorescence values from this type of instrument have been found to be less reliable near the top of the scale.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on four casts using a variety of settings for CELLTM to determine the best choice of parameters. CELLTM was run using (0.03, 7) for the primary and (0.03, 9) for the secondary conductivity for all casts which was also the choice for 2006-09.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The deepest cast was plotted to check for agreement between the pairs of T and C sensors. The following values are rough estimates from downcast data:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2 | 400 | -0.0002 | ~0 X Noisy | +0.0005 Noisy | High, fairly steady |

The temperature differences were small. The conductivity differences were similar to those noted during 2006-22 but much noisier than during 2006-09 which immediately preceded it.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and errors were found in the positions of two casts (5 and 22). For cast #5 the SeaBird headers were scrambled and once those were fixed conversion was rerun. For cast #22 there was no UTC latitude/longitude recorded but there are positions entered in the headers, so these were added to the IOS headers after conversion. The header check looked fine when it was rerun after the corrections.

The header summary was run but no checks could be made as there were no logs available.

The track plot was produced and looks reasonable; it was added to the end of this report.

The average surface pressure is 3.8db which is reasonable.

##### 10. SHIFT

Fluorescence

A few casts were examined to compare the offset between the upcast and downcast fluorescence with that of the temperature traces. That difference divided by the sum of the average descent rate and ascent rate gives an estimate of the shift needed to align the fluorescence. For this cruise estimates varied from 0.9 to 1.7s.  A shift of +24 records (1s) was applied as is usual for this equipment.

Conductivity

For 2006-09 and 2006-22 the primary sensors were selected for archiving and the conductivity was advanced by -0.5s to minimize unstable features in T-S plots. Tests were run on one cast from this cruise and the same setting was found appropriate.

All casts were put through SHIFT to advance the primary conductivity by -0.5s.

##### 11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Low Salt & Last Press Min

Minimum Surface Salinity:      5.000	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### 12. DETAILED EDITING

The primary salinity was chosen for further processing since it has been used for other recent uses of this equipment.

Page plots were produced using (T0, S0). These were used to guide the editing.

On-screen plots of descent rate and pump status were also used to guide editing.

All casts required some light editing except #13 and #14.

Note was made of the editing details in the headers of the relevant files.

##### 13. Other Comparisons

Previous experience with these sensors –

- The primary sensors were used for many cruises during 2005 and for 4 during 2006. Salinity was found to be within 0.002 except for one cruise when it was low by about 0.003 and in most cases it was within 0.001. There was a lot of scatter and few deep bottles for all previous comparisons.

- The secondary sensors were used during 2005-16 and 4 cruises during 2006. There were few bottles and a lot of scatter in the comparisons but for 2006-22 it was believed that the errors were <0.001. Earlier cruises suggest larger errors ~±0.003.

Historic ranges – Profile plots were made of T and S with local climatology superimposed where available. There were minor excursions, mostly with near surface temperature being higher than the historic maximum. Since the ranges do not contain data from the past 3 years, this is not surprising. The salinity data were occasionally higher or lower than the range extremes but there is no systematic difference. These variations do not suggest calibration errors.

##### 14. Initial Recalibration

Based on the results of 2006-09 and 2006-22 recalibration is not justified.

##### 15. Fluorescence Processing

The EDT files were put through a median filter, size 11, applied to the fluorescence channel only. (Output:FIL)

##### 16. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. No problems were noted.

Temperature, Salinity and Fluorescence profiles were plotted on-screen and no problems noted.

##### 17. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

HEADER EDIT was used to fix formats and channel names, to correct the geographic description and to add the following comments:

Fluorescence data are nominal and unedited except that some

records were removed in editing temperature and salinity.

There was no calibration sampling. Based on the history of the

conductivity sensor, salinity data are considered to be reliable,

and likely to be within +/- 0.001.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### 18. Producing final files

A cross-reference listing was produced for CTD files.

HEADER CHECK was rerun and a cruise track plotted and no errors found.

The sensor history was updated for the conductivity and pressure sensors.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2006-38 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 2 July 2006 End: 22 July 2006 |  |  |  |  |  |  |
| Location: NE Pacific |  |  |  |  |  |  |
| Vessel: W.E.Ricker Party Chief: Sweeting R. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0585 | No | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585         Cruise ID#:		2006-38

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19/03/05 | Factory |  |  |
| Conductivity | 3038 | 03/03/05 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 2102 | 07/06/05 | “ |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/2000 | Factory |  |  |