## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2006-39

Agency: PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, B.C.

Location: WCVI

Project: LaPerouse Survey / Sardine Survey

Party Chief: McFarlane G.

Platform: W.E. Ricker

Date: July 25, 2006 – August 7 2006

Processed by: Germaine Gatien

Date of Processing: January 31, 2010 –  February 3, 2009

Number of original CTD casts:  30  	Number of CTD casts processed:  28

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Seapoint Fluorometer (#2229) with a 10X cable.

# SUMMARY OF QUALITY AND CONCERNS

There was an error in the pressure sensor configuration used at sea; an offset of -0.6db was entered whereas +0.4db would be appropriate.

There were salinity calibration samples, but the height of the Niskin above the CTD was not recorded making a comparison with the CTD salinity impossible. Moreover, many of  the deep bottle salinity values are much higher than any CTD salinity in the corresponding profiles. Given that a post-cruise sensor calibration indicates that the primary salinity calibration was excellent, it seems likely there was a problem with the samples or analysis.

Fluorescence data are nominal and unedited except that some records were removed in editing temperature and salinity.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

All file names were non-standard.

The cruise had two legs and the event numbers were repeated. Moreover, there are errors in the log so that the file numbered #6 contains the second attempt at event #5 (at LC5). There is a note in the log explaining this, but there is no note about the fact that the next half dozen files have different contents than the log indicates. There are also 2 files named “2006-07-12” and “Sardine11b” which are not listed in the log, but there is a note saying that 11b is the “right file”. This is going to take very careful checking to make sure the right numbers are assigned.

For details of how files are named and which station they correspond to, see file 2006-39-cast-sample-info.xls. The data from the salinity analysis spreadsheet were added to the spreadsheet. There were no sample numbers in the log or on the salinity analysis sheets, but the correspondence was generally clear because of station names and depths on labels. The only discrepancy between the sampling indicated in the log book and that found in the analysis spreadsheet is that there are 2 samples said to be from the cast at LC2 in the spreadsheet, whereas no sampling is indicated at that site in the log book. On the other hand the log indicates there was sampling at LC7 but no results were found. This seems likely to be the result of misreading of the label with a 7 taken to be a 2. This looks correct since the cast at LC2 is not deep enough to produce a sample at the depth indicated on the sample label. The values were entered as being from LC7.

There are also discrepancies about dates in the headers and the log. It appears that the times were being converted to UTC but mistakes were made in the date in doing that conversion; this is noted in the log in a few cases. The header times and dates make sense.

##### Preliminary Steps

The cruise summary sheet was completed.

The histories of the conductivity and pressure sensors were obtained. Calibration constants were checked. As has been found many times for this CTD, the pressure offset was set to -0.6db whereas +0.4db has been found appropriate for the past 2 years, so +0.4db was entered for this cruise as well and the date of the pressure calibration was corrected. The new con file was named 2006-39-ctd.con.

##### Conversion of Raw Data

All data were converted using configuration file 2006-39-ctd.con.

The pumps did not come on until the CTD was down to at least 5db.

A few casts were examined and all expected channels are present.

The descent rate was fairly high but very noisy.

The two pairs of T and C channels look similar and the upcast data are reasonably close to the downcast data though they are much noisier.

At this point the file names were changed to match those in file 2006-39-cast-sample-info.xls. Each file was opened to ensure that the station name matched what was expected.

The first cast at (LC5) 2009-69-0005.cnv contained only 1 line of data so it will not be processed further.

The two casts at Sardine11 2009-69-0023.cnv and 2009-69-0024.cnv were examined and it is clear that the pumps misbehaved on the first cast, so only the second should be processed further. That is consistent with the log notes that the second cast should be used.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

CELLTM was run using (0.03, 7) for the primary and (0.03, 9) for the secondary conductivity for cruises 2006-09 and 2006-38 when the same equipment was used. Tests were done on a few casts to check that those settings worked well for this cruise as well and they did.

CELLTM was run using (0.03, 7) for the primary and (0.03, 9) for the secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A few casts were plotted to check for agreement between the pairs of T and C sensors and were compared with the results found for cruises before and after 2006-69. The differences were extremely noisy. There were no really deep casts, but it does look as though there is some pressure-dependence in the salinity differences.

| Month | Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- | --- |
| July | 2006-38 | 400 | -0.0002 | ~0 XN | +0.0005 | High, fairly steady |
| July | 2006-39 | 500 | -0.0002 | -0.00012 | -0.001 | High, fairly steady |
| Aug | 2006-26 | 400 | -0.0002 | -0.001 | -0.011 | High, fairly steady |

The results show little change from the cruise before this one. The cruise after had very large differences, but there was a lot of variability between casts suggesting a problem other than calibration drift.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

The station names were in the headers but the bottom depths were not, so those were added using a text editor.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and no errors were found.

The cross-reference list was checked against log records and there was one discrepancy in time and many in date. The dates in the file headers look right as explained in section 1. The time discrepancy is for file #19 at station Sardine7. This looks likely to be a typo in the log book since the header position is midway between Sardine6 and Sardine8.

The track plot was produced and looks reasonable.

The average surface pressure is 1.8db but the corresponding salinity values are very low and the pumps were not on for most casts until the CTD was 3 to 6db deep. For cast #3 the pumps did come on ~1db and the salinity values are low after the pumps have been on long enough to make the values believable. These are not so low as to suggest that the CTD was out of water, but that there was very low salinity surface water for at least some casts. There has been no suggestion of error in this pressure sensor from other cruises using this sensor from 2006, 2007 and early 2008.

##### 10. BOTTLES

The data from the spreadsheet 2006-39-cast-sample-info.xls were simplified and reorganized so there is a line for each bottle fired and saved as 2009-69-salinity.xls.

There is no record of how far above the CTD the Niskin bottle was mounted. To try to determine what the offset was, the CTD salinity channel was checked to see at what pressure the CTD salinity value equalled the bottle salinity. First the deep samples were checked. It is likely that the Niskin bottle was fired when the CTD was stopped at the bottom. If the separation between bottom of cast and apparent depth of Niskin firing were fairly consistent, then we could estimate the offset. Unfortunately this did not work out because, in most cases, the deep bottle salinity values were higher than the maximum value of CTD salinity by 0.2 to 0.4 (and this was looking at full data files before any records were removed). Looking at the surface samples the bottle values correspond to CTD salinity from a depth that could not reasonably be said to be “surface” samples. So either the CTD salinity is notably low or the bottle salinity is much too high.

Based on the history of the sensors there is no reason to believe there is a large error in the CTD sensors, particularly since the sensor pairs are in reasonable agreement. It is more likely there was a sampling or analysis problem. If the bottles were not sealed properly, then evaporation might explain the high salinity values. Given the doubts about the samples and the impossibility of finding a suitable pressure offset to enable a comparison with the CTD salinity, no further work was done to compare them.

##### 11. SHIFT

Fluorescence

A few casts were examined to compare the offset between the upcast and downcast fluorescence with that of the temperature traces. That difference divided by the sum of the average descent rate and ascent rate gives an estimate of the shift needed to align the fluorescence, though corruption of data in the upcast can lead to an estimate that is too high for the downcast. For this cruise estimates varied from 1s to 2s.  A shift of +24 records (1s) was applied as is usual for this equipment.

Conductivity

For 2006-09 and 2006-22 the primary sensors were selected for archiving and the conductivity was advanced by -0.5s to minimize unstable features in T-S plots. A check of a few casts shows that the secondary channels are very noisy, so once again the primary will be selected. Tests were run on one cast from this cruise using -0.5s and -0.8sand -0.5s was found to do a good job of smoothing salinity.

All casts were put through SHIFT to advance the primary conductivity by -0.5s.

##### 12. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Pressure Minimum

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### 13. DETAILED EDITING

The primary temperature and salinity were chosen for further processing since they were selected for other cruises that used this equipment in 2006 and the primary salinity is less noisy than the secondary.

On-screen T-S plots and plots of descent rate and pump status were also used to guide editing.

All casts required some editing.

Note was made of the editing details in the headers of the relevant files.

##### 14. Other Comparisons

Previous experience with these sensors –

- The primary sensors were used for many cruises during 2005 and for 6 during 2006. Salinity was found to be within 0.002 except for one cruise in 2005 when it was low by about 0.003; in most cases it was within 0.001. There was a lot of scatter and few deep bottles for all previous comparisons. A post-cruise calibration indicated that there was no significant drift in the conductivity sensor.

- The secondary sensors were used during 2005-16 and 5 cruises during 2006. There were few bottles and a lot of scatter in the comparisons, but for 2006-22 it was believed that the errors were <0.001. Earlier cruises suggest larger errors ~±0.003. A post-cruise calibration showed a drift of about 0.0001psu per month, so after 13 months, if the drift were linear, the salinity would be low by about 0.0013. This fits the observations of section 7 when the secondary was found to be lower than the primary by 0.001.

Historic ranges – Profile plots were made of T and S with local climatology superimposed; all data fell within the ranges.

##### 15. Initial Recalibration

The calibration sampling was not useful and the history of the instrument suggests that the salinity is good to ±0.001. Recalibration is not justified and none was applied.

##### 16. Fluorescence Processing

The EDT files were put through a median filter, size 11, applied to the fluorescence channel only. One cast was examined before and after, and the filter was found to have worked well. (Output:FIL)

##### 17. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. No problems were noted.

Temperature, Salinity and Fluorescence profiles were plotted on-screen and no problems noted.

##### 18. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

HEADER EDIT was used to fix formats and channel names, to correct the chief scientist’s name and to add the following comments:

Fluorescence data are nominal and unedited except that some

records were removed in editing temperature and salinity.

When the primary conductivity sensor was tested at the factory in

December 2006, no drift in calibration is noted, so the primary

salinity is probably within +/-0.002 except in high gradient regions

where larger errors are likely.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### 19. Producing final files

A cross-reference listing was produced for CTD files.

HEADER CHECK was rerun and a cruise track plotted and no errors found.

The sensor history was updated for the conductivity and pressure sensors.

Particulars (notes from log sheets)

5. Configuration file error – aborted and restarted as 6.

11. Use file 11b.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2006-39 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 25 July 2006 End: 7 August 2006 |  |  |  |  |  |  |
| Location: LaPerouse |  |  |  |  |  |  |
| Vessel: W.E.Ricker Party Chief: McFarlane S. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0585 | No | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585         Cruise ID#:		2006-39

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19/03/05 | Factory |  |  |
| Conductivity | 3038 | 03/03/05 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 2102 | 07/06/05 | “ |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/2000 | Factory |  |  |